## README
item | what it is / how it is computed
WHAT THIS WORKBOOK IS
Set A | The 12 printed test sheets as clean digital images (3200 x 2304 px), picked from the phone's gallery: the app with no camera in the way. 12 images x 2 tracers.
Set B | The same 12 sheets printed and photographed with the app in good indoor light ('ideal'): 24 photographs (2 per sheet) x 2 tracers, each pair from ONE Digitise tap in both-tracers mode (pair_id).
Tracers | learned = b3, the CNN (LiteRT FP16) that reads each row; classical = the threshold-and-centroid control. Everything else in the pipeline is identical for both.
WHAT THE PHONE WRITES (each CSV)
Header lines | '#' lines: device, photo, target sheet, tracer, corners, geometry (grid fit), calibration (px/mm, ticks found of 9, time-zero source), per-stage times (ms), per-stage peak memory, quality verdict and warnings, the classifier's probabilities, pair_id, photo_rotation.
Time series | time_s + 13 columns: I, II, III, aVR, aVL, aVF, V1-V6 and II_rhythm, at 500 Hz, 5,000 rows (10 s), in mV. A 3x4 lead has values only in its own 2.5 s cell (1,250 samples: I/II/III 0-2.5 s, aVR/aVL/aVF 2.5-5 s, V1-V3 5-7.5 s, V4-V6 7.5-10 s); empty elsewhere. II_rhythm has all 10 s. Empty = NaN = the tracer saw no trace there.
How the phone made them | Each traced row (page pixels) is cut on the row's time knots (the printed separator ticks), each pixel column mapped to seconds, then resampled to the 500 Hz grid by linear interpolation (np.interp / SheetLayout.cellSignal). Amplitude: mV = -(row - median row of the cell) / (px per mm x 10 mm/mV); so every cell is about its own median (a printout has no absolute baseline).
WHAT THE SCORER DOES (score_device_sheets.py -> score_12lead.score_lead)
Truth | The PTB-XL record's own 500 Hz signal (the one the kit printed), cut to exactly the same window: for a 3x4 lead the 1,250 samples the kit drew in that cell; for the rhythm strip all 5,000. No resampling of the truth: both are 500 Hz, sample k of the cell against sample k of the truth.
Gaps | Samples the tracer left empty are filled by linear interpolation between their neighbours (run_photo.fill_nan) before every metric except coverage. A lead with < 50 finite samples is not scored at all: it appears with an empty PCC and counts as a FAILURE.
coverage | Fraction of the cell's samples that are finite (before filling).
pcc_asis | Pearson correlation of truth and digitised over the cell, as cut (no shift, no rescale): np.corrcoef. Offset- and scale-invariant: it measures SHAPE AND TIMING, not amplitude.
snr_db_asis | 10 log10( sum truth^2 / sum (truth - digitised)^2 ), both with their own median removed. Sensitive to shape, timing AND amplitude.
shifted_snr_db_asis | The best SNR over integer shifts of -25..+25 samples (+-50 ms at 500 Hz), each scored on the overlapping part only: the PhysioNet Challenge 2024 convention, which forgives a small time offset.
rmse_mv_asis | sqrt(mean (truth - digitised)^2) in mV, AS THE SCORER COMPUTES IT: truth in its raw mV (its own baseline offset kept), digitised about its median. NOTE: this counts the truth window's baseline offset (median ~0.035 mV, up to 0.155) as error, which no digitiser could remove; SNR and the phone's own RMSE remove both medians. See rmse_mv_median_removed.
rmse_mv_median_removed | Computed in THIS workbook only: the same RMSE with both medians removed (consistent with SNR). The fair amplitude error.
gain_ratio | Least-squares slope of digitised on truth: cov(truth, dig) / var(truth). 1.0 = amplitude right; 0.9 = 10% low. Intercept fitted, so a baseline error does not affect it.
r_ratio | R-peak height ratio: truth peaks = local maxima above half the largest deflection (polarity taken from the truth, so a downward QRS is scored on its S wave), >= 0.25 s apart; heights from each signal's median; for each peak the digitised value is its largest same-polarity deviation within +-25 ms; the median of digitised / truth over the peaks.
pcc_aligned / align_* | The best PCC after letting the digitised signal be rescaled in time (x0.90 to x1.10 in 0.25% steps, np.interp) and shifted (+-100 ms for a 2.5 s cell; any lag by FFT cross-correlation for the 10 s strip). It shows what the trace would score with a perfect time base: high aligned PCC with lower plain PCC = a TIMING error, not a reading error. Withheld (empty, 'n.m.') when the aligned overlap is < 85% of the window.
shift_ms / shift_status | The integer lag (+-200 ms) maximising PCC of truth against digitised, medians removed, NaN pairs skipped, >= 85% overlap; status UNRELIABLE if that PCC < 0.5, FAILED at the search bound. A diagnostic of where the time error is.
HOW THE SUMMARIES ARE MADE
Unit | The LEAD: 12 per photo (the rhythm strip is reported separately). Medians and quartiles over leads; a lead with no score is a failure: counted in n, n_unavailable and as 'not >= 0.9', excluded from the median.
95% CI | Of a median: CLUSTER bootstrap over photos (leads of one photo are not independent), 2,000 resamples, seed 20261012 (make_report_figures.cluster_ci).
Learned vs classical | Paired by (photo, lead); Wilcoxon signed-rank on the PCC difference.
Times | Per export, from the header's timings_ms; 'pipeline' = every stage except score* (scoring against the truth is test harness, not the app). Memory: the peak native heap per stage (MB), from memory_peak_mb.
Photo quality (Set B) | Measured on each photograph here: sheet fill (quad area / frame), corner angles' largest departure from 90 degrees (the phone not held parallel), keystone (longer / shorter of top and bottom edges), light ratio (5th / 95th percentile of the paper's smoothed luma; 1 = perfectly even), sharpness (variance of the Laplacian), clipped fraction (pixels >= 250).

## Summary 12 leads
set_name | tracer_kind | photos | leads | leads_unavailable | pct_pcc_ge_0_9 | pcc_median | pcc_q1 | pcc_q3 | pcc_ci95_lo | pcc_ci95_hi | pcc_aligned_median | snr_db_asis_median | shifted_snr_db_asis_median | rmse_mv_asis_median | rmse_mv_median_removed_median | gain_ratio_median | r_ratio_median | coverage_median | pcc_aligned_not_measurable
A (clean images) | classical | 12 | 144 | 0 | 93.1 | 0.9712 | 0.95225 | 0.98443 | 0.9652 | 0.97981 | 0.98785 | 12.4975 | 15.518 | 0.06255 | 0.0356 | 0.9374 | 0.87865 | 0.9712 | 0
B (ideal photos) | classical | 24 | 288 | 14 | 20.8 | 0.6461 | 0.39263 | 0.8843 | 0.5436 | 0.7524 | 0.72005 | 2.005 | 2.408 | 0.1407 | 0.12815 | 0.58905 | 0.5978 | 0.9624 | 0
A (clean images) | learned | 12 | 144 | 0 | 98.6 | 0.994 | 0.99062 | 0.99663 | 0.9922 | 0.99585 | 0.99755 | 19.309 | 22.971 | 0.0458 | 0.0181 | 0.9811 | 0.9712 | 0.9984 | 0
B (ideal photos) | learned | 24 | 288 | 0 | 79.9 | 0.94725 | 0.90882 | 0.9729 | 0.9383 | 0.9561 | 0.9921 | 9.9955 | 15.503 | 0.0698 | 0.0492 | 0.91545 | 0.9406 | 0.9976 | 0

## Summary rhythm strip
set_name | tracer_kind | photos | leads | leads_unavailable | pct_pcc_ge_0_9 | pcc_median | pcc_q1 | pcc_q3 | pcc_ci95_lo | pcc_ci95_hi | pcc_aligned_median | snr_db_asis_median | shifted_snr_db_asis_median | rmse_mv_asis_median | rmse_mv_median_removed_median | gain_ratio_median | r_ratio_median | coverage_median | pcc_aligned_not_measurable
A (clean images) | classical | 12 | 12 | 0 | 100 | 0.97645 | 0.9699 | 0.98325 | 0.9685 | 0.9837 | 0.98585 | 12.819 | 15.4385 | 0.0389 | 0.02965 | 0.9383 | 0.82845 | 0.9992 | 0
B (ideal photos) | classical | 24 | 24 | 1 | 4.2 | 0.6266 | 0.47065 | 0.75165 | 0.5403 | 0.7335 | 0.7756 | 1.355 | 1.726 | 0.1211 | 0.1166 | 0.594 | 0.7167 | 0.9986 | 8
A (clean images) | learned | 12 | 12 | 0 | 100 | 0.99395 | 0.9894 | 0.99538 | 0.9884 | 0.99565 | 0.99395 | 19.227 | 19.227 | 0.0313 | 0.0166 | 0.9827 | 0.96375 | 0.9992 | 0
B (ideal photos) | learned | 24 | 24 | 0 | 54.2 | 0.9047 | 0.87248 | 0.9167 | 0.87525 | 0.913 | 0.9628 | 7.3725 | 7.455 | 0.07105 | 0.06505 | 0.86995 | 0.9344 | 0.9989 | 1

## Tracer comparison
set | n_pairs | statistic | p | median_diff
A (clean images) | 144 | 306 | 1.1251e-22 | 0.01995
B (ideal photos) | 274 | 1026.5 | 6.3316e-42 | 0.26865

## Per photo
set | ecg_id | ptbxl_label | photo_id | tracer | export | photo | photo_path | photo_rotation | corners | geometry | calibration | quality | warnings | pcc_median_12 | pcc_min_12 | leads_pcc_ge_0_9 | leads_unavailable | pcc_aligned_median_12 | snr_db_median_12 | shifted_snr_db_median_12 | rmse_mv_median_12 | rmse_mv_median_removed_median_12 | gain_median_12 | pcc_rhythm | pipeline_ms | trace_ms | peak_native_mb | pss_end_mb | classifier_flags | p_NORM | p_MI | p_STTC | p_CD | p_HYP
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9674 | 0.3354 | 11 | 0 | 0.9907 | 11.9535 | 16.942 | 0.0424 | 0.0322 | 0.919 | 0.9889 | 3777 | 448 | 145.3 | 180.0 | NORM | 0.814 | 0.1506 | 0.0285 | 0.0476 | 0.014
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.99585 | 0.9878 | 12 | 0 | 0.99705 | 20.6555 | 22.031 | 0.0245 | 0.0092 | 0.9849 | 0.9962 | 23855 | 20259 | 506.9 | 478.4 | NORM | 0.9408 | 0.0307 | 0.0167 | 0.0122 | 0.0059
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x37 rules, 10.2780 px/mm, inliers 1.000) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.92065 | 0.8371 | 9 | 0 | 0.94215 | 7.954 | 9.297 | 0.04445 | 0.0416 | 0.8913 | 0.9657 | 2899 | 405 | 148.3 | 205.3 | NORM CD | 0.7174 | 0.0806 | 0.0354 | 0.2348 | 0.0211
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x37 rules, 10.2780 px/mm, inliers 1.000) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.98 | 0.7924 | 10 | 0 | 0.9922 | 14.1025 | 18.3465 | 0.03805 | 0.0317 | 0.947 | 0.9864 | 22287 | 19743 | 518.4 | 539.6 | NORM CD | 0.8147 | 0.0438 | 0.0422 | 0.15 | 0.0239
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.98035 | 0.9507 | 12 | 0 | 0.9922 | 14.2705 | 18.018 | 0.0508 | 0.0316 | 0.94735 | 0.9715 | 2922 | 371 | 146.6 | 231.9 | NORM CD | 0.5944 | 0.0842 | 0.0977 | 0.0926 | 0.0724
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9963 | 0.9913 | 12 | 0 | 0.99825 | 21.5655 | 24.704 | 0.03775 | 0.0115 | 0.9859 | 0.9951 | 22294 | 19752 | 517.4 | 542.4 | NORM CD | 0.6551 | 0.053 | 0.1118 | 0.0918 | 0.0997
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.985) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.95855 | 0.8375 | 11 | 0 | 0.9795 | 11.077 | 14.0315 | 0.09105 | 0.03645 | 0.8937 | 0.9713 | 2929 | 372 | 151.3 | 194.5 | NORM | 0.9827 | 0.0264 | 0.0011 | 0.0017 | 0.0007
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.985) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.99025 | 0.9847 | 12 | 0 | 0.9948 | 17.4495 | 20.1135 | 0.0743 | 0.01805 | 0.9729 | 0.9939 | 22320 | 19765 | 515.8 | 540.8 | NORM | 0.9392 | 0.1173 | 0.0009 | 0.0013 | 0.0007
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.96455 | 0.6226 | 10 | 0 | 0.98725 | 11.4665 | 15.239 | 0.0364 | 0.0307 | 0.92935 | 0.956 | 2838 | 376 | 150.1 | 215.8 | MI | 0.1765 | 0.4823 | 0.3273 | 0.0134 | 0.0175
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.99345 | 0.9829 | 12 | 0 | 0.9972 | 18.897 | 21.908 | 0.0248 | 0.01305 | 0.9829 | 0.9904 | 22378 | 19887 | 518.7 | 539.9 | STTC | 0.0785 | 0.2947 | 0.6339 | 0.0112 | 0.0308
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.97995 | 0.95 | 12 | 0 | 0.9915 | 14.467 | 17.8655 | 0.0824 | 0.04065 | 0.94945 | 0.9869 | 2916 | 378 | 149.4 | 200.9 |  | 0.2526 | 0.0514 | 0.2444 | 0.0107 | 0.4507
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2760 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.99615 | 0.9909 | 12 | 0 | 0.9987 | 21.4855 | 25.7715 | 0.06845 | 0.0175 | 0.99085 | 0.9975 | 22470 | 19928 | 519 | 540.9 | HYP | 0.1433 | 0.0544 | 0.2933 | 0.0081 | 0.6597
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2759 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.97415 | 0.8483 | 11 | 0 | 0.98785 | 13.1205 | 16.1185 | 0.0581 | 0.03115 | 0.92655 | 0.9163 | 2860 | 365 | 147.1 | 197.6 | MI | 0.0139 | 0.9817 | 0.0522 | 0.0099 | 0.0092
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2759 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9949 | 0.9869 | 12 | 0 | 0.9982 | 20.1265 | 24.609 | 0.0505 | 0.0113 | 0.98035 | 0.9862 | 22430 | 19898 | 515.4 | 534.4 | MI | 0.0199 | 0.9742 | 0.0544 | 0.0097 | 0.0096
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2805 px/mm, inliers 0.989) | ticks 10.2815 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.98145 | 0.9395 | 12 | 0 | 0.9902 | 14.1145 | 16.7475 | 0.0467 | 0.03515 | 0.9361 | 0.9806 | 2930 | 378 | 146.2 | 197.5 | MI | 0.014 | 0.9813 | 0.1383 | 0.0322 | 0.0486
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2805 px/mm, inliers 0.989) | ticks 10.2815 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.99505 | 0.9861 | 12 | 0 | 0.9983 | 20.1525 | 24.3615 | 0.0325 | 0.01525 | 0.9796 | 0.9932 | 22543 | 19990 | 516.8 | 539.5 | MI | 0.0371 | 0.9419 | 0.1947 | 0.0303 | 0.0704
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2805 px/mm, inliers 0.989) | ticks 10.2815 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.98905 | 0.9462 | 12 | 0 | 0.9963 | 17.14 | 21.2685 | 0.0535 | 0.02655 | 0.9819 | 0.9828 | 2879 | 384 | 153.9 | 201.3 | MI CD HYP | 0.0001 | 0.9445 | 0.1007 | 0.6296 | 0.5064
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2805 px/mm, inliers 0.989) | ticks 10.2815 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9972 | 0.9825 | 12 | 0 | 0.9988 | 22.581 | 26.4105 | 0.03745 | 0.0142 | 0.9965 | 0.994 | 22439 | 19929 | 519.5 | 535.9 | MI CD HYP | 0.0001 | 0.9498 | 0.1002 | 0.6504 | 0.5124
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2780 px/mm, inliers 0.987) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.98595 | 0.9373 | 12 | 0 | 0.99475 | 15.49 | 19.258 | 0.1026 | 0.0391 | 0.93985 | 0.9816 | 2999 | 389 | 150 | 214.9 | HYP | 0.1077 | 0.1096 | 0.0859 | 0.0269 | 0.901
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2780 px/mm, inliers 0.987) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.99705 | 0.9836 | 12 | 0 | 0.99705 | 22.2055 | 22.2055 | 0.0852 | 0.02295 | 0.98495 | 0.9988 | 22802 | 20169 | 519.3 | 542.6 | HYP | 0.0162 | 0.0855 | 0.1792 | 0.0202 | 0.9881
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2761 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.96895 | 0.907 | 12 | 0 | 0.9839 | 12.0195 | 14.3265 | 0.09265 | 0.0513 | 0.89745 | 0.9723 | 2865 | 383 | 155.4 | 204.5 | STTC HYP | 0.1272 | 0.0882 | 0.5582 | 0.0404 | 0.5293
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x38 rules, 10.2761 px/mm, inliers 0.989) | ticks 10.2762 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9913 | 0.981 | 12 | 0 | 0.9957 | 17.7155 | 20.988 | 0.067 | 0.02635 | 0.97645 | 0.9862 | 22447 | 19984 | 520.8 | 544.5 | STTC HYP | 0.0786 | 0.064 | 0.5758 | 0.0336 | 0.7172
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x37 rules, 10.2791 px/mm, inliers 0.993) | ticks 10.2783 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.95075 | 0.8216 | 10 | 0 | 0.98205 | 10.4285 | 14.213 | 0.085 | 0.04215 | 0.9198 | 0.9846 | 2931 | 376 | 149.8 | 208.6 | CD | 0.0192 | 0.1733 | 0.0068 | 0.9797 | 0.4305
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | picked 3200x2304 |  |  | automatic (refined=true) | grid (fit rms 0.024 mm, 54x37 rules, 10.2791 px/mm, inliers 0.993) | ticks 10.2783 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9887 | 0.9743 | 12 | 0 | 0.99575 | 16.413 | 20.768 | 0.05775 | 0.0223 | 0.97275 | 0.995 | 22343 | 19781 | 518.1 | 543.6 | CD | 0.019 | 0.1279 | 0.0057 | 0.984 | 0.4153
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | picked 2396x3200 | data/kit_photos/setB/eacg_20261013_112807.jpg |  | automatic (refined=false) | grid (fit rms 0.193 mm, 53x37 rules, 7.9816 px/mm, inliers 0.985) | ticks 7.9900 px/mm, 8/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/III | 0.5803 | -0.0778 | 4 | 0 | 0.59595 | 1.779 | 1.8655 | 0.0934 | 0.09185 | 0.43545 | 0.7861 | 2340 | 330 | 131.5 | 199.2 | NORM | 0.6817 | 0.3402 | 0.0317 | 0.0465 | 0.0131
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | picked 2396x3200 | data/kit_photos/setB/eacg_20261013_112807.jpg |  | automatic (refined=false) | grid (fit rms 0.193 mm, 53x37 rules, 7.9816 px/mm, inliers 0.985) | ticks 7.9900 px/mm, 8/9 ticks, t0 from pulses | ok |  | 0.92875 | 0.6886 | 9 | 0 | 0.9887 | 8.5375 | 14.8105 | 0.05505 | 0.03955 | 0.8795 | 0.8759 | 22302 | 19678 | 502.1 | 533.4 | NORM | 0.8705 | 0.1058 | 0.0147 | 0.0137 | 0.0066
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_113507.jpg |  | automatic (refined=false) | grid (fit rms 0.187 mm, 54x37 rules, 7.1879 px/mm, inliers 0.997) | ticks 7.2192 px/mm, 8/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | 0.5473 | 0.2006 | 1 | 1 | 0.5517 | 1.399 | 1.422 | 0.1428 | 0.1454 | 0.3781 | 0.6852 | 1926 | 241 | 124.5 | 223.3 | MI CD | 0.2268 | 0.6977 | 0.0498 | 0.1674 | 0.0163
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_113507.jpg |  | automatic (refined=false) | grid (fit rms 0.187 mm, 54x37 rules, 7.1879 px/mm, inliers 0.997) | ticks 7.2192 px/mm, 8/9 ticks, t0 from pulses | ok |  | 0.92685 | 0.8749 | 8 | 0 | 0.9917 | 8.6305 | 15.43 | 0.0546 | 0.04225 | 0.8885 | 0.896 | 21934 | 19675 | 499.8 | 538.1 | NORM | 0.8163 | 0.1635 | 0.0165 | 0.0163 | 0.0062
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_115407.jpg |  | user | grid (fit rms 0.212 mm, 54x38 rules, 7.6518 px/mm, inliers 0.989) | ticks 7.6560 px/mm, 9/9 ticks, t0 from rhythm_locator | check | NO_PULSE/warn | 0.74625 | 0.3408 | 2 | 0 | 0.8272 | 3.0265 | 3.226 | 0.08825 | 0.08485 | 0.68505 | 0.5403 | 2182 | 232 | 132 | 196.7 | MI CD | 0.1329 | 0.6202 | 0.047 | 0.393 | 0.0632
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_115407.jpg |  | user | grid (fit rms 0.212 mm, 54x38 rules, 7.6518 px/mm, inliers 0.989) | ticks 7.6560 px/mm, 9/9 ticks, t0 from rhythm_locator | check | NO_PULSE/warn | 0.9283 | 0.8519 | 8 | 0 | 0.98445 | 8.6485 | 11.9475 | 0.06395 | 0.0469 | 0.9039 | 0.7536 | 21292 | 19522 | 504.1 | 538.3 | NORM CD | 0.8178 | 0.0495 | 0.0269 | 0.1745 | 0.0169
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_115513.jpg |  | automatic (refined=false) | grid (fit rms 0.216 mm, 55x38 rules, 7.6588 px/mm, inliers 0.979) | ticks 7.6837 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2 | 0.645 | 0.2329 | 0 | 4 | 0.68495 | 1.906 | 2.003 | 0.09345 | 0.0951 | 0.54645 | 0.6153 | 2063 | 218 | 125.5 | 195.1 | MI CD | 0.0314 | 0.9177 | 0.2261 | 0.1128 | 0.0195
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_115513.jpg |  | automatic (refined=false) | grid (fit rms 0.216 mm, 55x38 rules, 7.6588 px/mm, inliers 0.979) | ticks 7.6837 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.938 | 0.8835 | 10 | 0 | 0.97655 | 9.295 | 11.2505 | 0.0549 | 0.04305 | 0.91235 | 0.7638 | 21256 | 19520 | 501 | 538.9 | NORM CD | 0.7661 | 0.1004 | 0.0284 | 0.1747 | 0.0154
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | picked 2396x3200 | data/kit_photos/setB/eacg_20261013_132719.jpg |  | user | grid (fit rms 0.197 mm, 54x37 rules, 7.8156 px/mm, inliers 0.969) | ticks 7.8358 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/III | 0.8299 | 0.3963 | 6 | 0 | 0.84925 | 5.754 | 6.1065 | 0.14415 | 0.1397 | 0.72615 | 0.8695 | 2035 | 231 | 133.1 | 204.4 | CD | 0.3618 | 0.2427 | 0.1698 | 0.1123 | 0.0836
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | picked 2396x3200 | data/kit_photos/setB/eacg_20261013_132719.jpg |  | user | grid (fit rms 0.197 mm, 54x37 rules, 7.8156 px/mm, inliers 0.969) | ticks 7.8358 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.96725 | 0.8914 | 11 | 0 | 0.99395 | 12.2695 | 18.0755 | 0.06735 | 0.04235 | 0.93465 | 0.9289 | 21139 | 19378 | 502.8 | 545.1 | NORM CD | 0.5528 | 0.077 | 0.1329 | 0.1228 | 0.1266
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_133023.jpg |  | user | grid (fit rms 0.210 mm, 54x37 rules, 8.0161 px/mm, inliers 0.966) | ticks 8.0320 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5 | 0.92675 | 0.1227 | 7 | 0 | 0.951 | 8.487 | 9.831 | 0.0876 | 0.06265 | 0.8195 | 0.9054 | 2022 | 237 | 128 | 203.0 | CD | 0.3917 | 0.1301 | 0.1527 | 0.1579 | 0.0902
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_133023.jpg |  | user | grid (fit rms 0.210 mm, 54x37 rules, 8.0161 px/mm, inliers 0.966) | ticks 8.0320 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.95345 | 0.8894 | 11 | 0 | 0.994 | 10.453 | 18.1075 | 0.0662 | 0.03945 | 0.93215 | 0.9045 | 21448 | 19619 | 503.6 | 542.9 | NORM CD | 0.5332 | 0.0815 | 0.1533 | 0.1103 | 0.1241
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_133534.jpg |  | user | grid (fit rms 0.225 mm, 54x36 rules, 7.9346 px/mm, inliers 0.961) | ticks 7.9639 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V4 | 0.7633 | -0.1455 | 3 | 0 | 0.8194 | 3.6575 | 4.8395 | 0.14485 | 0.133 | 0.61725 | 0.6025 | 2048 | 238 | 126.2 | 193.8 | NORM | 0.7191 | 0.297 | 0.0272 | 0.0105 | 0.0036
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_133534.jpg |  | user | grid (fit rms 0.225 mm, 54x36 rules, 7.9346 px/mm, inliers 0.961) | ticks 7.9639 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.89765 | 0.5506 | 6 | 0 | 0.9806 | 7.3115 | 12.0705 | 0.11115 | 0.07605 | 0.82915 | 0.906 | 21738 | 19927 | 502.5 | 540.3 | NORM | 0.987 | 0.0231 | 0.0011 | 0.0019 | 0.001
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_133810.jpg |  | user | grid (fit rms 0.223 mm, 55x36 rules, 7.6481 px/mm, inliers 0.928) | ticks 7.6520 px/mm, 9/9 ticks, t0 from rhythm_locator | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5 | 0.74825 | 0.1752 | 2 | 0 | 0.8418 | 3.128 | 3.288 | 0.15675 | 0.13445 | 0.64285 | 0.2683 | 2117 | 248 | 128.6 | 204.9 | NORM | 0.7311 | 0.1025 | 0.0355 | 0.0106 | 0.0034
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_133810.jpg |  | user | grid (fit rms 0.223 mm, 55x36 rules, 7.6481 px/mm, inliers 0.928) | ticks 7.6520 px/mm, 9/9 ticks, t0 from rhythm_locator | check | NO_PULSE/warn | 0.90815 | 0.6476 | 7 | 0 | 0.9806 | 7.8465 | 10.1945 | 0.11275 | 0.06605 | 0.85955 | 0.8626 | 21355 | 19525 | 502.8 | 544.8 | NORM | 0.9906 | 0.0106 | 0.0013 | 0.0021 | 0.0009
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_145217.jpg |  | user | grid (fit rms 0.173 mm, 55x37 rules, 8.0037 px/mm, inliers 0.990) | ticks 8.0160 px/mm, 9/9 ticks, t0 from rhythm_locator | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6 | 0.46585 | -0.1609 | 0 | 0 | 0.47915 | 0.3415 | 0.7465 | 0.13055 | 0.13085 | 0.3077 | 0.7335 | 2342 | 235 | 133 | 210.6 | MI CD | 0.0175 | 0.9699 | 0.185 | 0.0749 | 0.0259
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_145217.jpg |  | user | grid (fit rms 0.173 mm, 55x37 rules, 8.0037 px/mm, inliers 0.990) | ticks 8.0160 px/mm, 9/9 ticks, t0 from rhythm_locator | check | NO_PULSE/warn | 0.95115 | 0.8509 | 10 | 0 | 0.99185 | 10.0515 | 15.707 | 0.0425 | 0.03925 | 0.9091 | 0.8773 | 21719 | 19760 | 507.3 | 552.9 | STTC | 0.0577 | 0.2818 | 0.6451 | 0.0102 | 0.026
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_145357.jpg |  | automatic (refined=true) | grid (fit rms 0.183 mm, 55x37 rules, 7.7737 px/mm, inliers 0.997) | ticks 7.7838 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/I | 0.74415 | 0.0831 | 3 | 0 | 0.8917 | 3.9065 | 4.5045 | 0.1207 | 0.11515 | 0.64885 | 0.8743 | 2140 | 222 | 131.6 | 200.5 | MI STTC | 0.0661 | 0.4259 | 0.599 | 0.0219 | 0.0336
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_145357.jpg |  | automatic (refined=true) | grid (fit rms 0.183 mm, 55x37 rules, 7.7737 px/mm, inliers 0.997) | ticks 7.7838 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.94605 | 0.8377 | 8 | 0 | 0.99255 | 9.7815 | 17.6675 | 0.0538 | 0.05055 | 0.90475 | 0.8745 | 21373 | 19492 | 508.2 | 548.2 | STTC | 0.0679 | 0.175 | 0.7129 | 0.0104 | 0.0394
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_144637.jpg |  | user | grid (fit rms 0.198 mm, 55x37 rules, 7.6752 px/mm, inliers 0.997) | ticks 7.6877 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 0.42875 | 0.079 | 0 | 0 | 0.4717 | -0.3605 | -0.196 | 0.2419 | 0.2313 | 0.5222 | 0.1067 | 2134 | 216 | 129.7 | 203.8 | STTC CD | 0.1216 | 0.2511 | 0.6411 | 0.0605 | 0.1878
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_144637.jpg |  | user | grid (fit rms 0.198 mm, 55x37 rules, 7.6752 px/mm, inliers 0.997) | ticks 7.6877 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9639 | 0.8498 | 10 | 0 | 0.99435 | 11.801 | 17.188 | 0.0875 | 0.0496 | 0.9329 | 0.9365 | 21151 | 19272 | 506.3 | 547.1 | HYP | 0.1452 | 0.0778 | 0.3108 | 0.0085 | 0.6119
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_145006.jpg |  | user | grid (fit rms 0.176 mm, 54x36 rules, 7.7160 px/mm, inliers 0.993) | ticks 7.7277 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V5 | 0.4559 | -0.0042 | 2 | 0 | 0.5129 | -0.2145 | -0.156 | 0.2036 | 0.1861 | 0.56705 | 0.3175 | 2084 | 227 | 132.8 | 202.5 |  | 0.3588 | 0.1382 | 0.1697 | 0.0445 | 0.1405
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_145006.jpg |  | user | grid (fit rms 0.176 mm, 54x36 rules, 7.7160 px/mm, inliers 0.993) | ticks 7.7277 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.966 | 0.8832 | 11 | 0 | 0.99455 | 12.324 | 17.6555 | 0.09335 | 0.05965 | 0.94215 | 0.876 | 21307 | 19429 | 504.7 | 550.1 | HYP | 0.186 | 0.0638 | 0.2639 | 0.0085 | 0.5861
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | camera 3200x2396 | data/kit_photos/setB/eacg_20261013_134006.jpg |  | automatic (refined=true) | grid (fit rms 0.175 mm, 55x37 rules, 9.0394 px/mm, inliers 0.990) | ticks 9.0570 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.41935 | 0.0654 | 1 | 0 | 0.461 | 0.166 | 0.229 | 0.18955 | 0.17655 | 0.30925 | 0.6428 | 2792 | 312 | 139 | 223.0 | MI CD | 0.0015 | 0.9956 | 0.1467 | 0.0703 | 0.0174
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | camera 3200x2396 | data/kit_photos/setB/eacg_20261013_134006.jpg |  | automatic (refined=true) | grid (fit rms 0.175 mm, 55x37 rules, 9.0394 px/mm, inliers 0.990) | ticks 9.0570 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.95375 | 0.8742 | 10 | 0 | 0.99405 | 10.7585 | 16.6835 | 0.066 | 0.048 | 0.9189 | 0.8542 | 22885 | 20466 | 508.7 | 550.6 | MI | 0.0181 | 0.9741 | 0.0782 | 0.0189 | 0.0152
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | camera 3200x2396 | data/kit_photos/setB/eacg_20261013_134159.jpg |  | automatic (refined=true) | grid (fit rms 0.153 mm, 56x37 rules, 9.9519 px/mm, inliers 0.991) | ticks 9.9860 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V2 | 0.6556 | 0.2955 | 2 | 0 | 0.70875 | 1.6925 | 2.314 | 0.1334 | 0.12335 | 0.61145 | 0.5819 | 3184 | 358 | 144.6 | 215.2 | MI | 0.0032 | 0.9958 | 0.0468 | 0.0246 | 0.0065
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | camera 3200x2396 | data/kit_photos/setB/eacg_20261013_134159.jpg |  | automatic (refined=true) | grid (fit rms 0.153 mm, 56x37 rules, 9.9519 px/mm, inliers 0.991) | ticks 9.9860 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.94655 | 0.7751 | 10 | 0 | 0.993 | 10.099 | 16.2335 | 0.0575 | 0.0461 | 0.91775 | 0.8075 | 23049 | 20226 | 515 | 559.3 | MI | 0.0231 | 0.9656 | 0.0769 | 0.0163 | 0.0145
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | camera 3200x2396 | data/kit_photos/setB/eacg_20261013_134510.jpg | 270 clockwise (auto 270) | user | grid (fit rms 0.137 mm, 55x37 rules, 10.6052 px/mm, inliers 0.985) | ticks 10.6266 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.95725 | 0.5964 | 10 | 0 | 0.98615 | 10.809 | 14.419 | 0.06215 | 0.04505 | 0.9011 | 0.7392 | 3636 | 394 | 149.4 | 229.1 | MI | 0.0157 | 0.9632 | 0.1595 | 0.033 | 0.0414
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | camera 3200x2396 | data/kit_photos/setB/eacg_20261013_134510.jpg | 270 clockwise (auto 270) | user | grid (fit rms 0.137 mm, 55x37 rules, 10.6052 px/mm, inliers 0.985) | ticks 10.6266 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.97525 | 0.8503 | 9 | 0 | 0.99405 | 13.1215 | 18.6835 | 0.056 | 0.0371 | 0.95145 | 0.913 | 22885 | 19746 | 514.4 | 558.8 | MI | 0.0329 | 0.9274 | 0.2012 | 0.0282 | 0.0664
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | camera 3200x2396 | data/kit_photos/setB/eacg_20261013_134642.jpg |  | automatic (refined=false) | grid (fit rms 0.145 mm, 56x37 rules, 10.1656 px/mm, inliers 0.975) | ticks 10.1782 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | 0.86435 | 0.2357 | 5 | 0 | 0.9238 | 6.1435 | 7.745 | 0.08465 | 0.07445 | 0.82555 | 0.6266 | 3354 | 358 | 151 | 220.3 | MI CD | 0.0049 | 0.9922 | 0.1226 | 0.0509 | 0.0522
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | camera 3200x2396 | data/kit_photos/setB/eacg_20261013_134642.jpg |  | automatic (refined=false) | grid (fit rms 0.145 mm, 56x37 rules, 10.1656 px/mm, inliers 0.975) | ticks 10.1782 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9545 | 0.7542 | 10 | 0 | 0.9943 | 10.724 | 17.6155 | 0.05325 | 0.0451 | 0.91115 | 0.9591 | 23126 | 20191 | 515.6 | 556.1 | MI | 0.0284 | 0.9396 | 0.2035 | 0.0293 | 0.0707
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_143255.jpg |  | automatic (refined=true) | grid (fit rms 0.192 mm, 55x37 rules, 8.1442 px/mm, inliers 0.940) | ticks 8.1881 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V4 | 0.81745 | -0.0546 | 3 | 0 | 0.8213 | 4.665 | 4.722 | 0.0979 | 0.0885 | 0.7172 | 0.483 | 2362 | 248 | 129.1 | 223.1 | MI CD | 0.006 | 0.9628 | 0.2523 | 0.2537 | 0.2109
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_143255.jpg |  | automatic (refined=true) | grid (fit rms 0.192 mm, 55x37 rules, 8.1442 px/mm, inliers 0.940) | ticks 8.1881 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.98545 | 0.8535 | 10 | 0 | 0.99605 | 15.4635 | 18.699 | 0.052 | 0.03165 | 0.9642 | 0.9094 | 21784 | 19648 | 504.9 | 549.5 | MI CD HYP | 0.0001 | 0.9487 | 0.1484 | 0.5663 | 0.4839
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_144103.jpg |  | automatic (refined=true) | grid (fit rms 0.213 mm, 56x37 rules, 7.8667 px/mm, inliers 0.962) | ticks 7.8979 px/mm, 8/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3 | 0.73405 | 0.226 | 2 | 0 | 0.8161 | 3.2985 | 3.539 | 0.12835 | 0.12805 | 0.6484 | 0.7194 | 2216 | 226 | 137.4 | 194.5 | MI CD | 0.0036 | 0.986 | 0.2454 | 0.1711 | 0.1351
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_144103.jpg |  | automatic (refined=true) | grid (fit rms 0.213 mm, 56x37 rules, 7.8667 px/mm, inliers 0.962) | ticks 7.8979 px/mm, 8/9 ticks, t0 from pulses | ok |  | 0.98405 | 0.8312 | 11 | 0 | 0.9938 | 15.0255 | 15.53 | 0.0561 | 0.03925 | 0.9591 | 0.9055 | 21571 | 19632 | 508.4 | 553.0 | MI CD HYP | 0.0001 | 0.9333 | 0.1491 | 0.5774 | 0.4976
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162629.jpg |  | automatic (refined=false) | grid (fit rms 0.184 mm, 55x38 rules, 7.8940 px/mm, inliers 0.969) | ticks 7.9239 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm | 0.15125 | 0.0086 | 0 | 2 | 0.18225 | -1.0455 | -0.921 | 0.4709 | 0.45195 | 0.21815 | 0.0572 | 2398 | 278 | 137.2 | 220.8 | MI CD | 0.0131 | 0.9424 | 0.3522 | 0.0876 | 0.0566
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162629.jpg |  | automatic (refined=false) | grid (fit rms 0.184 mm, 55x38 rules, 7.8940 px/mm, inliers 0.969) | ticks 7.9239 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.91295 | 0.7285 | 8 | 0 | 0.98815 | 7.8 | 14.8495 | 0.12515 | 0.1076 | 0.90265 | 0.9049 | 21470 | 19433 | 509.6 | 556.0 | HYP | 0.069 | 0.219 | 0.1798 | 0.0357 | 0.9434
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162743.jpg |  | user | grid (fit rms 0.211 mm, 55x38 rules, 6.9667 px/mm, inliers 0.973) | ticks 6.9870 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | 0.3575 | 0.1729 | 0 | 7 | 0.4153 | -0.372 | -0.353 | 0.3672 | 0.3621 | 0.2207 |  | 2185 | 214 | 136.7 | 215.9 | MI CD | 0.0124 | 0.9334 | 0.2193 | 0.0636 | 0.0282
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162743.jpg |  | user | grid (fit rms 0.211 mm, 55x38 rules, 6.9667 px/mm, inliers 0.973) | ticks 6.9870 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.93575 | 0.7623 | 9 | 0 | 0.98645 | 9.198 | 13.515 | 0.1323 | 0.1071 | 0.898 | 0.9424 | 21376 | 19387 | 506.7 | 555.0 | HYP | 0.0546 | 0.1883 | 0.1456 | 0.0317 | 0.9418
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162400.jpg |  | automatic (refined=true) | grid (fit rms 0.204 mm, 55x37 rules, 7.7869 px/mm, inliers 0.983) | ticks 7.8038 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V4 | 0.3852 | 0.182 | 1 | 0 | 0.4194 | -1.201 | -1.1185 | 0.3334 | 0.31135 | 0.4444 | 0.4583 | 2068 | 230 | 131.5 | 205.6 | MI | 0.1134 | 0.6065 | 0.2897 | 0.028 | 0.019
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162400.jpg |  | automatic (refined=true) | grid (fit rms 0.204 mm, 55x37 rules, 7.7869 px/mm, inliers 0.983) | ticks 7.8038 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.95545 | 0.8591 | 11 | 0 | 0.9867 | 10.4595 | 13.973 | 0.09485 | 0.0621 | 0.91585 | 0.8664 | 21239 | 19391 | 507.8 | 548.3 | STTC HYP | 0.0698 | 0.0798 | 0.6404 | 0.0465 | 0.7014
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162511.jpg |  | automatic (refined=true) | grid (fit rms 0.201 mm, 55x37 rules, 8.0197 px/mm, inliers 0.993) | ticks 8.0440 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.5288 | -0.0106 | 0 | 0 | 0.56325 | 0.9285 | 0.9855 | 0.20405 | 0.2036 | 0.1932 | 0.3075 | 2189 | 242 | 135.2 | 198.0 | MI CD | 0.1055 | 0.7799 | 0.1406 | 0.1712 | 0.0205
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162511.jpg |  | automatic (refined=true) | grid (fit rms 0.201 mm, 55x37 rules, 8.0197 px/mm, inliers 0.993) | ticks 8.0440 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9342 | 0.8736 | 10 | 0 | 0.98245 | 9.034 | 13.3745 | 0.1019 | 0.0718 | 0.88855 | 0.9292 | 21551 | 19695 | 507.1 | 557.7 | STTC HYP | 0.0927 | 0.0879 | 0.6101 | 0.041 | 0.6153
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_161907.jpg |  | automatic (refined=true) | grid (fit rms 0.209 mm, 55x36 rules, 7.3784 px/mm, inliers 0.990) | ticks 7.3994 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.75245 | 0.5677 | 3 | 0 | 0.85555 | 3.5775 | 4.2715 | 0.1165 | 0.1064 | 0.68135 | 0.7641 | 1919 | 203 | 130.8 | 206.5 | CD | 0.0183 | 0.2841 | 0.006 | 0.9658 | 0.2403
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_161907.jpg |  | automatic (refined=true) | grid (fit rms 0.209 mm, 55x36 rules, 7.3784 px/mm, inliers 0.990) | ticks 7.3994 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9423 | 0.9079 | 12 | 0 | 0.9915 | 9.6875 | 15.726 | 0.07975 | 0.0472 | 0.9041 | 0.9185 | 21194 | 19423 | 503.7 | 545.2 | CD | 0.0166 | 0.1332 | 0.0053 | 0.98 | 0.399
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162135.jpg |  | automatic (refined=false) | grid (fit rms 0.201 mm, 54x38 rules, 7.9995 px/mm, inliers 0.974) | ticks 8.0120 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4 | 0.78385 | 0.2845 | 3 | 0 | 0.8711 | 4.3425 | 5.6245 | 0.12075 | 0.11495 | 0.6874 | 0.7903 | 2306 | 239 | 135 | 230.4 | MI CD | 0.0312 | 0.6272 | 0.0137 | 0.8504 | 0.1059
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | camera 2396x3200 | data/kit_photos/setB/eacg_20261013_162135.jpg |  | automatic (refined=false) | grid (fit rms 0.201 mm, 54x38 rules, 7.9995 px/mm, inliers 0.974) | ticks 8.0120 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.95445 | 0.8724 | 11 | 0 | 0.9927 | 10.6635 | 16.2885 | 0.0761 | 0.0392 | 0.9124 | 0.9161 | 21719 | 19648 | 508.1 | 547.1 | CD | 0.0237 | 0.1166 | 0.0049 | 0.9717 | 0.347

## Per sheet
set_name | ecg_id | ptbxl_label | tracer_kind | photos | leads | leads_unavailable | pct_pcc_ge_0_9 | pcc_median | pcc_q1 | pcc_q3 | pcc_ci95_lo | pcc_ci95_hi | pcc_aligned_median | snr_db_asis_median | shifted_snr_db_asis_median | rmse_mv_asis_median | rmse_mv_median_removed_median | gain_ratio_median | r_ratio_median | coverage_median | pcc_aligned_not_measurable
A (clean images) | 9 | NORM | classical | 1 | 12 | 0 | 91.7 | 0.9674 | 0.953 | 0.9826 |  |  | 0.9907 | 11.9535 | 16.942 | 0.0424 | 0.0322 | 0.919 | 0.89955 | 0.9712 | 0
A (clean images) | 9 | NORM | learned | 1 | 12 | 0 | 100 | 0.99585 | 0.9906 | 0.99633 |  |  | 0.99705 | 20.6555 | 22.031 | 0.0245 | 0.0092 | 0.9849 | 0.972 | 0.9984 | 0
A (clean images) | 38 | NORM | classical | 1 | 12 | 0 | 75 | 0.92065 | 0.89807 | 0.9591 |  |  | 0.94215 | 7.954 | 9.297 | 0.04445 | 0.0416 | 0.8913 | 0.882 | 0.9712 | 0
A (clean images) | 38 | NORM | learned | 1 | 12 | 0 | 83.3 | 0.98 | 0.95087 | 0.993 |  |  | 0.9922 | 14.1025 | 18.3465 | 0.03805 | 0.0317 | 0.947 | 0.95625 | 0.9988 | 0
A (clean images) | 40 | NORM | classical | 1 | 12 | 0 | 100 | 0.98035 | 0.97013 | 0.98412 |  |  | 0.9922 | 14.2705 | 18.018 | 0.0508 | 0.0316 | 0.94735 | 0.88605 | 0.9712 | 0
A (clean images) | 40 | NORM | learned | 1 | 12 | 0 | 100 | 0.9963 | 0.9951 | 0.9974 |  |  | 0.99825 | 21.5655 | 24.704 | 0.03775 | 0.0115 | 0.9859 | 0.96545 | 0.998 | 0
A (clean images) | 57 | NORM | classical | 1 | 12 | 0 | 91.7 | 0.95855 | 0.94795 | 0.97 |  |  | 0.9795 | 11.077 | 14.0315 | 0.09105 | 0.03645 | 0.8937 | 0.7677 | 0.9712 | 0
A (clean images) | 57 | NORM | learned | 1 | 12 | 0 | 100 | 0.99025 | 0.98845 | 0.99225 |  |  | 0.9948 | 17.4495 | 20.1135 | 0.0743 | 0.01805 | 0.9729 | 0.92185 | 0.9984 | 0
A (clean images) | 160 | STTC | classical | 1 | 12 | 0 | 83.3 | 0.96455 | 0.91673 | 0.9795 |  |  | 0.98725 | 11.4665 | 15.239 | 0.0364 | 0.0307 | 0.92935 | 0.86035 | 0.9712 | 0
A (clean images) | 160 | STTC | learned | 1 | 12 | 0 | 100 | 0.99345 | 0.99095 | 0.99657 |  |  | 0.9972 | 18.897 | 21.908 | 0.0248 | 0.01305 | 0.9829 | 0.96755 | 0.9976 | 0
A (clean images) | 167 | STTC | classical | 1 | 12 | 0 | 100 | 0.97995 | 0.9687 | 0.98762 |  |  | 0.9915 | 14.467 | 17.8655 | 0.0824 | 0.04065 | 0.94945 | 0.88705 | 0.9712 | 0
A (clean images) | 167 | STTC | learned | 1 | 12 | 0 | 100 | 0.99615 | 0.9948 | 0.99772 |  |  | 0.9987 | 21.4855 | 25.7715 | 0.06845 | 0.0175 | 0.99085 | 0.97065 | 0.998 | 0
A (clean images) | 1737 | MI | classical | 1 | 12 | 0 | 91.7 | 0.97415 | 0.95457 | 0.98628 |  |  | 0.98785 | 13.1205 | 16.1185 | 0.0581 | 0.03115 | 0.92655 | 0.87595 | 0.9712 | 0
A (clean images) | 1737 | MI | learned | 1 | 12 | 0 | 100 | 0.9949 | 0.99155 | 0.99808 |  |  | 0.9982 | 20.1265 | 24.609 | 0.0505 | 0.0113 | 0.98035 | 0.9701 | 0.9984 | 0
A (clean images) | 3908 | MI | classical | 1 | 12 | 0 | 100 | 0.98145 | 0.96505 | 0.98585 |  |  | 0.9902 | 14.1145 | 16.7475 | 0.0467 | 0.03515 | 0.9361 | 0.8952 | 0.9708 | 0
A (clean images) | 3908 | MI | learned | 1 | 12 | 0 | 100 | 0.99505 | 0.99323 | 0.99625 |  |  | 0.9983 | 20.1525 | 24.3615 | 0.0325 | 0.01525 | 0.9796 | 0.9751 | 0.998 | 0
A (clean images) | 4016 | CD | classical | 1 | 12 | 0 | 100 | 0.98905 | 0.98047 | 0.99438 |  |  | 0.9963 | 17.14 | 21.2685 | 0.0535 | 0.02655 | 0.9819 | 0.96 | 0.9708 | 0
A (clean images) | 4016 | CD | learned | 1 | 12 | 0 | 100 | 0.9972 | 0.9944 | 0.99843 |  |  | 0.9988 | 22.581 | 26.4105 | 0.03745 | 0.0142 | 0.9965 | 0.99225 | 0.9984 | 0
A (clean images) | 5830 | HYP | classical | 1 | 12 | 0 | 100 | 0.98595 | 0.97372 | 0.9919 |  |  | 0.99475 | 15.49 | 19.258 | 0.1026 | 0.0391 | 0.93985 | 0.8639 | 0.9712 | 0
A (clean images) | 5830 | HYP | learned | 1 | 12 | 0 | 100 | 0.99705 | 0.99352 | 0.99873 |  |  | 0.99705 | 22.2055 | 22.2055 | 0.0852 | 0.02295 | 0.98495 | 0.9764 | 0.9984 | 0
A (clean images) | 7916 | HYP | classical | 1 | 12 | 0 | 100 | 0.96895 | 0.96275 | 0.97618 |  |  | 0.9839 | 12.0195 | 14.3265 | 0.09265 | 0.0513 | 0.89745 | 0.88165 | 0.9712 | 0
A (clean images) | 7916 | HYP | learned | 1 | 12 | 0 | 100 | 0.9913 | 0.99095 | 0.99362 |  |  | 0.9957 | 17.7155 | 20.988 | 0.067 | 0.02635 | 0.97645 | 0.97015 | 0.9976 | 0
A (clean images) | 8504 | CD | classical | 1 | 12 | 0 | 83.3 | 0.95075 | 0.90868 | 0.97408 |  |  | 0.98205 | 10.4285 | 14.213 | 0.085 | 0.04215 | 0.9198 | 0.871 | 0.9716 | 0
A (clean images) | 8504 | CD | learned | 1 | 12 | 0 | 100 | 0.9887 | 0.98028 | 0.99425 |  |  | 0.99575 | 16.413 | 20.768 | 0.05775 | 0.0223 | 0.97275 | 0.97715 | 0.998 | 0
B (ideal photos) | 9 | NORM | classical | 2 | 24 | 1 | 20.8 | 0.5473 | 0.2501 | 0.8592 | 0.5473 | 0.5803 | 0.5517 | 1.399 | 1.422 | 0.1168 | 0.1179 | 0.3781 | 0.2805 | 0.9632 | 0
B (ideal photos) | 9 | NORM | learned | 2 | 24 | 0 | 70.8 | 0.92745 | 0.89498 | 0.97025 | 0.92685 | 0.92875 | 0.9908 | 8.539 | 14.987 | 0.0546 | 0.04225 | 0.88485 | 0.9383 | 0.9976 | 0
B (ideal photos) | 38 | NORM | classical | 2 | 24 | 4 | 8.3 | 0.73035 | 0.5548 | 0.77873 | 0.645 | 0.74625 | 0.7873 | 2.7535 | 2.8625 | 0.0913 | 0.0921 | 0.64995 | 0.6614 | 0.9472 | 0
B (ideal photos) | 38 | NORM | learned | 2 | 24 | 0 | 75 | 0.9335 | 0.9105 | 0.95493 | 0.9283 | 0.938 | 0.9831 | 9.076 | 11.514 | 0.0625 | 0.04305 | 0.9071 | 0.91745 | 0.9976 | 0
B (ideal photos) | 40 | NORM | classical | 2 | 24 | 0 | 54.2 | 0.9063 | 0.58215 | 0.96008 | 0.8299 | 0.92675 | 0.93365 | 7.452 | 7.697 | 0.11595 | 0.0853 | 0.7867 | 0.7306 | 0.9624 | 0
B (ideal photos) | 40 | NORM | learned | 2 | 24 | 0 | 91.7 | 0.96015 | 0.9447 | 0.97673 | 0.95345 | 0.96725 | 0.99395 | 11.2235 | 18.1075 | 0.06735 | 0.0395 | 0.93465 | 0.93675 | 0.9976 | 0
B (ideal photos) | 57 | NORM | classical | 2 | 24 | 0 | 20.8 | 0.7633 | 0.35225 | 0.8832 | 0.74825 | 0.7633 | 0.83785 | 3.6575 | 4.0055 | 0.15675 | 0.13445 | 0.61725 | 0.60855 | 0.962 | 0
B (ideal photos) | 57 | NORM | learned | 2 | 24 | 0 | 54.2 | 0.90805 | 0.83522 | 0.92275 | 0.89765 | 0.90815 | 0.9806 | 7.7765 | 11.6625 | 0.11245 | 0.0714 | 0.85545 | 0.86395 | 0.9976 | 0
B (ideal photos) | 160 | STTC | classical | 2 | 24 | 0 | 12.5 | 0.5367 | 0.24397 | 0.85435 | 0.46585 | 0.74415 | 0.61955 | 0.7445 | 1.33 | 0.1207 | 0.11515 | 0.49285 | 0.5086 | 0.9632 | 0
B (ideal photos) | 160 | STTC | learned | 2 | 24 | 0 | 75 | 0.94915 | 0.89848 | 0.9743 | 0.94605 | 0.95115 | 0.9923 | 9.8835 | 16.0035 | 0.0438 | 0.0405 | 0.90495 | 0.9405 | 0.9976 | 0
B (ideal photos) | 167 | STTC | classical | 2 | 24 | 0 | 8.3 | 0.43565 | 0.34555 | 0.77443 | 0.42875 | 0.4559 | 0.48025 | -0.2145 | -0.156 | 0.2251 | 0.21525 | 0.56705 | 0.58615 | 0.9616 | 0
B (ideal photos) | 167 | STTC | learned | 2 | 24 | 0 | 87.5 | 0.9656 | 0.9493 | 0.98028 | 0.9639 | 0.966 | 0.99435 | 12.056 | 17.575 | 0.0902 | 0.0548 | 0.93925 | 0.9444 | 0.9976 | 0
B (ideal photos) | 1737 | MI | classical | 2 | 24 | 0 | 12.5 | 0.5315 | 0.31148 | 0.7517 | 0.41935 | 0.6556 | 0.66005 | 1.0825 | 1.578 | 0.1549 | 0.14435 | 0.40435 | 0.3275 | 0.9636 | 0
B (ideal photos) | 1737 | MI | learned | 2 | 24 | 0 | 83.3 | 0.9527 | 0.92483 | 0.97508 | 0.94655 | 0.95375 | 0.9936 | 10.5025 | 16.395 | 0.0613 | 0.0477 | 0.9189 | 0.9375 | 0.9976 | 0
B (ideal photos) | 3908 | MI | classical | 2 | 24 | 0 | 62.5 | 0.92265 | 0.81827 | 0.96065 | 0.86435 | 0.95725 | 0.9809 | 7.9875 | 12.229 | 0.06855 | 0.06155 | 0.8716 | 0.8313 | 0.964 | 0
B (ideal photos) | 3908 | MI | learned | 2 | 24 | 0 | 79.2 | 0.96255 | 0.9088 | 0.98495 | 0.9545 | 0.97525 | 0.99405 | 11.48 | 18.433 | 0.0552 | 0.042 | 0.9306 | 0.95465 | 0.9984 | 0
B (ideal photos) | 4016 | CD | classical | 2 | 24 | 0 | 20.8 | 0.7938 | 0.56657 | 0.88707 | 0.73405 | 0.81745 | 0.8161 | 4.23 | 4.349 | 0.1265 | 0.12695 | 0.665 | 0.8283 | 0.964 | 0
B (ideal photos) | 4016 | CD | learned | 2 | 24 | 0 | 87.5 | 0.9848 | 0.95655 | 0.99065 | 0.98405 | 0.98545 | 0.99505 | 15.238 | 17.264 | 0.05385 | 0.03705 | 0.9634 | 0.9771 | 0.9976 | 0
B (ideal photos) | 5830 | HYP | classical | 2 | 24 | 9 | 0 | 0.1948 | 0.14945 | 0.44925 | 0.15125 | 0.3575 | 0.205 | -0.656 | -0.526 | 0.4551 | 0.4434 | 0.2207 | 0.0793 | 0.59 | 0
B (ideal photos) | 5830 | HYP | learned | 2 | 24 | 0 | 70.8 | 0.9227 | 0.88775 | 0.95932 | 0.91295 | 0.93575 | 0.98815 | 8.076 | 14.171 | 0.12515 | 0.1071 | 0.90205 | 0.9543 | 0.9972 | 0
B (ideal photos) | 7916 | HYP | classical | 2 | 24 | 0 | 4.2 | 0.4588 | 0.33415 | 0.61652 | 0.3852 | 0.5288 | 0.49395 | 0.419 | 0.485 | 0.25645 | 0.2618 | 0.37295 | 0.3447 | 0.9604 | 0
B (ideal photos) | 7916 | HYP | learned | 2 | 24 | 0 | 87.5 | 0.944 | 0.91583 | 0.95995 | 0.9342 | 0.95545 | 0.9842 | 9.7965 | 13.973 | 0.09975 | 0.0658 | 0.9021 | 0.94115 | 0.998 | 0
B (ideal photos) | 8504 | CD | classical | 2 | 24 | 0 | 25 | 0.75245 | 0.64185 | 0.88088 | 0.75245 | 0.78385 | 0.8711 | 3.7685 | 4.9875 | 0.1174 | 0.1074 | 0.68135 | 0.61695 | 0.9608 | 0
B (ideal photos) | 8504 | CD | learned | 2 | 24 | 0 | 95.8 | 0.94775 | 0.92258 | 0.9675 | 0.9423 | 0.95445 | 0.9923 | 10.058 | 15.726 | 0.07615 | 0.04245 | 0.90875 | 0.92545 | 0.9976 | 0

## PCC by lead
set_name | tracer_kind | I | II | III | aVR | aVL | aVF | V1 | V2 | V3 | V4 | V5 | V6 | II rhythm
A (clean images) | classical | 0.9855 | 0.96715 | 0.9384 | 0.97405 | 0.98045 | 0.98325 | 0.977 | 0.9696 | 0.95565 | 0.9666 | 0.97205 | 0.9751 | 0.97645
A (clean images) | learned | 0.99655 | 0.99645 | 0.99555 | 0.99625 | 0.99245 | 0.9945 | 0.99415 | 0.99465 | 0.9911 | 0.9922 | 0.9929 | 0.9921 | 0.99395
B (ideal photos) | classical | 0.6932 | 0.726 | 0.74415 | 0.7902 | 0.88385 | 0.8901 | 0.50055 | 0.61 | 0.4298 | 0.4196 | 0.5594 | 0.5964 | 0.6266
B (ideal photos) | learned | 0.96935 | 0.9542 | 0.92855 | 0.95985 | 0.94875 | 0.9575 | 0.94875 | 0.95565 | 0.9418 | 0.9136 | 0.90435 | 0.95355 | 0.9047

## SNR by lead
set_name | tracer_kind | I | II | III | aVR | aVL | aVF | V1 | V2 | V3 | V4 | V5 | V6 | II rhythm
A (clean images) | classical | 15.3025 | 12.126 | 9.1965 | 13.1085 | 13.85 | 14.4525 | 13.5905 | 12.469 | 10.932 | 12.006 | 12.7195 | 13.134 | 12.819
A (clean images) | learned | 21.8935 | 21.602 | 20.1405 | 21.424 | 18.19 | 19.633 | 19.421 | 19.995 | 17.829 | 18.2695 | 18.772 | 17.7795 | 19.227
B (ideal photos) | classical | 1.852 | 2.7815 | 3.314 | 3.603 | 5.7815 | 7.059 | -0.332 | 1.974 | 0.085 | 0.092 | 1.3635 | 1.543 | 1.355
B (ideal photos) | learned | 12.3115 | 10.6645 | 8.48 | 11.2735 | 9.862 | 10.8355 | 10.073 | 10.5975 | 9.796 | 8.0215 | 7.594 | 10.693 | 7.3725

## Learned vs classical
photo_id | leads | learned_pcc_median | classical_pcc_median | diff_median | learned_better_leads | set
pair_20261013_113100 | 12 | 0.92875 | 0.5803 | 0.1861 | 10 | B (ideal photos)
pair_20261013_113548 | 11 | 0.929 | 0.5473 | 0.413 | 10 | B (ideal photos)
pair_20261013_115454 | 12 | 0.9283 | 0.74625 | 0.2023 | 12 | B (ideal photos)
pair_20261013_115548 | 8 | 0.9315 | 0.645 | 0.2897 | 8 | B (ideal photos)
pair_20261013_132946 | 12 | 0.96725 | 0.8299 | 0.14815 | 11 | B (ideal photos)
pair_20261013_133135 | 12 | 0.95345 | 0.92675 | 0.04185 | 9 | B (ideal photos)
pair_20261013_133702 | 12 | 0.89765 | 0.7633 | 0.1287 | 10 | B (ideal photos)
pair_20261013_133915 | 12 | 0.90815 | 0.74825 | 0.1206 | 10 | B (ideal photos)
pair_20261013_134146 | 12 | 0.95375 | 0.41935 | 0.5217 | 11 | B (ideal photos)
pair_20261013_134236 | 12 | 0.94655 | 0.6556 | 0.2666 | 11 | B (ideal photos)
pair_20261013_134604 | 12 | 0.97525 | 0.95725 | 0.01145 | 8 | B (ideal photos)
pair_20261013_134722 | 12 | 0.9545 | 0.86435 | 0.03735 | 8 | B (ideal photos)
pair_20261013_143426 | 12 | 0.98545 | 0.81745 | 0.16905 | 9 | B (ideal photos)
pair_20261013_144140 | 12 | 0.98405 | 0.73405 | 0.2361 | 11 | B (ideal photos)
pair_20261013_144827 | 12 | 0.9639 | 0.42875 | 0.50825 | 12 | B (ideal photos)
pair_20261013_145129 | 12 | 0.966 | 0.4559 | 0.48535 | 12 | B (ideal photos)
pair_20261013_145325 | 12 | 0.95115 | 0.46585 | 0.47965 | 12 | B (ideal photos)
pair_20261013_145443 | 12 | 0.94605 | 0.74415 | 0.151 | 11 | B (ideal photos)
pair_20261013_162024 | 12 | 0.9423 | 0.75245 | 0.21735 | 9 | B (ideal photos)
pair_20261013_162240 | 12 | 0.95445 | 0.78385 | 0.17905 | 10 | B (ideal photos)
pair_20261013_162442 | 12 | 0.95545 | 0.3852 | 0.5761 | 11 | B (ideal photos)
pair_20261013_162603 | 12 | 0.9342 | 0.5288 | 0.4024 | 11 | B (ideal photos)
pair_20261013_162729 | 10 | 0.9048 | 0.15125 | 0.7062 | 10 | B (ideal photos)
pair_20261013_162839 | 5 | 0.9589 | 0.3575 | 0.6031 | 5 | B (ideal photos)
ptbxl160_clean | 12 | 0.99345 | 0.96455 | 0.0289 | 12 | A (clean images)
ptbxl167_clean | 12 | 0.99615 | 0.97995 | 0.0163 | 12 | A (clean images)
ptbxl1737_clean | 12 | 0.9949 | 0.97415 | 0.02075 | 12 | A (clean images)
ptbxl38_clean | 12 | 0.98 | 0.92065 | 0.02745 | 10 | A (clean images)
ptbxl3908_clean | 12 | 0.99505 | 0.98145 | 0.01325 | 12 | A (clean images)
ptbxl4016_clean | 12 | 0.9972 | 0.98905 | 0.00745 | 11 | A (clean images)
ptbxl40_clean | 12 | 0.9963 | 0.98035 | 0.0154 | 12 | A (clean images)
ptbxl57_clean | 12 | 0.99025 | 0.95855 | 0.0293 | 12 | A (clean images)
ptbxl5830_clean | 12 | 0.99705 | 0.98595 | 0.01065 | 10 | A (clean images)
ptbxl7916_clean | 12 | 0.9913 | 0.96895 | 0.02315 | 12 | A (clean images)
ptbxl8504_clean | 12 | 0.9887 | 0.95075 | 0.02955 | 12 | A (clean images)
ptbxl9_clean | 12 | 0.99585 | 0.9674 | 0.0252 | 12 | A (clean images)

## Timing & memory
set | ecg_id | photo_id | tracer | ms_corners | ms_rectify | ms_grid | ms_locate | ms_calibrate+segment | ms_trace | ms_cut | ms_quality | ms_score* | ms_classify | native_mb_corners | native_mb_rectify | native_mb_grid | native_mb_locate | native_mb_calibrate+segment | native_mb_trace | native_mb_cut | native_mb_quality | native_mb_score* | native_mb_classify
A (clean images) | 9 | ptbxl9_clean | classical | 0 | 179 | 2700 | 319 | 21 | 448 | 34 | 5 | 38 | 71 | 69.6 | 125.7 | 145.3 | 120.3 | 92 | 97.1 | 92.6 | 92.5 | 92.6 | 101.2
A (clean images) | 9 | ptbxl9_clean | learned | 0 | 192 | 2875 | 301 | 47 | 20259 | 46 | 13 | 80 | 122 | 58.1 | 114.9 | 138.4 | 113.3 | 85 | 506.9 | 444.8 | 445 | 447.6 | 454.3
A (clean images) | 38 | ptbxl38_clean | classical | 0 | 190 | 1923 | 309 | 5 | 405 | 10 | 0 | 7 | 57 | 72 | 128 | 148.3 | 123.5 | 95.1 | 102.1 | 95.2 | 95.2 | 95.2 | 104.4
A (clean images) | 38 | ptbxl38_clean | learned | 0 | 203 | 1957 | 311 | 5 | 19743 | 11 | 1 | 7 | 56 | 73.1 | 129.2 | 149.8 | 124.3 | 96.4 | 518.4 | 455.5 | 455.2 | 455.4 | 464.6
A (clean images) | 40 | ptbxl40_clean | classical | 0 | 201 | 1988 | 297 | 4 | 371 | 9 | 0 | 6 | 52 | 70 | 126.1 | 146.6 | 121.8 | 93.4 | 100.5 | 94.1 | 94.1 | 94.1 | 102.2
A (clean images) | 40 | ptbxl40_clean | learned | 0 | 187 | 1965 | 300 | 4 | 19752 | 11 | 1 | 8 | 74 | 72.5 | 128.5 | 148 | 123.2 | 95.2 | 517.4 | 454.4 | 454.2 | 454.2 | 465.5
A (clean images) | 57 | ptbxl57_clean | classical | 0 | 199 | 1974 | 312 | 5 | 372 | 12 | 0 | 7 | 55 | 71.3 | 127.4 | 151.3 | 122.7 | 94.4 | 101.6 | 94.8 | 94.8 | 94.8 | 103
A (clean images) | 57 | ptbxl57_clean | learned | 0 | 183 | 1986 | 313 | 5 | 19765 | 10 | 0 | 8 | 58 | 67.4 | 123.4 | 147.2 | 122.2 | 93.9 | 515.8 | 453.7 | 453.7 | 453.8 | 465.1
A (clean images) | 160 | ptbxl160_clean | classical | 0 | 179 | 1936 | 276 | 6 | 376 | 10 | 0 | 6 | 55 | 73.8 | 129.8 | 150.1 | 125.2 | 97.1 | 103.7 | 98 | 97.9 | 98 | 107.2
A (clean images) | 160 | ptbxl160_clean | learned | 0 | 177 | 1959 | 288 | 4 | 19887 | 10 | 0 | 7 | 53 | 73.8 | 129.7 | 153.5 | 125.1 | 96.8 | 518.7 | 455.6 | 455.6 | 455.6 | 466
A (clean images) | 167 | ptbxl167_clean | classical | 0 | 177 | 1981 | 313 | 4 | 378 | 12 | 0 | 6 | 51 | 73 | 129 | 149.4 | 124.4 | 96.2 | 101.7 | 96.6 | 96.6 | 96.7 | 105.7
A (clean images) | 167 | ptbxl167_clean | learned | 0 | 183 | 1977 | 313 | 4 | 19928 | 10 | 0 | 7 | 55 | 74 | 130 | 154 | 125.5 | 97.1 | 519 | 455.3 | 455.3 | 455.4 | 465.2
A (clean images) | 1737 | ptbxl1737_clean | classical | 0 | 170 | 1949 | 304 | 4 | 365 | 11 | 0 | 7 | 57 | 71 | 127 | 147.1 | 122.6 | 94.2 | 100.2 | 95.1 | 95.1 | 95.1 | 103.2
A (clean images) | 1737 | ptbxl1737_clean | learned | 0 | 190 | 1970 | 301 | 4 | 19898 | 10 | 0 | 8 | 57 | 70.4 | 126.4 | 146.9 | 121.5 | 93.4 | 515.4 | 452.8 | 452.3 | 452.5 | 461.4
A (clean images) | 3908 | ptbxl3908_clean | classical | 0 | 182 | 1981 | 310 | 4 | 378 | 13 | 0 | 6 | 62 | 69.8 | 125.8 | 146.2 | 121.1 | 93.1 | 98.2 | 94.2 | 94.2 | 94.3 | 103
A (clean images) | 3908 | ptbxl3908_clean | learned | 0 | 185 | 1993 | 308 | 4 | 19990 | 10 | 0 | 7 | 53 | 71.5 | 127.5 | 151.5 | 122.7 | 94.7 | 516.8 | 455.3 | 455.3 | 455.3 | 463.4
A (clean images) | 4016 | ptbxl4016_clean | classical | 0 | 179 | 1947 | 300 | 4 | 384 | 11 | 0 | 9 | 54 | 73.6 | 129.7 | 153.9 | 125.1 | 96.7 | 103.9 | 97.5 | 97.5 | 97.5 | 106.6
A (clean images) | 4016 | ptbxl4016_clean | learned | 0 | 176 | 1963 | 302 | 4 | 19929 | 10 | 0 | 7 | 55 | 71.2 | 127.3 | 151.1 | 125.8 | 97.6 | 519.5 | 451 | 450.9 | 450.9 | 461.4
A (clean images) | 5830 | ptbxl5830_clean | classical | 0 | 170 | 2020 | 354 | 4 | 389 | 9 | 0 | 6 | 53 | 74 | 130.1 | 150 | 125.1 | 97.2 | 103.3 | 98.2 | 98.2 | 98.2 | 106.5
A (clean images) | 5830 | ptbxl5830_clean | learned | 0 | 160 | 2048 | 356 | 4 | 20169 | 11 | 0 | 7 | 54 | 74.7 | 130.7 | 151.3 | 125.9 | 97.8 | 519.3 | 457.6 | 457.6 | 457.6 | 466.1
A (clean images) | 7916 | ptbxl7916_clean | classical | 0 | 174 | 1974 | 270 | 4 | 383 | 9 | 0 | 7 | 51 | 74.9 | 130.9 | 155.4 | 126.4 | 98.3 | 103.5 | 98.7 | 98.8 | 98.8 | 107
A (clean images) | 7916 | ptbxl7916_clean | learned | 0 | 186 | 1931 | 276 | 4 | 19984 | 10 | 0 | 7 | 56 | 75.7 | 131.7 | 152 | 126.9 | 98.8 | 520.8 | 458.2 | 458.2 | 458.4 | 468
A (clean images) | 8504 | ptbxl8504_clean | classical | 0 | 180 | 1992 | 316 | 4 | 376 | 9 | 0 | 6 | 54 | 73.7 | 129.7 | 149.8 | 125.2 | 97 | 102 | 97.8 | 97.7 | 97.8 | 106.9
A (clean images) | 8504 | ptbxl8504_clean | learned | 0 | 200 | 1973 | 322 | 4 | 19781 | 10 | 0 | 7 | 53 | 73 | 129.1 | 153.1 | 124 | 95.9 | 518.1 | 455.9 | 455.6 | 455.7 | 464.1
B (ideal photos) | 9 | pair_20261013_113100 | classical | 0 | 165 | 1495 | 267 | 4 | 330 | 9 | 2 | 24 | 68 | 72.1 | 118 | 131.5 | 112.3 | 83.8 | 88.2 | 84.6 | 84.6 | 84.6 | 93.2
B (ideal photos) | 9 | pair_20261013_113100 | learned | 0 | 137 | 2055 | 261 | 46 | 19678 | 37 | 8 | 29 | 80 | 62.4 | 109 | 122.1 | 108.1 | 80.2 | 502.1 | 439.9 | 439.6 | 439.8 | 448.7
B (ideal photos) | 9 | pair_20261013_113548 | classical | 0 | 142 | 1312 | 159 | 3 | 241 | 6 | 1 | 6 | 62 | 72.1 | 116.2 | 124.5 | 111.3 | 82.9 | 85.7 | 83.7 | 83.7 | 83.7 | 92.5
B (ideal photos) | 9 | pair_20261013_113548 | learned | 0 | 146 | 1800 | 163 | 26 | 19675 | 44 | 6 | 32 | 74 | 67 | 111.1 | 118.3 | 105.9 | 77.8 | 499.8 | 436.8 | 438.2 | 438.2 | 445.3
B (ideal photos) | 38 | pair_20261013_115454 | classical | 0 | 137 | 1532 | 195 | 5 | 232 | 10 | 0 | 8 | 71 | 72.7 | 118.7 | 132 | 113.7 | 85.3 | 88 | 86.1 | 86.1 | 86.1 | 94.9
B (ideal photos) | 38 | pair_20261013_115454 | learned | 0 | 107 | 1389 | 202 | 4 | 19522 | 9 | 1 | 7 | 58 | 68.6 | 115 | 125 | 110.8 | 82.4 | 504.1 | 441.3 | 441.3 | 441.4 | 449.8
B (ideal photos) | 38 | pair_20261013_115548 | classical | 0 | 126 | 1463 | 192 | 3 | 218 | 5 | 0 | 5 | 56 | 70.7 | 114.8 | 125.5 | 111.3 | 83.1 | 87.5 | 83.8 | 83.8 | 83.8 | 92.2
B (ideal photos) | 38 | pair_20261013_115548 | learned | 0 | 114 | 1371 | 185 | 3 | 19520 | 9 | 0 | 7 | 54 | 66.1 | 110.1 | 121.8 | 107.6 | 79.3 | 501 | 438.3 | 438.3 | 438.3 | 446.8
B (ideal photos) | 40 | pair_20261013_132946 | classical | 0 | 187 | 1365 | 186 | 3 | 231 | 8 | 0 | 6 | 55 | 73.6 | 119.8 | 133.1 | 115.1 | 86.8 | 91.2 | 87.5 | 87.5 | 87.5 | 96.3
B (ideal photos) | 40 | pair_20261013_132946 | learned | 0 | 128 | 1395 | 170 | 3 | 19378 | 9 | 0 | 16 | 56 | 67.9 | 113.9 | 127.9 | 109.4 | 81 | 502.8 | 439.7 | 439.3 | 439.5 | 449.1
B (ideal photos) | 40 | pair_20261013_133135 | classical | 0 | 148 | 1362 | 211 | 4 | 237 | 8 | 0 | 7 | 52 | 72.8 | 118.8 | 128 | 113.6 | 85.2 | 89.4 | 86.4 | 85.8 | 85.8 | 94.3
B (ideal photos) | 40 | pair_20261013_133135 | learned | 0 | 121 | 1408 | 230 | 4 | 19619 | 8 | 0 | 8 | 58 | 68.6 | 114.7 | 124.5 | 110.3 | 81.9 | 503.6 | 440.7 | 440.7 | 440.6 | 450.5
B (ideal photos) | 57 | pair_20261013_133702 | classical | 0 | 149 | 1368 | 222 | 4 | 238 | 10 | 0 | 7 | 57 | 71.1 | 117.2 | 126.2 | 112 | 83.7 | 87.4 | 85.1 | 85.1 | 85.1 | 93.8
B (ideal photos) | 57 | pair_20261013_133702 | learned | 0 | 133 | 1395 | 216 | 3 | 19927 | 9 | 0 | 7 | 55 | 67.3 | 113.4 | 127.1 | 109.2 | 80.8 | 502.5 | 437.6 | 437.4 | 437.5 | 446.6
B (ideal photos) | 57 | pair_20261013_133915 | classical | 0 | 141 | 1336 | 322 | 3 | 248 | 6 | 0 | 6 | 61 | 72.9 | 116.9 | 128.6 | 114.6 | 85.8 | 89.4 | 87.6 | 87.6 | 87.6 | 96.7
B (ideal photos) | 57 | pair_20261013_133915 | learned | 0 | 113 | 1349 | 297 | 3 | 19525 | 8 | 0 | 7 | 60 | 68.1 | 112.1 | 124.2 | 110 | 81.2 | 502.8 | 440 | 440 | 440.2 | 450
B (ideal photos) | 160 | pair_20261013_145325 | classical | 0 | 182 | 1665 | 198 | 3 | 235 | 8 | 0 | 6 | 51 | 77.5 | 123.6 | 133 | 118.9 | 90.6 | 95 | 91.2 | 91.2 | 91.2 | 100
B (ideal photos) | 160 | pair_20261013_145325 | learned | 0 | 138 | 1563 | 193 | 3 | 19760 | 9 | 0 | 7 | 53 | 72.2 | 118.3 | 128.4 | 114.2 | 85.7 | 507.3 | 444.5 | 444.1 | 444.1 | 452.8
B (ideal photos) | 160 | pair_20261013_145443 | classical | 0 | 131 | 1532 | 182 | 4 | 222 | 9 | 0 | 7 | 60 | 76.2 | 122.3 | 131.6 | 117.4 | 89.1 | 92.9 | 89.5 | 89.5 | 89.5 | 97.7
B (ideal photos) | 160 | pair_20261013_145443 | learned | 0 | 109 | 1500 | 207 | 3 | 19492 | 8 | 0 | 7 | 54 | 73.4 | 119.4 | 133.5 | 114.9 | 86.5 | 508.2 | 444.4 | 444.5 | 444.5 | 453.3
B (ideal photos) | 167 | pair_20261013_144827 | classical | 0 | 143 | 1506 | 203 | 2 | 216 | 8 | 0 | 6 | 56 | 74.7 | 118.8 | 129.7 | 116 | 87.5 | 90.4 | 88.8 | 88.8 | 89 | 97.6
B (ideal photos) | 167 | pair_20261013_144827 | learned | 0 | 97 | 1522 | 189 | 2 | 19272 | 8 | 0 | 7 | 61 | 71.3 | 115.2 | 125.2 | 113.1 | 84.7 | 506.3 | 442.7 | 442.7 | 442.9 | 451.2
B (ideal photos) | 167 | pair_20261013_145129 | classical | 0 | 142 | 1424 | 225 | 2 | 227 | 8 | 0 | 6 | 56 | 77.4 | 121.4 | 132.8 | 116.5 | 88.2 | 91.2 | 89.7 | 89.6 | 89.6 | 98.2
B (ideal photos) | 167 | pair_20261013_145129 | learned | 0 | 126 | 1476 | 213 | 3 | 19429 | 8 | 0 | 7 | 52 | 71.6 | 115.6 | 127.7 | 111.4 | 82.9 | 504.7 | 441.6 | 441.6 | 441.6 | 449.6
B (ideal photos) | 1737 | pair_20261013_134146 | classical | 0 | 178 | 1999 | 237 | 3 | 312 | 8 | 0 | 6 | 55 | 72.3 | 124.4 | 139 | 118.7 | 90.5 | 95.2 | 90.7 | 90.7 | 90.7 | 99
B (ideal photos) | 1737 | pair_20261013_134146 | learned | 0 | 141 | 1958 | 254 | 3 | 20466 | 9 | 0 | 7 | 54 | 69.1 | 121.1 | 140.8 | 115.1 | 86.9 | 508.7 | 445.5 | 445.5 | 445.5 | 454.9
B (ideal photos) | 1737 | pair_20261013_134236 | classical | 0 | 198 | 2245 | 307 | 4 | 358 | 10 | 0 | 7 | 62 | 73.7 | 125.7 | 144.6 | 124.3 | 96 | 100.5 | 96.2 | 96.2 | 96.2 | 104.5
B (ideal photos) | 1737 | pair_20261013_134236 | learned | 0 | 181 | 2272 | 303 | 4 | 20226 | 10 | 0 | 8 | 53 | 69.9 | 122 | 142.4 | 121.5 | 93.2 | 515 | 452 | 452 | 452 | 462.5
B (ideal photos) | 3908 | pair_20261013_134604 | classical | 0 | 212 | 2634 | 319 | 4 | 394 | 12 | 0 | 6 | 61 | 74.4 | 130.4 | 149.4 | 124.8 | 96.6 | 102.8 | 97.5 | 97.5 | 97.5 | 105.2
B (ideal photos) | 3908 | pair_20261013_134604 | learned | 0 | 182 | 2567 | 322 | 4 | 19746 | 10 | 0 | 7 | 54 | 65.9 | 121.9 | 145.9 | 120.4 | 92.4 | 514.4 | 451.6 | 451.6 | 451.8 | 460.8
B (ideal photos) | 3908 | pair_20261013_134722 | classical | 0 | 203 | 2432 | 293 | 5 | 358 | 11 | 0 | 6 | 52 | 72.9 | 129 | 151 | 122.9 | 94.7 | 100.8 | 95.2 | 95.2 | 95.3 | 103.9
B (ideal photos) | 3908 | pair_20261013_134722 | learned | 0 | 181 | 2365 | 292 | 5 | 20191 | 10 | 0 | 13 | 82 | 70.3 | 126.3 | 150.7 | 121.7 | 93.6 | 515.6 | 451.9 | 450.7 | 451 | 459.5
B (ideal photos) | 4016 | pair_20261013_143426 | classical | 0 | 157 | 1659 | 235 | 3 | 248 | 8 | 0 | 10 | 52 | 74 | 120 | 129.1 | 114.6 | 86.1 | 94.4 | 87.2 | 87.2 | 87.2 | 95.8
B (ideal photos) | 4016 | pair_20261013_143426 | learned | 0 | 122 | 1683 | 249 | 3 | 19648 | 9 | 0 | 7 | 70 | 69.5 | 115.6 | 129.8 | 111.4 | 83.1 | 504.9 | 442 | 442 | 442.1 | 451.8
B (ideal photos) | 4016 | pair_20261013_144140 | classical | 0 | 140 | 1583 | 201 | 2 | 226 | 8 | 0 | 7 | 56 | 77.8 | 123.9 | 137.4 | 118.7 | 90.4 | 94.8 | 92.1 | 92.1 | 92.2 | 101.1
B (ideal photos) | 4016 | pair_20261013_144140 | learned | 0 | 114 | 1538 | 222 | 3 | 19632 | 9 | 0 | 7 | 53 | 72.8 | 118.7 | 129.1 | 114.9 | 86.6 | 508.4 | 445.5 | 445.5 | 445.6 | 454.2
B (ideal photos) | 5830 | pair_20261013_162729 | classical | 0 | 165 | 1583 | 308 | 3 | 278 | 5 | 0 | 12 | 56 | 77.8 | 123.8 | 137.2 | 119.2 | 90.8 | 94.5 | 90.8 | 90.8 | 90.8 | 99
B (ideal photos) | 5830 | pair_20261013_162729 | learned | 0 | 111 | 1585 | 274 | 2 | 19433 | 9 | 0 | 7 | 56 | 74.4 | 120.4 | 134.7 | 116.2 | 87.8 | 509.6 | 446.1 | 446.1 | 446.1 | 454.9
B (ideal photos) | 5830 | pair_20261013_162839 | classical | 0 | 140 | 1547 | 220 | 2 | 214 | 2 | 0 | 2 | 60 | 78.2 | 124.3 | 136.7 | 116.1 | 87.8 | 90.4 | 88 | 88 | 88 | 96.1
B (ideal photos) | 5830 | pair_20261013_162839 | learned | 0 | 122 | 1579 | 222 | 2 | 19387 | 8 | 0 | 7 | 56 | 73.7 | 119.8 | 128 | 113.4 | 85.1 | 506.7 | 444 | 444 | 444 | 452.5
B (ideal photos) | 7916 | pair_20261013_162442 | classical | 0 | 138 | 1462 | 173 | 4 | 230 | 8 | 0 | 7 | 53 | 76.2 | 120.3 | 131.5 | 117.5 | 89 | 92.5 | 90.1 | 90.1 | 90.1 | 98.4
B (ideal photos) | 7916 | pair_20261013_162442 | learned | 0 | 103 | 1471 | 178 | 3 | 19391 | 10 | 0 | 13 | 83 | 72.8 | 116.8 | 129.2 | 114.5 | 86.1 | 507.8 | 444.6 | 444.6 | 444.8 | 453.1
B (ideal photos) | 7916 | pair_20261013_162603 | classical | 0 | 147 | 1547 | 183 | 3 | 242 | 7 | 0 | 8 | 60 | 76.1 | 122.1 | 135.2 | 117.3 | 88.9 | 91.8 | 89.3 | 89.3 | 89.3 | 97.6
B (ideal photos) | 7916 | pair_20261013_162603 | learned | 0 | 111 | 1491 | 190 | 3 | 19695 | 9 | 0 | 7 | 52 | 71.8 | 117.8 | 131.9 | 113.8 | 85.4 | 507.1 | 444.1 | 443.9 | 443.9 | 452.5
B (ideal photos) | 8504 | pair_20261013_162024 | classical | 0 | 137 | 1348 | 166 | 2 | 203 | 7 | 0 | 6 | 56 | 75.4 | 119.4 | 130.8 | 114.4 | 86 | 89.3 | 87.9 | 87.9 | 88 | 96.4
B (ideal photos) | 8504 | pair_20261013_162024 | learned | 0 | 107 | 1377 | 174 | 2 | 19423 | 8 | 0 | 13 | 103 | 70.5 | 114.5 | 122.6 | 110.1 | 81.8 | 503.7 | 440.8 | 440.8 | 442.3 | 450.4
B (ideal photos) | 8504 | pair_20261013_162240 | classical | 0 | 157 | 1606 | 242 | 3 | 239 | 8 | 0 | 6 | 51 | 79.5 | 125.6 | 135 | 119.9 | 91.6 | 95.6 | 93.1 | 93.1 | 93.1 | 102.4
B (ideal photos) | 8504 | pair_20261013_162240 | learned | 0 | 116 | 1636 | 250 | 3 | 19648 | 8 | 0 | 7 | 58 | 72.7 | 118.7 | 129 | 114.7 | 86.3 | 508.1 | 445.2 | 444.9 | 444.9 | 455.6

## Photo quality (B)
frame | orientation_shot | sheet_fill | max_corner_angle_dev_deg | keystone_top_bottom | light_ratio_p5_p95 | paper_luma_median | sharpness_lap_var | clipped_fraction | photo | ecg_id | px_per_mm | learned_pcc_median | classical_pcc_median
2396x3200 | portrait | 0.432 | 6.02 | 1.065 | 0.787 | 176 | 864 | 0 | data/kit_photos/setB/eacg_20261013_112807.jpg | 9 | 7.99 | 0.92875 | 0.5803
2396x3200 | portrait | 0.401 | 1.62 | 1.038 | 0.882 | 174 | 716.7 | 0 | data/kit_photos/setB/eacg_20261013_113507.jpg | 9 | 7.2192 | 0.92685 | 0.5473
2396x3200 | portrait | 0.403 | 9.19 | 1.153 | 0.873 | 172 | 974.3 | 0 | data/kit_photos/setB/eacg_20261013_115407.jpg | 38 | 7.656 | 0.9283 | 0.74625
2396x3200 | portrait | 0.401 | 8.84 | 1.204 | 0.906 | 172 | 274.2 | 0 | data/kit_photos/setB/eacg_20261013_115513.jpg | 38 | 7.6837 | 0.938 | 0.645
2396x3200 | portrait | 0.41 | 4.19 | 1.087 | 0.895 | 164 | 1072.5 | 0 | data/kit_photos/setB/eacg_20261013_132719.jpg | 40 | 7.8358 | 0.96725 | 0.8299
2396x3200 | portrait | 0.438 | 4.57 | 1.086 | 0.892 | 160 | 751.3 | 0 | data/kit_photos/setB/eacg_20261013_133023.jpg | 40 | 8.032 | 0.95345 | 0.92675
2396x3200 | portrait | 0.418 | 5.39 | 1.077 | 0.895 | 164 | 874.8 | 0 | data/kit_photos/setB/eacg_20261013_133534.jpg | 57 | 7.9639 | 0.89765 | 0.7633
2396x3200 | portrait | 0.377 | 9.31 | 1.102 | 0.903 | 167 | 597.6 | 0 | data/kit_photos/setB/eacg_20261013_133810.jpg | 57 | 7.652 | 0.90815 | 0.74825
2396x3200 | portrait | 0.442 | 2 | 1.034 | 0.85 | 157 | 709.8 | 0 | data/kit_photos/setB/eacg_20261013_145217.jpg | 160 | 8.016 | 0.95115 | 0.46585
2396x3200 | portrait | 0.413 | 2.53 | 1.056 | 0.839 | 161 | 1037.5 | 0 | data/kit_photos/setB/eacg_20261013_145357.jpg | 160 | 7.7838 | 0.94605 | 0.74415
2396x3200 | portrait | 0.399 | 8.85 | 1.146 | 0.871 | 163 | 751.1 | 0 | data/kit_photos/setB/eacg_20261013_144637.jpg | 167 | 7.6877 | 0.9639 | 0.42875
2396x3200 | portrait | 0.407 | 2.84 | 1.065 | 0.853 | 163 | 1344.4 | 0 | data/kit_photos/setB/eacg_20261013_145006.jpg | 167 | 7.7277 | 0.966 | 0.4559
3200x2396 | landscape | 0.559 | 2.08 | 1.035 | 0.888 | 152 | 632.6 | 0 | data/kit_photos/setB/eacg_20261013_134006.jpg | 1737 | 9.057 | 0.95375 | 0.41935
3200x2396 | landscape | 0.671 | 3.21 | 1.066 | 0.881 | 134 | 353.6 | 0 | data/kit_photos/setB/eacg_20261013_134159.jpg | 1737 | 9.986 | 0.94655 | 0.6556
2396x3200 | portrait | 0.763 | 2.49 | 1.004 | 0.867 | 133 | 306.6 | 0 | data/kit_photos/setB/eacg_20261013_134510.jpg | 3908 | 10.6266 | 0.97525 | 0.95725
3200x2396 | landscape | 0.704 | 4.73 | 1.101 | 0.866 | 132 | 144.2 | 0 | data/kit_photos/setB/eacg_20261013_134642.jpg | 3908 | 10.1782 | 0.9545 | 0.86435
2396x3200 | portrait | 0.456 | 2.31 | 1.05 | 0.823 | 150 | 1330.8 | 0 | data/kit_photos/setB/eacg_20261013_143255.jpg | 4016 | 8.1881 | 0.98545 | 0.81745
2396x3200 | portrait | 0.419 | 3.04 | 1.057 | 0.891 | 164 | 1100.5 | 0 | data/kit_photos/setB/eacg_20261013_144103.jpg | 4016 | 7.8979 | 0.98405 | 0.73405
2396x3200 | portrait | 0.412 | 8.28 | 1.163 | 0.904 | 160 | 223.6 | 0 | data/kit_photos/setB/eacg_20261013_162629.jpg | 5830 | 7.9239 | 0.91295 | 0.15125
2396x3200 | portrait | 0.419 | 8.51 | 1.16 | 0.892 | 174 | 565.5 | 0 | data/kit_photos/setB/eacg_20261013_162743.jpg | 5830 | 6.987 | 0.93575 | 0.3575
2396x3200 | portrait | 0.407 | 5.23 | 1.111 | 0.908 | 167 | 647.3 | 0 | data/kit_photos/setB/eacg_20261013_162400.jpg | 7916 | 7.8038 | 0.95545 | 0.3852
2396x3200 | portrait | 0.433 | 4.3 | 1.098 | 0.893 | 161 | 204 | 0 | data/kit_photos/setB/eacg_20261013_162511.jpg | 7916 | 8.044 | 0.9342 | 0.5288
2396x3200 | portrait | 0.368 | 4.73 | 1.098 | 0.879 | 165 | 1025 | 0 | data/kit_photos/setB/eacg_20261013_161907.jpg | 8504 | 7.3994 | 0.9423 | 0.75245
2396x3200 | portrait | 0.434 | 4.52 | 1.094 | 0.879 | 166 | 794.2 | 0 | data/kit_photos/setB/eacg_20261013_162135.jpg | 8504 | 8.012 | 0.95445 | 0.78385

## Quality vs PCC (B)
photo measure | tracer | spearman_rho | p | n photos
max_corner_angle_dev_deg | learned | -0.451 | 0.027 | 24
max_corner_angle_dev_deg | classical | -0.107 | 0.6174 | 24
keystone_top_bottom | learned | -0.348 | 0.0952 | 24
keystone_top_bottom | classical | -0.233 | 0.2738 | 24
sheet_fill | learned | 0.362 | 0.0822 | 24
sheet_fill | classical | 0.271 | 0.2009 | 24
px_per_mm | learned | 0.387 | 0.0618 | 24
px_per_mm | classical | 0.336 | 0.1088 | 24
light_ratio_p5_p95 | learned | -0.321 | 0.1261 | 24
light_ratio_p5_p95 | classical | -0.204 | 0.3378 | 24
sharpness_lap_var | learned | 0.314 | 0.1352 | 24
sharpness_lap_var | classical | 0.228 | 0.2843 | 24

## Per lead (all)
set_name | ecg_id | ptbxl_label | photo_id | tracer_kind | export | lead_label | row | col | coverage | pcc_asis | pcc_aligned | snr_db_asis | shifted_snr_db_asis | rmse_mv_asis | rmse_mv_median_removed | gain_ratio | r_ratio | r_ratio_aligned | align_scale_pct | align_lag_ms | align_overlap_s | shift_ms | shift_status | quality | warnings
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | I | 0 | 0 | 0.98 | 0.9868 | 0.9899 | 15.829 | 16.925 | 0.0464 | 0.0133 | 0.9487 | 0.9414 | 0.9409 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | aVR | 0 | 1 | 0.9608 | 0.3354 | 0.3759 | -6.258 | -6.172 | 0.2076 | 0.1975 | 0.7148 | 0.898 | 0.8972 | 1 | 20 | 2.476 | -6 | UNRELIABLE | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | V1 | 0 | 2 | 0.9624 | 0.9668 | 0.9941 | 11.827 | 18.275 | 0.0409 | 0.0405 | 0.9064 | 0.9011 | 0.9007 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | V4 | 0 | 3 | 0.9808 | 0.968 | 0.9915 | 12.08 | 16.959 | 0.0581 | 0.057 | 0.9004 | 0.8486 | 0.8487 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | II | 1 | 0 | 0.98 | 0.9941 | 0.9961 | 19.142 | 20.796 | 0.0439 | 0.0117 | 0.9681 | 0.9248 | 0.9248 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | aVL | 1 | 1 | 0.9608 | 0.9623 | 0.9753 | 11.268 | 12.985 | 0.021 | 0.0155 | 0.9044 | 0.8016 | 0.8016 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | V2 | 1 | 2 | 0.9624 | 0.9521 | 0.9823 | 9.889 | 12.93 | 0.1124 | 0.103 | 0.8153 | 0.7423 | 0.7422 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | V5 | 1 | 3 | 0.9808 | 0.9772 | 0.9947 | 13.553 | 19.199 | 0.0361 | 0.0359 | 0.9316 | 0.9062 | 0.9062 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | III | 2 | 0 | 0.98 | 0.9114 | 0.954 | 7.215 | 7.374 | 0.0283 | 0.0285 | 0.961 | 0.8922 | 0.889 | -0.5 | -8 | 2.5 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | aVF | 2 | 1 | 0.9608 | 0.985 | 0.993 | 15.285 | 18.211 | 0.0223 | 0.015 | 0.9538 | 0.9174 | 0.9183 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | V3 | 2 | 2 | 0.9624 | 0.9533 | 0.989 | 10.452 | 15.706 | 0.1213 | 0.0772 | 0.8693 | 0.833 | 0.8322 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | V6 | 2 | 3 | 0.9808 | 0.9818 | 0.9958 | 14.547 | 20.257 | 0.0282 | 0.023 | 0.942 | 0.9223 | 0.9223 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261012_134254.csv | II rhythm | 3 | 0 | 0.9992 | 0.9889 | 0.9956 | 16.868 | 20.692 | 0.0408 | 0.0171 | 0.9621 | 0.9223 | 0.9224 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | I | 0 | 0 | 0.9992 | 0.9959 | 0.9959 | 20.956 | 20.956 | 0.0464 | 0.0073 | 0.985 | 0.9902 | 0.9902 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | aVR | 0 | 1 | 0.9976 | 0.9967 | 0.9967 | 21.877 | 21.877 | 0.0257 | 0.0077 | 0.9918 | 0.963 | 0.963 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | V1 | 0 | 2 | 0.9984 | 0.991 | 0.9989 | 17.493 | 26.457 | 0.0233 | 0.0211 | 0.9815 | 0.9726 | 0.9726 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | V4 | 0 | 3 | 0.9992 | 0.9942 | 0.9981 | 19.448 | 24 | 0.0281 | 0.0244 | 0.9753 | 0.9634 | 0.9635 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | II | 1 | 0 | 0.9976 | 0.9988 | 0.9988 | 26.339 | 26.339 | 0.0407 | 0.0051 | 0.9927 | 0.9824 | 0.9824 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | aVL | 1 | 1 | 0.9976 | 0.9888 | 0.9899 | 16.454 | 16.454 | 0.0167 | 0.0085 | 0.9648 | 0.9358 | 0.9034 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | V2 | 1 | 2 | 0.9984 | 0.9894 | 0.9974 | 16.686 | 22.185 | 0.0626 | 0.0471 | 0.9593 | 0.9475 | 0.9475 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | V5 | 1 | 3 | 0.9992 | 0.9959 | 0.9986 | 20.98 | 25.639 | 0.0171 | 0.0153 | 0.9848 | 0.9742 | 0.9742 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | III | 2 | 0 | 0.9976 | 0.9958 | 0.9958 | 20.355 | 20.355 | 0.007 | 0.0063 | 0.988 | 0.9714 | 0.9714 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | aVF | 2 | 1 | 0.9976 | 0.9962 | 0.9962 | 21.303 | 21.303 | 0.0172 | 0.0075 | 0.9858 | 0.9786 | 0.9786 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | V3 | 2 | 2 | 0.9984 | 0.9878 | 0.9961 | 16.046 | 20.219 | 0.1071 | 0.0405 | 0.94 | 0.9476 | 0.9476 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | V6 | 2 | 3 | 0.9992 | 0.9967 | 0.9989 | 21.899 | 26.497 | 0.0184 | 0.0099 | 0.9853 | 0.9888 | 0.989 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 9 | NORM | ptbxl9_clean | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261012_134218.csv | II rhythm | 3 | 0 | 0.9992 | 0.9962 | 0.9962 | 21.452 | 21.452 | 0.0426 | 0.0101 | 0.9947 | 0.9936 | 0.9936 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | I | 0 | 0 | 0.9792 | 0.8371 | 0.8586 | 5.078 | 5.608 | 0.04 | 0.0292 | 0.8165 | 0.8705 | 0.8641 | 0.25 | 6 | 2.494 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | aVR | 0 | 1 | 0.9632 | 0.9145 | 0.9488 | 7.89 | 10.044 | 0.0399 | 0.0343 | 0.8529 | 0.8042 | 0.8038 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | V1 | 0 | 2 | 0.9624 | 0.9189 | 0.934 | 7.557 | 8.48 | 0.0987 | 0.096 | 0.9663 | 0.9774 | 0.9777 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | V4 | 0 | 3 | 0.9792 | 0.9795 | 0.9956 | 14.012 | 20.369 | 0.041 | 0.0407 | 0.9594 | 0.8977 | 0.8972 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | II | 1 | 0 | 0.98 | 0.9017 | 0.9349 | 7.336 | 8.25 | 0.0649 | 0.0536 | 0.8243 | 0.7764 | 0.7735 | 0.5 | 8 | 2.488 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | aVL | 1 | 1 | 0.9616 | 0.9523 | 0.9539 | 10.236 | 10.381 | 0.0163 | 0.0165 | 0.8944 | 0.8157 | 0.8157 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | V2 | 1 | 2 | 0.9632 | 0.9269 | 0.938 | 8.018 | 8.775 | 0.1173 | 0.1167 | 0.9748 | 0.9045 | 0.9047 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | V5 | 1 | 3 | 0.9792 | 0.9828 | 0.9964 | 14.865 | 21.409 | 0.0334 | 0.0294 | 0.9685 | 0.8935 | 0.8945 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | III | 2 | 0 | 0.9792 | 0.8872 | 0.9307 | 6.464 | 7.852 | 0.0667 | 0.0425 | 0.8882 | 0.9049 | 0.9054 | 0.5 | 8 | 2.488 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | aVF | 2 | 1 | 0.9632 | 0.9224 | 0.9463 | 8.267 | 9.819 | 0.0479 | 0.0463 | 0.8616 | 0.7677 | 0.7678 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | V3 | 2 | 2 | 0.9624 | 0.8845 | 0.9286 | 6.387 | 8.491 | 0.1062 | 0.1065 | 0.8792 | 0.9002 | 0.9 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | V6 | 2 | 3 | 0.9792 | 0.9824 | 0.996 | 14.751 | 20.925 | 0.0218 | 0.0202 | 0.9728 | 0.7978 | 0.7978 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261012_144052.csv | II rhythm | 3 | 0 | 0.9992 | 0.9657 | 0.9851 | 11.772 | 15.256 | 0.037 | 0.0346 | 0.9245 | 0.763 | 0.7631 | 0 | 4 | 9.996 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | I | 0 | 0 | 0.9936 | 0.7924 | 0.8931 | 4.058 | 4.59 | 0.0446 | 0.0328 | 0.7608 | 0.9353 | 0.9329 | 0.5 | 8 | 2.488 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | aVR | 0 | 1 | 1 | 0.9721 | 0.9877 | 12.591 | 14.876 | 0.0309 | 0.0199 | 0.9299 | 0.9455 | 0.9455 | -0.25 | -2 | 2.5 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | V1 | 0 | 2 | 0.9992 | 0.9907 | 0.9987 | 17.268 | 25.553 | 0.0361 | 0.0314 | 0.9909 | 0.9927 | 0.9927 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | V4 | 0 | 3 | 0.9968 | 0.9933 | 0.9994 | 18.864 | 29.043 | 0.0233 | 0.0233 | 0.9901 | 0.9593 | 0.9586 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | II | 1 | 0 | 0.9992 | 0.9529 | 0.9681 | 10.417 | 10.417 | 0.0517 | 0.0376 | 0.9031 | 0.9104 | 0.9069 | 0.25 | 2 | 2.494 | 0 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | aVL | 1 | 1 | 0.9984 | 0.9848 | 0.987 | 15.085 | 15.085 | 0.0096 | 0.0094 | 0.9435 | 0.8103 | 0.8098 | -0.25 | -4 | 2.5 | 0 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | V2 | 1 | 2 | 1 | 0.994 | 0.9994 | 19.245 | 28.809 | 0.04 | 0.032 | 0.9914 | 0.9824 | 0.9824 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | V5 | 1 | 3 | 0.9968 | 0.9939 | 0.9995 | 19.299 | 29.247 | 0.0208 | 0.0177 | 0.9904 | 0.9616 | 0.9616 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | III | 2 | 0 | 0.9952 | 0.8312 | 0.9086 | 4.92 | 5.557 | 0.0737 | 0.0508 | 0.7872 | 0.9602 | 0.9474 | 0.75 | 10 | 2.482 | -4 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | aVF | 2 | 1 | 1 | 0.9448 | 0.9641 | 9.698 | 10.602 | 0.0431 | 0.0393 | 0.8879 | 0.9392 | 0.9351 | -0.5 | -6 | 2.5 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | V3 | 2 | 2 | 0.9992 | 0.9752 | 0.9968 | 13.12 | 21.608 | 0.0504 | 0.0491 | 0.9505 | 0.9792 | 0.9787 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | V6 | 2 | 3 | 0.9968 | 0.9929 | 0.9967 | 18.621 | 21.935 | 0.0141 | 0.0129 | 0.9962 | 0.9532 | 0.9533 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 38 | NORM | ptbxl38_clean | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261012_144014.csv | II rhythm | 3 | 0 | 0.9992 | 0.9864 | 0.9921 | 15.743 | 18.125 | 0.0249 | 0.0219 | 0.9768 | 0.9221 | 0.9223 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | I | 0 | 0 | 0.98 | 0.9784 | 0.982 | 13.873 | 14.656 | 0.0555 | 0.0332 | 0.9513 | 0.8928 | 0.8921 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | aVR | 0 | 1 | 0.9608 | 0.9842 | 0.9927 | 15.178 | 18.179 | 0.0299 | 0.0232 | 0.9375 | 0.8958 | 0.8958 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | V1 | 0 | 2 | 0.9624 | 0.9827 | 0.9963 | 14.585 | 20.607 | 0.065 | 0.0524 | 0.9455 | 0.9211 | 0.9211 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | V4 | 0 | 3 | 0.9808 | 0.9782 | 0.9917 | 13.956 | 17.857 | 0.0461 | 0.0359 | 0.9328 | 0.8406 | 0.8406 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | II | 1 | 0 | 0.98 | 0.9597 | 0.9744 | 11.065 | 11.343 | 0.031 | 0.03 | 0.9631 | 0.7815 | 0.7756 | -0.5 | -6 | 2.5 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | aVL | 1 | 1 | 0.9608 | 0.9841 | 0.993 | 15.096 | 18.428 | 0.0276 | 0.0274 | 0.9503 | 0.8333 | 0.8333 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | V2 | 1 | 2 | 0.9624 | 0.9859 | 0.9939 | 15.686 | 18.906 | 0.1333 | 0.065 | 0.9475 | 0.8793 | 0.8787 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | V5 | 1 | 3 | 0.9808 | 0.9823 | 0.9952 | 14.863 | 20.143 | 0.0674 | 0.0293 | 0.948 | 0.9023 | 0.9023 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | III | 2 | 0 | 0.98 | 0.9507 | 0.9605 | 10.201 | 10.541 | 0.0419 | 0.0358 | 0.9472 | 0.8194 | 0.8173 | -0.5 | -6 | 2.5 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | aVF | 2 | 1 | 0.9608 | 0.971 | 0.9776 | 12.473 | 13.506 | 0.023 | 0.0195 | 0.9175 | 0.9049 | 0.9041 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | V3 | 2 | 2 | 0.9624 | 0.9675 | 0.988 | 12.347 | 16.213 | 0.0931 | 0.0581 | 0.9066 | 0.8254 | 0.8254 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | V6 | 2 | 3 | 0.9808 | 0.9844 | 0.996 | 15.295 | 20.71 | 0.0778 | 0.0233 | 0.948 | 0.9457 | 0.9457 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261012_145119.csv | II rhythm | 3 | 0 | 0.9992 | 0.9715 | 0.9811 | 12.699 | 14.456 | 0.0335 | 0.0244 | 0.9536 | 0.7563 | 0.7563 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | I | 0 | 0 | 0.9976 | 0.9992 | 0.9992 | 27.509 | 27.509 | 0.0416 | 0.0069 | 0.9825 | 0.973 | 0.973 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | aVR | 0 | 1 | 0.9976 | 0.9963 | 0.9969 | 21.556 | 21.556 | 0.0215 | 0.0112 | 0.9834 | 0.9742 | 0.969 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | V1 | 0 | 2 | 0.9984 | 0.9954 | 0.9994 | 20.334 | 28.678 | 0.0469 | 0.027 | 0.9897 | 0.977 | 0.9769 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | V4 | 0 | 3 | 0.9992 | 0.9959 | 0.9982 | 21.232 | 24.735 | 0.0339 | 0.0155 | 0.9926 | 0.9643 | 0.9643 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | II | 1 | 0 | 0.9976 | 0.9983 | 0.9983 | 24.673 | 24.673 | 0.0106 | 0.0063 | 0.9884 | 0.9366 | 0.9366 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | aVL | 1 | 1 | 0.9976 | 0.9972 | 0.9972 | 22.48 | 22.48 | 0.0132 | 0.0117 | 0.9803 | 0.9522 | 0.9522 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | V2 | 1 | 2 | 0.9984 | 0.9968 | 0.9995 | 22.291 | 30.227 | 0.1234 | 0.0304 | 0.9957 | 0.9666 | 0.966 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | V5 | 1 | 3 | 0.9992 | 0.9922 | 0.9955 | 18.358 | 19.592 | 0.0617 | 0.0196 | 0.9979 | 0.9733 | 0.9705 | -0.25 | -2 | 2.5 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | III | 2 | 0 | 0.9976 | 0.998 | 0.998 | 23.28 | 23.28 | 0.0247 | 0.0079 | 0.9666 | 0.9486 | 0.9486 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | aVF | 2 | 1 | 0.9976 | 0.9942 | 0.9942 | 19.315 | 19.315 | 0.0138 | 0.0089 | 0.9675 | 0.9462 | 0.9462 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | V3 | 2 | 2 | 0.9984 | 0.9913 | 0.9985 | 18.044 | 25.17 | 0.0798 | 0.0301 | 0.9718 | 0.9345 | 0.9346 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | V6 | 2 | 3 | 0.9992 | 0.9963 | 0.9983 | 21.575 | 25.022 | 0.0757 | 0.0113 | 0.9947 | 0.9908 | 0.9909 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 40 | NORM | ptbxl40_clean | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261012_144921.csv | II rhythm | 3 | 0 | 0.9996 | 0.9951 | 0.9951 | 20.192 | 20.192 | 0.0247 | 0.0103 | 0.9828 | 0.9263 | 0.9263 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | I | 0 | 0 | 0.98 | 0.9855 | 0.9912 | 15.703 | 17.741 | 0.111 | 0.0226 | 0.9492 | 0.8132 | 0.8132 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | aVR | 0 | 1 | 0.9608 | 0.9703 | 0.9857 | 12.366 | 15.112 | 0.0841 | 0.0341 | 0.9052 | 0.8074 | 0.8074 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | V1 | 0 | 2 | 0.9624 | 0.9627 | 0.993 | 11.279 | 17.412 | 0.0401 | 0.0358 | 0.8952 | 0.9302 | 0.9302 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | V4 | 0 | 3 | 0.9808 | 0.9363 | 0.9611 | 9.332 | 11.185 | 0.1274 | 0.0791 | 0.8272 | 0.7224 | 0.7224 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | II | 1 | 0 | 0.98 | 0.979 | 0.9867 | 13.985 | 15.816 | 0.0486 | 0.0344 | 0.9252 | 0.7822 | 0.782 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | aVL | 1 | 1 | 0.9608 | 0.9559 | 0.9664 | 10.576 | 11.673 | 0.0206 | 0.0203 | 0.8872 | 0.6741 | 0.6741 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | V2 | 1 | 2 | 0.9624 | 0.9699 | 0.9942 | 12.465 | 19.055 | 0.0927 | 0.0581 | 0.9282 | 0.851 | 0.851 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | V5 | 1 | 3 | 0.9808 | 0.9403 | 0.9687 | 9.765 | 12.288 | 0.1077 | 0.0738 | 0.8539 | 0.7241 | 0.724 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | III | 2 | 0 | 0.98 | 0.8375 | 0.8554 | 5.094 | 5.593 | 0.1439 | 0.0371 | 0.7905 | 0.6167 | 0.6166 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | aVF | 2 | 1 | 0.9608 | 0.9612 | 0.9785 | 11.525 | 13.938 | 0.0816 | 0.0322 | 0.9122 | 0.7446 | 0.7441 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | V3 | 2 | 2 | 0.9624 | 0.9554 | 0.9796 | 10.875 | 14.022 | 0.0894 | 0.0598 | 0.8922 | 0.827 | 0.8266 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | V6 | 2 | 3 | 0.9808 | 0.9505 | 0.9794 | 10.517 | 14.041 | 0.1068 | 0.0642 | 0.8848 | 0.7532 | 0.7531 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261012_144623.csv | II rhythm | 3 | 0 | 0.9992 | 0.9713 | 0.9865 | 12.561 | 15.621 | 0.049 | 0.0403 | 0.9222 | 0.7815 | 0.7815 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | I | 0 | 0 | 0.9992 | 0.9972 | 0.9972 | 22.831 | 22.831 | 0.1082 | 0.0099 | 0.9911 | 0.9508 | 0.9508 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | aVR | 0 | 1 | 0.9976 | 0.9927 | 0.9927 | 18.55 | 18.55 | 0.0737 | 0.0167 | 0.9784 | 0.9549 | 0.9549 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | V1 | 0 | 2 | 0.9984 | 0.989 | 0.999 | 16.575 | 26.717 | 0.0286 | 0.0194 | 0.9792 | 0.9779 | 0.9779 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | V4 | 0 | 3 | 0.9992 | 0.9888 | 0.9926 | 16.839 | 18.542 | 0.1065 | 0.0333 | 0.9647 | 0.9244 | 0.9243 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | II | 1 | 0 | 0.9976 | 0.9974 | 0.9974 | 23.087 | 23.087 | 0.034 | 0.0121 | 0.9838 | 0.9112 | 0.9112 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | aVL | 1 | 1 | 0.9976 | 0.9847 | 0.9847 | 15.103 | 15.103 | 0.0125 | 0.0121 | 0.962 | 0.8668 | 0.8668 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | V2 | 1 | 2 | 0.9984 | 0.9921 | 0.9992 | 18.247 | 28.11 | 0.0749 | 0.0298 | 0.9819 | 0.9647 | 0.9648 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | V5 | 1 | 3 | 0.9992 | 0.9874 | 0.9941 | 16.395 | 19.54 | 0.0858 | 0.0344 | 0.9597 | 0.9184 | 0.9183 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | III | 2 | 0 | 0.9976 | 0.9851 | 0.9851 | 15.128 | 15.128 | 0.1421 | 0.0117 | 0.9465 | 0.7551 | 0.7551 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | aVF | 2 | 1 | 0.9976 | 0.991 | 0.991 | 17.803 | 17.803 | 0.072 | 0.0156 | 0.9731 | 0.8943 | 0.8943 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | V3 | 2 | 2 | 0.9984 | 0.9904 | 0.9975 | 17.474 | 23.13 | 0.0671 | 0.028 | 0.9727 | 0.9537 | 0.9541 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | V6 | 2 | 3 | 0.9992 | 0.9901 | 0.9955 | 17.425 | 20.687 | 0.0917 | 0.029 | 0.9684 | 0.9193 | 0.9193 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 57 | NORM | ptbxl57_clean | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261012_144548.csv | II rhythm | 3 | 0 | 0.9992 | 0.9939 | 0.9939 | 19.274 | 19.274 | 0.0357 | 0.0186 | 0.9794 | 0.9266 | 0.9266 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | I | 0 | 0 | 0.98 | 0.9885 | 0.9934 | 16.21 | 18.284 | 0.0215 | 0.0186 | 0.9381 | 0.8813 | 0.8813 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | aVR | 0 | 1 | 0.9608 | 0.9791 | 0.9899 | 13.798 | 16.529 | 0.0301 | 0.0185 | 0.9213 | 0.8426 | 0.8426 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | V1 | 0 | 2 | 0.9624 | 0.9801 | 0.9965 | 14.044 | 20.949 | 0.0355 | 0.0328 | 0.9461 | 0.9201 | 0.9201 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | V4 | 0 | 3 | 0.9808 | 0.6226 | 0.7672 | 1.414 | 1.933 | 0.1769 | 0.1758 | 0.5893 | 0.8086 | 0.8173 | 1 | 22 | 2.476 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | II | 1 | 0 | 0.98 | 0.9205 | 0.9483 | 7.984 | 8.198 | 0.0319 | 0.0286 | 0.9374 | 0.7993 | 0.7982 | -0.5 | -6 | 2.5 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | aVL | 1 | 1 | 0.9608 | 0.9855 | 0.993 | 15.393 | 18.234 | 0.0191 | 0.0182 | 0.9426 | 0.9042 | 0.9042 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | V2 | 1 | 2 | 0.9624 | 0.962 | 0.9925 | 11.22 | 17.341 | 0.0691 | 0.0549 | 0.8995 | 0.867 | 0.867 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | V5 | 1 | 3 | 0.9808 | 0.7443 | 0.7856 | 3.517 | 3.856 | 0.108 | 0.108 | 0.5853 | 0.8266 | 0.8275 | 1.25 | 26 | 2.47 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | III | 2 | 0 | 0.98 | 0.9671 | 0.9781 | 11.713 | 11.95 | 0.035 | 0.0228 | 0.9858 | 0.8603 | 0.8598 | -0.5 | -8 | 2.5 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | aVF | 2 | 1 | 0.9608 | 0.9793 | 0.9857 | 13.872 | 15.324 | 0.0373 | 0.0102 | 0.9449 | 0.8661 | 0.8655 | -0.25 | 0 | 2.5 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | V3 | 2 | 2 | 0.9624 | 0.9419 | 0.9888 | 9.402 | 15.154 | 0.0701 | 0.0683 | 0.8473 | 0.8604 | 0.8604 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | V6 | 2 | 3 | 0.9808 | 0.9054 | 0.9194 | 7.471 | 8.054 | 0.0565 | 0.0555 | 0.7632 | 0.8398 | 0.8392 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261012_154854.csv | II rhythm | 3 | 0 | 0.9992 | 0.956 | 0.9696 | 10.769 | 12.315 | 0.0348 | 0.0213 | 0.9235 | 0.8124 | 0.8124 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | I | 0 | 0 | 0.9944 | 0.9983 | 0.9983 | 24.603 | 24.603 | 0.0113 | 0.0071 | 0.9851 | 0.9782 | 0.9782 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | aVR | 0 | 1 | 0.9976 | 0.9965 | 0.9965 | 21.713 | 21.713 | 0.0241 | 0.0074 | 0.9952 | 0.9679 | 0.9679 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | V1 | 0 | 2 | 0.9984 | 0.9943 | 0.9995 | 19.464 | 29.773 | 0.0213 | 0.0176 | 0.9881 | 0.9779 | 0.9778 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | V4 | 0 | 3 | 0.9968 | 0.992 | 0.992 | 17.734 | 17.734 | 0.0294 | 0.0269 | 0.9557 | 0.96 | 0.96 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | II | 1 | 0 | 0.9976 | 0.9963 | 0.9963 | 21.349 | 21.349 | 0.0146 | 0.0061 | 0.9807 | 0.9423 | 0.9423 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | aVL | 1 | 1 | 0.9976 | 0.9968 | 0.9976 | 22.059 | 22.059 | 0.0111 | 0.0085 | 0.9928 | 0.9804 | 0.9795 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | V2 | 1 | 2 | 0.9984 | 0.9893 | 0.9992 | 16.708 | 27.343 | 0.053 | 0.0292 | 0.9729 | 0.9717 | 0.9716 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | V5 | 1 | 3 | 0.9992 | 0.9915 | 0.9968 | 17.698 | 21.757 | 0.0211 | 0.0211 | 0.9732 | 0.9672 | 0.9672 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | III | 2 | 0 | 0.9944 | 0.9976 | 0.9976 | 22.855 | 22.855 | 0.027 | 0.0063 | 0.995 | 0.937 | 0.937 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | aVF | 2 | 1 | 0.9976 | 0.9926 | 0.9926 | 18.33 | 18.33 | 0.035 | 0.0061 | 0.9761 | 0.9428 | 0.9428 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | V3 | 2 | 2 | 0.9984 | 0.9835 | 0.999 | 14.809 | 25.192 | 0.0407 | 0.0367 | 0.9532 | 0.9686 | 0.9686 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | V6 | 2 | 3 | 0.9936 | 0.9829 | 0.9899 | 14.835 | 17.104 | 0.0255 | 0.0238 | 0.9874 | 0.9561 | 0.956 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 160 | STTC | ptbxl160_clean | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261012_154835.csv | II rhythm | 3 | 0 | 0.998 | 0.9904 | 0.9904 | 17.368 | 17.368 | 0.0269 | 0.01 | 0.978 | 0.9549 | 0.9549 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | I | 0 | 0 | 0.98 | 0.9886 | 0.9925 | 16.303 | 18.004 | 0.0274 | 0.0106 | 0.951 | 0.8643 | 0.8643 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | aVR | 0 | 1 | 0.9608 | 0.9854 | 0.9941 | 15.725 | 19.334 | 0.0546 | 0.0265 | 0.9472 | 0.8861 | 0.8861 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | V1 | 0 | 2 | 0.9624 | 0.9753 | 0.9892 | 13.328 | 16.821 | 0.0926 | 0.0421 | 0.9519 | 0.9274 | 0.9281 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | V4 | 0 | 3 | 0.9808 | 0.9652 | 0.9862 | 11.932 | 15.492 | 0.0837 | 0.0539 | 0.8978 | 0.8675 | 0.8675 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | II | 1 | 0 | 0.98 | 0.9936 | 0.9971 | 19.152 | 22.269 | 0.0675 | 0.0263 | 0.9668 | 0.8922 | 0.8922 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | aVL | 1 | 1 | 0.9608 | 0.9835 | 0.9905 | 15.443 | 17.727 | 0.0537 | 0.0193 | 0.9479 | 0.9241 | 0.9242 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | V2 | 1 | 2 | 0.9624 | 0.9693 | 0.9836 | 12.518 | 15.241 | 0.168 | 0.0635 | 0.9736 | 0.809 | 0.809 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | V5 | 1 | 3 | 0.9808 | 0.9669 | 0.9862 | 11.886 | 14.918 | 0.1025 | 0.0647 | 0.8719 | 0.7787 | 0.7787 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | III | 2 | 0 | 0.98 | 0.9946 | 0.9974 | 19.916 | 22.943 | 0.0336 | 0.0211 | 0.972 | 0.9102 | 0.9102 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | aVF | 2 | 1 | 0.9608 | 0.9873 | 0.9951 | 16.356 | 20.238 | 0.0929 | 0.0392 | 0.9551 | 0.888 | 0.888 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | V3 | 2 | 2 | 0.9624 | 0.95 | 0.9724 | 10.591 | 13.168 | 0.1481 | 0.0717 | 0.9374 | 0.9357 | 0.9351 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | V6 | 2 | 3 | 0.9808 | 0.9764 | 0.9938 | 13.491 | 18.755 | 0.0811 | 0.049 | 0.9299 | 0.8803 | 0.8803 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261012_154730.csv | II rhythm | 3 | 0 | 0.9992 | 0.9869 | 0.9956 | 16.219 | 20.525 | 0.0854 | 0.0388 | 0.954 | 0.8897 | 0.8897 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | I | 0 | 0 | 0.9976 | 0.9977 | 0.9977 | 23.375 | 23.375 | 0.0261 | 0.0047 | 0.9914 | 0.9619 | 0.9619 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | aVR | 0 | 1 | 0.9976 | 0.9966 | 0.9966 | 22.133 | 22.133 | 0.0494 | 0.0127 | 0.9906 | 0.9701 | 0.9701 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | V1 | 0 | 2 | 0.9984 | 0.9952 | 0.9994 | 20.394 | 28.787 | 0.0745 | 0.0187 | 0.9858 | 0.9845 | 0.9849 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | V4 | 0 | 3 | 0.9992 | 0.9914 | 0.9981 | 18.009 | 24.301 | 0.0687 | 0.0268 | 0.9767 | 0.9792 | 0.9791 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | II | 1 | 0 | 0.9976 | 0.9995 | 0.9995 | 30.525 | 30.525 | 0.0641 | 0.0071 | 0.9935 | 0.969 | 0.969 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | aVL | 1 | 1 | 0.9976 | 0.9968 | 0.9968 | 22.6 | 22.6 | 0.0495 | 0.0085 | 0.9936 | 0.9698 | 0.9698 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | V2 | 1 | 2 | 0.9984 | 0.9953 | 0.9992 | 20.745 | 28.405 | 0.1514 | 0.0246 | 0.993 | 0.9587 | 0.9586 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | V5 | 1 | 3 | 0.9992 | 0.9936 | 0.9987 | 19.186 | 25.716 | 0.084 | 0.0279 | 0.9744 | 0.9712 | 0.9712 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | III | 2 | 0 | 0.9944 | 0.9996 | 0.9996 | 31.224 | 31.224 | 0.0304 | 0.0057 | 0.9951 | 0.9812 | 0.9812 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | aVF | 2 | 1 | 0.9976 | 0.9978 | 0.9978 | 23.976 | 23.976 | 0.0886 | 0.0163 | 0.9911 | 0.9743 | 0.9743 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | V3 | 2 | 2 | 0.9984 | 0.9909 | 0.9988 | 17.964 | 26.681 | 0.1269 | 0.0307 | 0.9788 | 0.9744 | 0.9744 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | V6 | 2 | 3 | 0.9992 | 0.9957 | 0.9987 | 20.838 | 25.827 | 0.0682 | 0.021 | 0.9791 | 0.9685 | 0.9685 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 167 | STTC | ptbxl167_clean | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261012_154711.csv | II rhythm | 3 | 0 | 0.9984 | 0.9975 | 0.9975 | 23.366 | 23.366 | 0.0806 | 0.017 | 0.9892 | 0.9726 | 0.9726 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | I | 0 | 0 | 0.98 | 0.99 | 0.9942 | 16.965 | 18.939 | 0.0657 | 0.0288 | 0.9451 | 0.8774 | 0.8773 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | aVR | 0 | 1 | 0.9608 | 0.9696 | 0.9823 | 12.388 | 14.53 | 0.0269 | 0.0248 | 0.908 | 0.8627 | 0.8627 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | V1 | 0 | 2 | 0.9624 | 0.9787 | 0.9964 | 13.853 | 20.785 | 0.0435 | 0.0324 | 0.9473 | 0.9463 | 0.9463 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | V4 | 0 | 3 | 0.9808 | 0.9506 | 0.9871 | 10.29 | 15.243 | 0.0803 | 0.054 | 0.8674 | 0.828 | 0.828 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | II | 1 | 0 | 0.98 | 0.8483 | 0.8779 | 5.175 | 5.427 | 0.0321 | 0.0317 | 0.853 | 0.8427 | 0.8313 | -0.25 | -4 | 2.5 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | aVL | 1 | 1 | 0.9608 | 0.9859 | 0.993 | 15.915 | 18.793 | 0.0791 | 0.0278 | 0.952 | 0.8918 | 0.8918 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | V2 | 1 | 2 | 0.9624 | 0.9874 | 0.9968 | 16.098 | 21.924 | 0.0487 | 0.0354 | 0.9719 | 0.8745 | 0.8745 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | V5 | 1 | 3 | 0.9808 | 0.9468 | 0.9857 | 9.959 | 14.744 | 0.063 | 0.0413 | 0.8564 | 0.8301 | 0.8301 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | III | 2 | 0 | 0.98 | 0.985 | 0.9873 | 15.648 | 16.346 | 0.0714 | 0.0306 | 0.9648 | 0.931 | 0.9309 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | aVF | 2 | 1 | 0.9608 | 0.9896 | 0.9948 | 17.096 | 19.898 | 0.0532 | 0.0118 | 0.9621 | 0.9179 | 0.9181 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | V3 | 2 | 2 | 0.9624 | 0.9559 | 0.9854 | 10.989 | 15.477 | 0.0874 | 0.0652 | 0.9041 | 0.908 | 0.908 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | V6 | 2 | 3 | 0.9808 | 0.9579 | 0.9884 | 10.936 | 15.891 | 0.026 | 0.0252 | 0.8907 | 0.8477 | 0.8479 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261012_145445.csv | II rhythm | 3 | 0 | 0.9992 | 0.9163 | 0.9437 | 8.01 | 9.672 | 0.0234 | 0.0231 | 0.8536 | 0.8226 | 0.8226 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | I | 0 | 0 | 0.9976 | 0.9985 | 0.9985 | 25.16 | 25.16 | 0.0625 | 0.0112 | 0.9844 | 0.9695 | 0.9695 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | aVR | 0 | 1 | 0.9976 | 0.9958 | 0.9967 | 20.875 | 20.875 | 0.0138 | 0.0093 | 0.9784 | 0.9707 | 0.9693 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | V1 | 0 | 2 | 0.9984 | 0.994 | 0.9993 | 19.378 | 28.635 | 0.0313 | 0.0172 | 0.9871 | 0.9799 | 0.98 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | V4 | 0 | 3 | 0.9992 | 0.9908 | 0.9961 | 17.521 | 21.095 | 0.063 | 0.0235 | 0.9674 | 0.9639 | 0.9639 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | II | 1 | 0 | 0.9992 | 0.9932 | 0.9932 | 18.424 | 18.424 | 0.0071 | 0.0069 | 0.9619 | 0.9809 | 0.9809 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | aVL | 1 | 1 | 0.9976 | 0.9983 | 0.9984 | 24.999 | 24.999 | 0.0688 | 0.0098 | 0.9823 | 0.9664 | 0.9648 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | V2 | 1 | 2 | 0.9984 | 0.9959 | 0.9991 | 20.919 | 27.287 | 0.0402 | 0.0203 | 0.9977 | 0.9721 | 0.9722 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | V5 | 1 | 3 | 0.9992 | 0.9901 | 0.9966 | 17.108 | 21.379 | 0.0511 | 0.0182 | 0.957 | 0.9528 | 0.9528 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | III | 2 | 0 | 0.9976 | 0.9988 | 0.9988 | 26.075 | 26.075 | 0.0678 | 0.0092 | 0.9921 | 0.9809 | 0.9809 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | aVF | 2 | 1 | 0.9976 | 0.998 | 0.998 | 24.434 | 24.434 | 0.0499 | 0.0051 | 0.9914 | 0.9759 | 0.9759 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | V3 | 2 | 2 | 0.9984 | 0.9869 | 0.9985 | 16.175 | 24.784 | 0.066 | 0.0359 | 0.9621 | 0.9614 | 0.9614 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | V6 | 2 | 3 | 0.9992 | 0.9918 | 0.9972 | 17.829 | 22.003 | 0.0142 | 0.0114 | 0.9643 | 0.9603 | 0.9603 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 1737 | MI | ptbxl1737_clean | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261012_145356.csv | II rhythm | 3 | 0 | 0.9992 | 0.9862 | 0.9862 | 15.633 | 15.633 | 0.0111 | 0.0096 | 0.9578 | 0.9727 | 0.9727 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | I | 0 | 0 | 0.98 | 0.9855 | 0.989 | 14.902 | 14.902 | 0.0326 | 0.0205 | 0.9072 | 0.8939 | 0.8915 | 0.25 | 4 | 2.494 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | aVR | 0 | 1 | 0.9616 | 0.9845 | 0.9871 | 15.032 | 15.482 | 0.0282 | 0.0212 | 0.9221 | 0.8965 | 0.892 | -0.25 | -2 | 2.5 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | V1 | 0 | 2 | 0.9632 | 0.9869 | 0.998 | 15.805 | 23.479 | 0.0417 | 0.0353 | 0.9659 | 0.9225 | 0.9225 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | V4 | 0 | 3 | 0.9784 | 0.9728 | 0.9913 | 13.044 | 17.869 | 0.0714 | 0.0432 | 0.9501 | 0.9802 | 0.9804 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | II | 1 | 0 | 0.98 | 0.9646 | 0.9667 | 11.687 | 11.687 | 0.0681 | 0.035 | 0.9568 | 0.8381 | 0.8361 | -0.25 | -4 | 2.5 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | aVL | 1 | 1 | 0.9616 | 0.9787 | 0.9844 | 13.327 | 13.856 | 0.0186 | 0.0182 | 0.8935 | 0.8516 | 0.8463 | -0.25 | -2 | 2.5 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | V2 | 1 | 2 | 0.9632 | 0.9875 | 0.9979 | 16.07 | 23.395 | 0.072 | 0.0452 | 0.9663 | 0.9354 | 0.9344 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | V5 | 1 | 3 | 0.9784 | 0.9546 | 0.9899 | 10.625 | 16.508 | 0.0449 | 0.0449 | 0.8952 | 0.8726 | 0.8726 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | III | 2 | 0 | 0.98 | 0.9395 | 0.9453 | 9.393 | 9.393 | 0.0485 | 0.0321 | 0.9204 | 0.9113 | 0.9066 | -0.25 | -4 | 2.5 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | aVF | 2 | 1 | 0.9616 | 0.9889 | 0.9905 | 16.647 | 16.987 | 0.0175 | 0.0145 | 0.9563 | 0.8709 | 0.8698 | -0.25 | -2 | 2.5 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | V3 | 2 | 2 | 0.9632 | 0.9842 | 0.9964 | 15.121 | 21.212 | 0.0993 | 0.0392 | 0.9578 | 0.8972 | 0.8968 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | V6 | 2 | 3 | 0.9784 | 0.9652 | 0.9934 | 11.875 | 18.522 | 0.0485 | 0.0385 | 0.9212 | 0.8782 | 0.8782 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261012_145623.csv | II rhythm | 3 | 0 | 0.999 | 0.9806 | 0.9904 | 14.431 | 17.332 | 0.0472 | 0.0274 | 0.9465 | 0.8379 | 0.8379 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | I | 0 | 0 | 0.9976 | 0.994 | 0.9968 | 19.212 | 19.212 | 0.029 | 0.0125 | 0.9756 | 0.9754 | 0.9711 | 0.25 | 2 | 2.494 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | aVR | 0 | 1 | 0.9984 | 0.9962 | 0.9962 | 21.292 | 21.292 | 0.021 | 0.0103 | 0.9806 | 0.9754 | 0.9754 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | V1 | 0 | 2 | 1 | 0.9966 | 0.9997 | 21.657 | 31.83 | 0.0262 | 0.018 | 0.988 | 0.983 | 0.983 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | V4 | 0 | 3 | 0.9976 | 0.9924 | 0.9994 | 18.53 | 29.323 | 0.055 | 0.023 | 0.9843 | 0.9793 | 0.9795 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | II | 1 | 0 | 0.9944 | 0.9966 | 0.9975 | 21.855 | 21.855 | 0.0569 | 0.0109 | 0.9896 | 0.963 | 0.9601 | 0.25 | 2 | 2.494 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | aVL | 1 | 1 | 0.9984 | 0.9935 | 0.9935 | 18.708 | 18.708 | 0.0102 | 0.0098 | 0.9661 | 0.9779 | 0.9779 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | V2 | 1 | 2 | 1 | 0.9964 | 0.9993 | 21.387 | 28.433 | 0.0586 | 0.0245 | 0.9841 | 0.9748 | 0.9744 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | V5 | 1 | 3 | 0.996 | 0.9861 | 0.9991 | 15.712 | 26.868 | 0.025 | 0.025 | 0.972 | 0.9781 | 0.9781 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | III | 2 | 0 | 0.996 | 0.9953 | 0.9955 | 20.344 | 20.344 | 0.036 | 0.0091 | 0.9786 | 0.9686 | 0.9669 | 0.25 | 2 | 2.494 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | aVF | 2 | 1 | 0.9984 | 0.9952 | 0.9961 | 20.316 | 20.316 | 0.0151 | 0.0095 | 0.9718 | 0.9466 | 0.9378 | -0.25 | -4 | 2.5 | 0 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | V3 | 2 | 2 | 1 | 0.9949 | 0.9994 | 19.989 | 28.853 | 0.0947 | 0.0224 | 0.9843 | 0.9712 | 0.9709 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | V6 | 2 | 3 | 0.9976 | 0.989 | 0.9993 | 16.839 | 28.504 | 0.0371 | 0.0217 | 0.9761 | 0.974 | 0.974 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 3908 | MI | ptbxl3908_clean | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261012_145536.csv | II rhythm | 3 | 0 | 0.9982 | 0.9932 | 0.9932 | 19.007 | 19.007 | 0.0435 | 0.0162 | 0.9806 | 0.9548 | 0.9548 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | I | 0 | 0 | 0.98 | 0.9977 | 0.9982 | 22.819 | 23.001 | 0.061 | 0.0155 | 0.978 | 0.9408 | 0.9394 | 0.25 | 4 | 2.494 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | aVR | 0 | 1 | 0.9616 | 0.9462 | 0.9537 | 8.962 | 9.402 | 0.0672 | 0.06 | 1.0382 | 0.971 | 0.9713 | 0.25 | 6 | 2.494 | 0 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | V1 | 0 | 2 | 0.9632 | 0.9941 | 0.9992 | 19.296 | 27.337 | 0.0511 | 0.0405 | 0.985 | 0.9903 | 0.9903 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | V4 | 0 | 3 | 0.9784 | 0.9937 | 0.9992 | 19.105 | 27.759 | 0.0559 | 0.0302 | 0.981 | 0.989 | 0.9891 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | II | 1 | 0 | 0.98 | 0.9827 | 0.9874 | 14.66 | 15.406 | 0.0793 | 0.0229 | 0.9828 | 0.938 | 0.9354 | 0.25 | 4 | 2.494 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | aVL | 1 | 1 | 0.9616 | 0.9582 | 0.9691 | 10.491 | 11.182 | 0.0728 | 0.0615 | 1.0107 | 0.9216 | 0.9184 | 0.5 | 10 | 2.488 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | V2 | 1 | 2 | 0.9632 | 0.9943 | 0.9992 | 19.511 | 27.885 | 0.0595 | 0.0484 | 0.985 | 0.9837 | 0.9837 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | V5 | 1 | 3 | 0.9784 | 0.9844 | 0.9959 | 15.175 | 20.75 | 0.02 | 0.0198 | 0.9596 | 0.9708 | 0.9714 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | III | 2 | 0 | 0.98 | 0.9966 | 0.9967 | 21.658 | 21.787 | 0.0176 | 0.0147 | 0.9733 | 0.9492 | 0.9492 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | aVF | 2 | 1 | 0.9616 | 0.9841 | 0.986 | 15.033 | 15.091 | 0.028 | 0.0201 | 0.9425 | 0.927 | 0.927 | -0.25 | -2 | 2.5 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | V3 | 2 | 2 | 0.9632 | 0.9946 | 0.9991 | 19.955 | 27.752 | 0.0376 | 0.0349 | 0.9864 | 0.9872 | 0.9872 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | V6 | 2 | 3 | 0.9784 | 0.9738 | 0.9957 | 12.777 | 19.881 | 0.041 | 0.0213 | 0.9422 | 0.8685 | 0.8692 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261012_150214.csv | II rhythm | 3 | 0 | 0.999 | 0.9828 | 0.9911 | 14.661 | 17.411 | 0.0248 | 0.0248 | 0.9561 | 0.9387 | 0.9387 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | I | 0 | 0 | 0.996 | 0.998 | 0.9982 | 23.799 | 23.799 | 0.0583 | 0.0138 | 0.9864 | 0.9775 | 0.9753 | 0.25 | 2 | 2.494 | 0 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | aVR | 0 | 1 | 0.9984 | 0.9968 | 0.9968 | 21.909 | 21.909 | 0.0246 | 0.0135 | 1.0049 | 0.9969 | 0.9969 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | V1 | 0 | 2 | 1 | 0.9985 | 0.9997 | 25.136 | 32.092 | 0.0406 | 0.0206 | 0.9988 | 0.9956 | 0.9955 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | V4 | 0 | 3 | 0.9976 | 0.9976 | 0.999 | 23.253 | 27.012 | 0.0483 | 0.0187 | 0.989 | 0.9991 | 0.9991 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | II | 1 | 0 | 0.9992 | 0.9932 | 0.9948 | 18.567 | 18.567 | 0.078 | 0.0146 | 1.0091 | 0.995 | 0.992 | 0.25 | 2 | 2.494 | 0 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | aVL | 1 | 1 | 0.9984 | 0.9986 | 0.9986 | 25.809 | 25.809 | 0.0333 | 0.0105 | 0.9956 | 0.9803 | 0.9803 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | V2 | 1 | 2 | 1 | 0.9986 | 0.9999 | 25.458 | 35.851 | 0.0435 | 0.0244 | 0.9974 | 0.9934 | 0.9934 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | V5 | 1 | 3 | 0.9976 | 0.9956 | 0.9996 | 20.618 | 30.007 | 0.0112 | 0.0106 | 0.9939 | 0.9911 | 0.9911 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | III | 2 | 0 | 0.9864 | 0.9825 | 0.9845 | 13.945 | 13.945 | 0.0397 | 0.0357 | 1.0469 | 0.988 | 0.9882 | 0.25 | 2 | 2.494 | 0 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | aVF | 2 | 1 | 0.9984 | 0.9948 | 0.9956 | 19.951 | 19.951 | 0.0232 | 0.0114 | 0.9793 | 0.9822 | 0.9775 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | V3 | 2 | 2 | 1 | 0.9984 | 0.9995 | 25.088 | 30.158 | 0.0193 | 0.0193 | 0.9978 | 0.9937 | 0.9937 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | V6 | 2 | 3 | 0.9976 | 0.9915 | 0.9993 | 17.73 | 28.366 | 0.0352 | 0.012 | 0.9836 | 0.9634 | 0.963 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 4016 | CD | ptbxl4016_clean | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261012_150148.csv | II rhythm | 3 | 0 | 0.9986 | 0.994 | 0.994 | 19.18 | 19.18 | 0.0147 | 0.0147 | 0.9875 | 0.9891 | 0.9891 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | I | 0 | 0 | 0.98 | 0.9534 | 0.9554 | 10.338 | 10.42 | 0.0349 | 0.0275 | 0.9391 | 0.8319 | 0.8317 | 0.25 | 4 | 2.494 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | aVR | 0 | 1 | 0.9608 | 0.9912 | 0.9953 | 17.691 | 20.062 | 0.0426 | 0.0288 | 0.9508 | 0.9443 | 0.9443 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | V1 | 0 | 2 | 0.9624 | 0.9855 | 0.9954 | 15.296 | 19.575 | 0.0574 | 0.0493 | 0.936 | 0.8959 | 0.8959 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | V4 | 0 | 3 | 0.9808 | 0.994 | 0.9963 | 19.215 | 20.953 | 0.1292 | 0.0279 | 0.952 | 0.9836 | 0.9835 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | II | 1 | 0 | 0.98 | 0.9585 | 0.9607 | 10.899 | 11.084 | 0.2095 | 0.1393 | 0.8415 | 0.774 | 0.7741 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | aVL | 1 | 1 | 0.9608 | 0.9901 | 0.9949 | 16.705 | 19.312 | 0.0657 | 0.0339 | 1.0271 | 1.041 | 1.0411 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | V2 | 1 | 2 | 0.9624 | 0.9788 | 0.9868 | 12.473 | 13.681 | 0.114 | 0.105 | 0.8405 | 0.7447 | 0.7448 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | V5 | 1 | 3 | 0.9808 | 0.9942 | 0.9967 | 19.241 | 21.261 | 0.1034 | 0.0281 | 0.9508 | 0.9561 | 0.9561 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | III | 2 | 0 | 0.98 | 0.9373 | 0.9389 | 9 | 9.077 | 0.2143 | 0.1772 | 0.7795 | 0.6978 | 0.695 | 0.25 | 4 | 2.494 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | aVF | 2 | 1 | 0.9608 | 0.9824 | 0.9861 | 12.142 | 12.558 | 0.1286 | 0.1073 | 0.7956 | 0.6936 | 0.6936 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | V3 | 2 | 2 | 0.9624 | 0.9864 | 0.9946 | 15.684 | 19.204 | 0.1018 | 0.0443 | 0.9406 | 0.8227 | 0.8227 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | V6 | 2 | 3 | 0.9808 | 0.994 | 0.9968 | 19.047 | 21.264 | 0.08 | 0.0211 | 0.9513 | 0.9625 | 0.9625 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261012_155118.csv | II rhythm | 3 | 0 | 0.9992 | 0.9816 | 0.9852 | 11.847 | 12.214 | 0.1485 | 0.1135 | 0.787 | 0.6811 | 0.6812 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | I | 0 | 0 | 0.9992 | 0.9836 | 0.9836 | 14.481 | 14.481 | 0.0277 | 0.0171 | 1.0133 | 0.9551 | 0.9551 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | aVR | 0 | 1 | 0.9976 | 0.9987 | 0.9987 | 26.265 | 26.265 | 0.0312 | 0.0107 | 0.9882 | 0.9775 | 0.9775 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | V1 | 0 | 2 | 0.9984 | 0.9985 | 0.9985 | 25.035 | 25.035 | 0.0338 | 0.0161 | 0.9826 | 0.963 | 0.963 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | V4 | 0 | 3 | 0.9992 | 0.9995 | 0.9995 | 29.667 | 29.667 | 0.1235 | 0.0084 | 0.9886 | 0.9925 | 0.9925 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | II | 1 | 0 | 0.9976 | 0.9979 | 0.9979 | 23.173 | 23.173 | 0.1544 | 0.0339 | 0.9646 | 0.9753 | 0.9753 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | aVL | 1 | 1 | 0.9976 | 0.9895 | 0.9899 | 16.733 | 16.733 | 0.061 | 0.0338 | 0.9504 | 0.98 | 0.9634 | -0.25 | -4 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | V2 | 1 | 2 | 0.9984 | 0.9988 | 0.9988 | 25.847 | 25.847 | 0.0444 | 0.0225 | 0.9873 | 0.9793 | 0.9793 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | V5 | 1 | 3 | 0.9992 | 0.9996 | 0.9996 | 30.909 | 30.909 | 0.1013 | 0.0073 | 0.9894 | 0.9896 | 0.9896 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | III | 2 | 0 | 0.9976 | 0.9962 | 0.9962 | 19.937 | 19.937 | 0.1235 | 0.0503 | 0.9384 | 0.9753 | 0.9753 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | aVF | 2 | 1 | 0.9976 | 0.9939 | 0.9939 | 16.152 | 16.152 | 0.097 | 0.0676 | 0.8745 | 0.8702 | 0.8702 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | V3 | 2 | 2 | 0.9984 | 0.9962 | 0.9962 | 21.238 | 21.238 | 0.0937 | 0.0234 | 0.9973 | 0.9477 | 0.9477 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | V6 | 2 | 3 | 0.9992 | 0.9924 | 0.9924 | 17.479 | 17.479 | 0.0767 | 0.0253 | 0.9218 | 0.9875 | 0.9875 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 5830 | HYP | ptbxl5830_clean | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261012_155102.csv | II rhythm | 3 | 0 | 0.9992 | 0.9988 | 0.9988 | 26.386 | 26.386 | 0.1123 | 0.0213 | 0.9886 | 0.9838 | 0.9838 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | I | 0 | 0 | 0.98 | 0.976 | 0.9804 | 13.365 | 14.198 | 0.0862 | 0.0551 | 0.9335 | 0.9219 | 0.922 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | aVR | 0 | 1 | 0.9608 | 0.9767 | 0.9842 | 13.667 | 15.293 | 0.0748 | 0.047 | 0.9401 | 0.8909 | 0.8908 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | V1 | 0 | 2 | 0.9624 | 0.9682 | 0.9859 | 10.419 | 11.929 | 0.1773 | 0.1502 | 0.7654 | 0.6971 | 0.697 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | V4 | 0 | 3 | 0.9808 | 0.9608 | 0.988 | 11.246 | 15.544 | 0.1131 | 0.0812 | 0.8772 | 0.8965 | 0.8962 | 0 | 4 | 2.496 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | II | 1 | 0 | 0.98 | 0.9697 | 0.9729 | 12.565 | 13.044 | 0.064 | 0.0475 | 0.9299 | 0.8987 | 0.8988 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | aVL | 1 | 1 | 0.9608 | 0.9707 | 0.9822 | 12.477 | 14.455 | 0.0435 | 0.0381 | 0.9091 | 0.8632 | 0.8633 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | V2 | 1 | 2 | 0.9624 | 0.9634 | 0.9836 | 11.186 | 13.981 | 0.1232 | 0.1235 | 0.8654 | 0.7344 | 0.734 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | V5 | 1 | 3 | 0.9808 | 0.982 | 0.9949 | 14.731 | 19.745 | 0.0991 | 0.0557 | 0.9426 | 0.9402 | 0.9402 | 0 | 4 | 2.496 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | III | 2 | 0 | 0.98 | 0.907 | 0.9204 | 7.531 | 8.191 | 0.048 | 0.0375 | 0.8541 | 0.7508 | 0.7508 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | aVF | 2 | 1 | 0.9608 | 0.9513 | 0.9631 | 10.62 | 11.76 | 0.0498 | 0.0271 | 0.8858 | 0.8724 | 0.8724 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | V3 | 2 | 2 | 0.9624 | 0.9666 | 0.9877 | 11.562 | 14.931 | 0.1196 | 0.1082 | 0.8709 | 0.849 | 0.8489 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | V6 | 2 | 3 | 0.9808 | 0.984 | 0.9951 | 15.254 | 20.07 | 0.1026 | 0.0396 | 0.9528 | 0.9648 | 0.9648 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261012_155008.csv | II rhythm | 3 | 0 | 0.9992 | 0.9723 | 0.9801 | 12.939 | 14.309 | 0.0832 | 0.0461 | 0.9301 | 0.9183 | 0.9183 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | I | 0 | 0 | 0.996 | 0.9943 | 0.9943 | 19.596 | 19.596 | 0.0661 | 0.0269 | 0.9795 | 0.9818 | 0.9818 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | aVR | 0 | 1 | 0.9976 | 0.9902 | 0.9909 | 17.387 | 17.387 | 0.0679 | 0.0306 | 0.9749 | 0.9647 | 0.9557 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | V1 | 0 | 2 | 0.9984 | 0.9934 | 0.9993 | 18.972 | 28.158 | 0.1247 | 0.0561 | 0.9772 | 0.9843 | 0.9843 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | V4 | 0 | 3 | 0.9992 | 0.9912 | 0.9971 | 17.737 | 22.38 | 0.0893 | 0.0385 | 0.9695 | 0.9718 | 0.9718 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | II | 1 | 0 | 0.9944 | 0.9912 | 0.9912 | 17.873 | 17.873 | 0.0478 | 0.0258 | 0.9816 | 0.9794 | 0.9794 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | aVL | 1 | 1 | 0.9976 | 0.9912 | 0.992 | 17.672 | 17.672 | 0.0278 | 0.0209 | 0.9674 | 0.9644 | 0.9599 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | V2 | 1 | 2 | 0.9888 | 0.9896 | 0.9978 | 16.852 | 23.481 | 0.0643 | 0.0643 | 0.978 | 0.9507 | 0.9503 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | V5 | 1 | 3 | 0.9968 | 0.9963 | 0.9985 | 21.55 | 25.426 | 0.0871 | 0.0254 | 0.9848 | 0.987 | 0.987 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | III | 2 | 0 | 0.9976 | 0.9919 | 0.9919 | 17.521 | 17.521 | 0.0329 | 0.0119 | 0.9434 | 0.9243 | 0.9243 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | aVF | 2 | 1 | 0.9976 | 0.981 | 0.981 | 14.663 | 14.663 | 0.0452 | 0.017 | 0.9543 | 0.9652 | 0.9652 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | V3 | 2 | 2 | 0.9984 | 0.9914 | 0.9989 | 17.694 | 26.09 | 0.0792 | 0.0534 | 0.9757 | 0.9685 | 0.9683 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | V6 | 2 | 3 | 0.9992 | 0.9965 | 0.9983 | 21.859 | 24.974 | 0.0978 | 0.0185 | 0.9852 | 0.9932 | 0.9932 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 7916 | HYP | ptbxl7916_clean | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261012_154945.csv | II rhythm | 3 | 0 | 0.9976 | 0.9862 | 0.9862 | 15.949 | 15.949 | 0.0769 | 0.0326 | 0.9826 | 0.9867 | 0.9867 | 0 | 0 | 10 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | I | 0 | 0 | 0.98 | 0.9669 | 0.9669 | 11.926 | 11.926 | 0.0402 | 0.0342 | 0.941 | 0.876 | 0.876 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | aVR | 0 | 1 | 0.9624 | 0.9714 | 0.9819 | 12.55 | 14.053 | 0.0575 | 0.0341 | 0.9389 | 0.8791 | 0.8787 | -0.25 | 0 | 2.5 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | V1 | 0 | 2 | 0.9624 | 0.8917 | 0.9642 | 6.526 | 8.68 | 0.1063 | 0.1033 | 0.9143 | 0.9463 | 0.9441 | 1 | 20 | 2.476 | -6 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | V4 | 0 | 3 | 0.9792 | 0.9462 | 0.9895 | 10.04 | 15.628 | 0.171 | 0.116 | 0.848 | 0.8147 | 0.8148 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | II | 1 | 0 | 0.98 | 0.9821 | 0.9822 | 14.638 | 14.638 | 0.0621 | 0.0282 | 0.9611 | 0.8279 | 0.8176 | 0.25 | 4 | 2.494 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | aVL | 1 | 1 | 0.9608 | 0.9822 | 0.9829 | 14.373 | 14.373 | 0.0255 | 0.0141 | 0.9365 | 0.8116 | 0.8112 | -0.25 | -2 | 2.5 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | V2 | 1 | 2 | 0.964 | 0.9044 | 0.94 | 7.718 | 8.864 | 0.086 | 0.0671 | 0.7718 | 0.866 | 0.8618 | 0.5 | 10 | 2.488 | -4 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | V5 | 1 | 3 | 0.9792 | 0.9483 | 0.9944 | 10.242 | 18.609 | 0.1423 | 0.0896 | 0.8867 | 0.8995 | 0.8995 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | III | 2 | 0 | 0.98 | 0.9101 | 0.9194 | 7.411 | 7.411 | 0.084 | 0.0228 | 0.9253 | 0.7791 | 0.7791 | -0.25 | -4 | 2.5 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | aVF | 2 | 1 | 0.9608 | 0.9892 | 0.992 | 16.636 | 16.673 | 0.0573 | 0.0209 | 0.9875 | 0.7785 | 0.7748 | 0.25 | 4 | 2.494 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | V3 | 2 | 2 | 0.964 | 0.8216 | 0.9269 | 5.153 | 6.444 | 0.1558 | 0.1544 | 0.7003 | 0.8782 | 0.8835 | 1 | 18 | 2.476 | -4 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | V6 | 2 | 3 | 0.9792 | 0.9532 | 0.9957 | 10.615 | 19.591 | 0.089 | 0.0501 | 0.8976 | 0.9049 | 0.9049 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261012_154516.csv | II rhythm | 3 | 0 | 0.9992 | 0.9846 | 0.992 | 15.149 | 17.954 | 0.0322 | 0.0319 | 0.9634 | 0.8343 | 0.8343 | 0 | 2 | 9.998 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | I | 0 | 0 | 0.9976 | 0.9959 | 0.9959 | 20.899 | 20.899 | 0.0234 | 0.0122 | 0.9806 | 0.9781 | 0.9781 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | aVR | 0 | 1 | 0.9992 | 0.9843 | 0.9876 | 15.135 | 15.135 | 0.0545 | 0.0254 | 0.9751 | 0.9762 | 0.9743 | -0.25 | -2 | 2.5 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | V1 | 0 | 2 | 0.9992 | 0.9743 | 0.991 | 12.908 | 17.483 | 0.0573 | 0.0496 | 0.9716 | 0.9857 | 0.9851 | 0 | 4 | 2.496 | -4 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | V4 | 0 | 3 | 0.9968 | 0.9781 | 0.9977 | 13.969 | 23.574 | 0.1441 | 0.0738 | 0.9642 | 0.9798 | 0.9798 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | II | 1 | 0 | 0.9976 | 0.9956 | 0.9956 | 20.637 | 20.637 | 0.0582 | 0.0141 | 0.9816 | 0.9607 | 0.9607 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | aVL | 1 | 1 | 0.9976 | 0.9914 | 0.9915 | 17.654 | 17.654 | 0.0218 | 0.0096 | 0.9734 | 0.9577 | 0.9534 | -0.25 | -4 | 2.5 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | V2 | 1 | 2 | 1 | 0.9925 | 0.995 | 18.6 | 20.401 | 0.0627 | 0.0192 | 0.9721 | 0.9751 | 0.975 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | V5 | 1 | 3 | 0.9984 | 0.9784 | 0.9978 | 13.972 | 23.579 | 0.1257 | 0.0583 | 0.9632 | 0.9798 | 0.9799 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | III | 2 | 0 | 0.996 | 0.986 | 0.986 | 15.172 | 15.172 | 0.0839 | 0.0093 | 0.9407 | 0.8971 | 0.8971 | 0 | 0 | 2.5 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | aVF | 2 | 1 | 0.9976 | 0.9973 | 0.9978 | 22.655 | 23.632 | 0.0529 | 0.0105 | 0.9951 | 0.917 | 0.9171 | 0 | -2 | 2.498 | 0 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | V3 | 2 | 2 | 1 | 0.9938 | 0.9978 | 19.303 | 23.748 | 0.0477 | 0.0303 | 0.9736 | 0.9792 | 0.9791 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | V6 | 2 | 3 | 0.9984 | 0.9809 | 0.998 | 14.461 | 24.09 | 0.0814 | 0.0322 | 0.9655 | 0.9811 | 0.9811 | 0 | 2 | 2.498 | -2 | ok | ok
A (clean images) | 8504 | CD | ptbxl8504_clean | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261012_154443.csv | II rhythm | 3 | 0 | 0.9992 | 0.995 | 0.995 | 20.032 | 20.032 | 0.0191 | 0.0182 | 0.9884 | 0.9463 | 0.9463 | 0 | 0 | 10 | 0 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | I | 0 | 0 | 0.9792 | 0.9353 | 0.9718 | 9.049 | 10.357 | 0.0569 | 0.0289 | 0.9227 | 0.947 | 0.947 | 0.75 | 14 | 2.482 | -4 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | aVR | 0 | 1 | 0.9632 | 0.9561 | 0.9644 | 10.759 | 11.593 | 0.0365 | 0.0278 | 0.891 | 0.8603 | 0.8516 | 0.25 | 0 | 2.494 | 2 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | V1 | 0 | 2 | 0.9632 | 0.1611 | 0.1926 | -7.064 | -6.987 | 0.3741 | 0.3566 | 0.3032 | 0.8072 | 0.8047 | -0.75 | -2 | 2.5 | -6 | UNRELIABLE | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | V4 | 0 | 3 | 0.9056 | 0.4342 | 0.4538 | 0.333 | 0.483 | 0.2239 | 0.2205 | 0.3589 | 0.0331 | 0.0332 | 0 | 4 | 2.496 | -4 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | II | 1 | 0 | 0.9792 | 0.8762 | 0.9032 | 6.498 | 7.394 | 0.0655 | 0.05 | 0.7322 | 0.5656 | 0.5656 | 0.25 | 10 | 2.49 | -6 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | aVL | 1 | 1 | 0.9648 | 0.9607 | 0.9607 | 10.058 | 10.058 | 0.0177 | 0.0179 | 0.8656 | 0.8306 | 0.8306 | 0 | 0 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | V2 | 1 | 2 | 0.9624 | 0.2518 | 0.2989 | -0.289 | -0.278 | 0.3416 | 0.3326 | 0.145 | -0.0416 | 0.1024 | 6.25 | 100 | 2.352 | -120 | UNRELIABLE | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | V5 | 1 | 3 | 0.9792 | 0.7264 | 0.7381 | 3.225 | 3.248 | 0.1168 | 0.1179 | 0.4968 | 0.0248 | 0.018 | 1 | 6 | 2.476 | -2 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | III | 2 | 0 | 0.044 | -0.0778 | 0 | -0.049 | -0.024 | 0.07 | 0.0658 | -0.0051 | 0 | 0 | -10 | 88 | 2.412 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | aVF | 2 | 1 | 0.9648 | 0.9322 | 0.9429 | 8.98 | 9.422 | 0.0318 | 0.031 | 0.8312 | 0.8893 | 0.888 | 0.25 | 6 | 2.494 | -2 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | V3 | 2 | 2 | 0.964 | 0.3758 | 0.4348 | -1.319 | -0.968 | 0.3601 | 0.2992 | 0.3741 | 0.5573 | 0.5541 | 0.25 | 10 | 2.49 | -6 | UNRELIABLE | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | V6 | 2 | 3 | 0.9032 | 0.1217 | 0.1939 | -4.137 | -4.052 | 0.2027 | 0.1975 | 0.1707 | 0.2805 | 0.2752 | 3 | 36 | 2.428 | -4 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv | II rhythm | 3 | 0 | 0.855 | 0.7861 | 0.8885 | 4.439 | 4.443 | 0.0802 | 0.0714 | 0.6988 | 0.9072 | 0.9007 | 0.5 | 28 | 9.95 | 2 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | I | 0 | 0 | 0.9968 | 0.9696 | 0.9841 | 12.276 | 13.939 | 0.0499 | 0.02 | 0.98 | 0.9415 | 0.9394 | 0.25 | 6 | 2.494 | -2 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | aVR | 0 | 1 | 0.9984 | 0.9722 | 0.9863 | 12.669 | 15.682 | 0.0329 | 0.0223 | 0.9647 | 0.9461 | 0.9457 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | V1 | 0 | 2 | 0.9984 | 0.9108 | 0.9965 | 7.593 | 17.981 | 0.0658 | 0.066 | 0.8838 | 0.9525 | 0.9464 | -0.25 | 2 | 2.498 | -6 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | V4 | 0 | 3 | 0.9976 | 0.6886 | 0.9822 | 2.393 | 3.808 | 0.1744 | 0.1739 | 0.6627 | 0.9175 | 0.9095 | 1.25 | 6 | 2.47 | 8 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | II | 1 | 0 | 0.9968 | 0.994 | 0.9962 | 19.154 | 19.154 | 0.0442 | 0.0116 | 0.978 | 0.9555 | 0.9517 | 0.25 | 4 | 2.494 | 0 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | aVL | 1 | 1 | 0.9992 | 0.9316 | 0.9697 | 8.611 | 11.798 | 0.0231 | 0.0211 | 0.8752 | 0.9187 | 0.906 | -0.25 | -6 | 2.5 | 4 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | V2 | 1 | 2 | 0.9968 | 0.9259 | 0.9882 | 8.464 | 15.795 | 0.1281 | 0.1214 | 0.8652 | 0.8814 | 0.8814 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | V5 | 1 | 3 | 0.9976 | 0.7221 | 0.9958 | 2.864 | 4.789 | 0.1229 | 0.1229 | 0.6942 | 0.9383 | 0.9393 | 1.25 | 4 | 2.47 | 10 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | III | 2 | 0 | 0.9968 | 0.9776 | 0.9871 | 13.562 | 13.562 | 0.0144 | 0.0137 | 0.9424 | 0.916 | 0.9096 | -0.5 | -6 | 2.5 | 0 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | aVF | 2 | 1 | 0.9992 | 0.9892 | 0.9892 | 16.928 | 16.928 | 0.0177 | 0.0124 | 0.9794 | 0.9593 | 0.9593 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | V3 | 2 | 2 | 0.9976 | 0.9128 | 0.9913 | 7.898 | 16.972 | 0.1419 | 0.1036 | 0.8535 | 0.8931 | 0.8845 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | V6 | 2 | 3 | 0.9968 | 0.8859 | 0.991 | 6.68 | 7.749 | 0.0602 | 0.0568 | 0.8599 | 0.9571 | 0.9539 | 1 | 8 | 2.476 | 4 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113100 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv | II rhythm | 3 | 0 | 0.9986 | 0.8759 | 0.9628 | 6.529 | 6.529 | 0.0697 | 0.0561 | 0.8561 | 0.9739 | 0.9704 | 0.25 | 12 | 9.976 | 0 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | I | 0 | 0 | 0.9768 | 0.8189 | 0.9626 | 3.983 | 8.125 | 0.0687 | 0.0518 | 0.9089 | 0.9178 | 0.921 | -0.5 | -12 | 2.5 | 8 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | aVR | 0 | 1 | 0.96 | 0.3927 | 0.4181 | -4.862 | -4.831 | 0.1753 | 0.1681 | 0.7326 | 0.8612 | 0.8598 | 0.75 | 14 | 2.482 | -4 | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | V1 | 0 | 2 | 0.6704 | 0.2074 | 0.331 | -4.193 | -4.022 | 0.2667 | 0.2562 | 0.2763 | 0.1427 | 0.2742 | -10 | -66 | 2.5 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | V4 | 0 | 3 | 0.9776 | 0.2659 | 0.2707 | -1.845 | -1.695 | 0.2754 | 0.2833 | 0.2697 | -0.0617 | -0.0577 | 0 | 8 | 2.492 | -10 | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | II | 1 | 0 | 0.9768 | 0.8422 | 0.9099 | 5.504 | 7.817 | 0.0672 | 0.0561 | 0.7645 | 0.9157 | 0.9157 | 0 | -8 | 2.492 | 8 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | aVL | 1 | 1 | 0.96 | 0.962 | 0.962 | 10.699 | 10.699 | 0.0193 | 0.0166 | 0.8584 | 0.7489 | 0.7489 | 0 | 0 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | V2 | 1 | 2 | 0.9584 | 0.2006 | 0.2419 | -1.835 | -1.734 | 0.4579 | 0.3974 | 0.1629 | 0.0337 | 0.0337 | -2.5 | -24 | 2.5 | -26 | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | V5 | 1 | 3 | 0.98 | 0.5473 | 0.5517 | 1.399 | 1.422 | 0.1428 | 0.1454 | 0.34 | 0 | 0 | 0.25 | 6 | 2.494 | -4 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | III | 2 | 0 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | aVF | 2 | 1 | 0.712 | 0.5676 | 0.5711 | 1.876 | 1.898 | 0.0702 | 0.0702 | 0.3781 | 0.0344 | 0.036 | -0.75 | -18 | 2.5 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | V3 | 2 | 2 | 0.96 | 0.2484 | 0.2941 | -3.793 | -3.622 | 0.3721 | 0.3978 | 0.3268 | -0.1301 | -0.151 | 2.25 | 42 | 2.444 | -14 | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | V6 | 2 | 3 | 0.9768 | 0.674 | 0.7265 | 2.709 | 2.709 | 0.0893 | 0.0898 | 0.4011 | 0.2458 | 0.2458 | 1.5 | 16 | 2.464 | 0 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113551.csv | II rhythm | 3 | 0 | 0.691 | 0.6852 | 0.6993 | 3.113 | 3.113 | 0.0894 | 0.0831 | 0.4813 | 0.1786 | 0.1755 | 0.5 | 30 | 9.95 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | I | 0 | 0 | 0.9952 | 0.9469 | 0.9849 | 9.811 | 14.292 | 0.0524 | 0.0265 | 0.9647 | 0.95 | 0.9478 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | aVR | 0 | 1 | 0.9984 | 0.9916 | 0.9916 | 17.759 | 17.759 | 0.0261 | 0.0124 | 0.9622 | 0.9383 | 0.9383 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | V1 | 0 | 2 | 0.9992 | 0.8793 | 0.9949 | 6.275 | 18.695 | 0.0764 | 0.0768 | 0.854 | 0.9503 | 0.9486 | -0.25 | 4 | 2.496 | -6 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | V4 | 0 | 3 | 0.9976 | 0.9459 | 0.9846 | 9.94 | 10.209 | 0.0733 | 0.073 | 0.8913 | 0.9 | 0.8965 | 0.5 | 4 | 2.488 | 2 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | II | 1 | 0 | 0.9952 | 0.9247 | 0.9918 | 8.467 | 18.045 | 0.0568 | 0.0399 | 0.9168 | 0.9608 | 0.9593 | -0.25 | -10 | 2.496 | 8 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | aVL | 1 | 1 | 0.9984 | 0.9498 | 0.979 | 9.832 | 11.863 | 0.0206 | 0.0183 | 0.8857 | 0.8859 | 0.8575 | -0.25 | -6 | 2.5 | 2 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | V2 | 1 | 2 | 0.9968 | 0.8773 | 0.9941 | 6.353 | 17.642 | 0.1598 | 0.1548 | 0.8065 | 0.8585 | 0.8587 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | V5 | 1 | 3 | 0.9992 | 0.898 | 0.9961 | 7.221 | 9.144 | 0.0743 | 0.0744 | 0.8486 | 0.9329 | 0.9178 | 0.75 | 4 | 2.482 | 4 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | III | 2 | 0 | 0.9952 | 0.9095 | 0.9834 | 7.552 | 13.536 | 0.0276 | 0.0274 | 0.8941 | 0.953 | 0.946 | -0.5 | -12 | 2.5 | 6 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | aVF | 2 | 1 | 0.9984 | 0.9806 | 0.9906 | 14.387 | 17.303 | 0.0221 | 0.0166 | 0.9511 | 0.948 | 0.9445 | -0.25 | -6 | 2.5 | 2 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | V3 | 2 | 2 | 0.9984 | 0.8749 | 0.9928 | 6.419 | 16.568 | 0.1558 | 0.1228 | 0.7886 | 0.8402 | 0.8396 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | V6 | 2 | 3 | 0.9976 | 0.929 | 0.9952 | 8.794 | 10.109 | 0.0476 | 0.0446 | 0.884 | 0.9311 | 0.9299 | 0.75 | 6 | 2.482 | 4 | ok | ok
B (ideal photos) | 9 | NORM | pair_20261013_113548 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv | II rhythm | 3 | 0 | 0.999 | 0.896 | 0.9689 | 7.323 | 7.323 | 0.0666 | 0.0512 | 0.8662 | 0.9737 | 0.963 | 0.25 | 12 | 9.976 | 0 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | I | 0 | 0 | 0.98 | 0.7959 | 0.8477 | 3.925 | 4.13 | 0.0481 | 0.0333 | 0.7823 | 0.8649 | 0.864 | 0.75 | 10 | 2.482 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | aVR | 0 | 1 | 0.9616 | 0.9195 | 0.9285 | 8.071 | 8.453 | 0.0376 | 0.0335 | 0.8735 | 0.7511 | 0.7501 | -0.25 | -2 | 2.5 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | V1 | 0 | 2 | 0.9512 | 0.3408 | 0.368 | -3.95 | -3.811 | 0.4002 | 0.3609 | 0.4977 | 0.9211 | 0.9216 | 0 | 10 | 2.49 | -6 | UNRELIABLE | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | V4 | 0 | 3 | 0.9368 | 0.5476 | 0.6097 | 1.441 | 1.506 | 0.1708 | 0.1732 | 0.2883 | 0.0814 | 0.0736 | 5 | 42 | 2.38 | 4 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | II | 1 | 0 | 0.9784 | 0.773 | 0.8834 | 3.943 | 5.297 | 0.0858 | 0.0792 | 0.654 | 0.3337 | 0.3293 | 0.75 | 12 | 2.482 | -8 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | aVL | 1 | 1 | 0.9632 | 0.9365 | 0.9398 | 8.996 | 9.061 | 0.0192 | 0.019 | 0.8255 | 0.7968 | 0.7964 | -0.25 | -4 | 2.5 | 0 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | V2 | 1 | 2 | 0.9608 | 0.7503 | 0.8067 | 3.03 | 3.276 | 0.2077 | 0.2073 | 0.7472 | 0.6955 | 0.6951 | 2 | 26 | 2.45 | -6 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | V5 | 1 | 3 | 0.9424 | 0.6334 | 0.8002 | 2.178 | 2.517 | 0.1249 | 0.1267 | 0.3957 | 0.3894 | 0.7403 | 6.25 | 38 | 2.352 | 8 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | III | 2 | 0 | 0.9808 | 0.7422 | 0.8639 | 3.023 | 3.176 | 0.0808 | 0.0632 | 0.7161 | 0.9078 | 0.9048 | 2 | 18 | 2.45 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | aVF | 2 | 1 | 0.9536 | 0.8399 | 0.8483 | 5.226 | 5.452 | 0.0702 | 0.0658 | 0.769 | 0.6702 | 0.6701 | 0 | -2 | 2.498 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | V3 | 2 | 2 | 0.9544 | 0.5596 | 0.58 | 1.629 | 1.671 | 0.1849 | 0.1842 | 0.3459 | 0.0771 | 0.2233 | 5 | 58 | 2.38 | -4 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | V6 | 2 | 3 | 0.932 | 0.5572 | 0.6962 | 1.709 | 1.94 | 0.0907 | 0.0905 | 0.3792 | 0.6893 | 0.7713 | 6.75 | 46 | 2.342 | 4 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115458.csv | II rhythm | 3 | 0 | 0.9892 | 0.5403 | 0.7504 | 0.45 | 0.736 | 0.1277 | 0.1275 | 0.5383 | 0.6524 | 0.6936 | 0.5 | 26 | 9.95 | -10 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | I | 0 | 0 | 0.9984 | 0.8928 | 0.937 | 6.456 | 6.704 | 0.0411 | 0.0249 | 0.9202 | 0.9387 | 0.9296 | 0.5 | 8 | 2.488 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | aVR | 0 | 1 | 0.9984 | 0.9716 | 0.9853 | 12.558 | 12.558 | 0.0299 | 0.02 | 0.9378 | 0.9218 | 0.9201 | -0.25 | -4 | 2.5 | 0 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | V1 | 0 | 2 | 0.9984 | 0.9626 | 0.9957 | 11.268 | 17.481 | 0.0659 | 0.0626 | 0.9585 | 0.9886 | 0.9886 | 0.25 | 10 | 2.49 | -6 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | V4 | 0 | 3 | 0.9992 | 0.8519 | 0.9953 | 5.581 | 10.095 | 0.1075 | 0.1075 | 0.8023 | 0.9126 | 0.8966 | 1.25 | 6 | 2.47 | 10 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | II | 1 | 0 | 0.9976 | 0.8994 | 0.9801 | 7.154 | 7.808 | 0.064 | 0.0547 | 0.8646 | 0.8568 | 0.8464 | 0.5 | 8 | 2.488 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | aVL | 1 | 1 | 0.9984 | 0.9676 | 0.9794 | 11.758 | 12.431 | 0.0139 | 0.0138 | 0.8854 | 0.8821 | 0.8795 | -0.25 | -6 | 2.5 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | V2 | 1 | 2 | 0.9976 | 0.9728 | 0.9963 | 12.712 | 17.493 | 0.072 | 0.068 | 0.9505 | 0.9455 | 0.9441 | 0.25 | 8 | 2.492 | -4 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | V5 | 1 | 3 | 0.9984 | 0.8903 | 0.9951 | 6.905 | 11.231 | 0.073 | 0.0735 | 0.8338 | 0.8745 | 0.8903 | 1.25 | 6 | 2.47 | 10 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | III | 2 | 0 | 0.9992 | 0.9243 | 0.9649 | 8.18 | 8.408 | 0.0639 | 0.0349 | 0.9275 | 0.9717 | 0.9711 | 0.5 | 6 | 2.488 | 0 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | aVF | 2 | 1 | 0.9976 | 0.9562 | 0.9836 | 10.679 | 12.795 | 0.0387 | 0.0351 | 0.9141 | 0.8092 | 0.8074 | -0.25 | -8 | 2.498 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | V3 | 2 | 2 | 0.9968 | 0.9237 | 0.9826 | 8.304 | 14.38 | 0.0874 | 0.0854 | 0.877 | 0.9391 | 0.9387 | 0 | 4 | 2.496 | -4 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | V6 | 2 | 3 | 0.9976 | 0.9323 | 0.9945 | 8.993 | 11.464 | 0.0397 | 0.0391 | 0.8937 | 0.9414 | 0.9289 | 1 | 6 | 2.476 | 6 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115454 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv | II rhythm | 3 | 0 | 0.9998 | 0.7536 | 0.9532 | 3.085 | 3.257 | 0.095 | 0.0941 | 0.7639 | 0.8918 | 0.8982 | 0.5 | 24 | 9.95 | -4 | ok | check | NO_PULSE/warn
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | I | 0 | 0 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | aVR | 0 | 1 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | V1 | 0 | 2 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | V4 | 0 | 3 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | II | 1 | 0 | 0.9808 | 0.7185 | 0.8837 | 2.484 | 2.549 | 0.0919 | 0.0937 | 0.7006 | 0.2728 | 0.7328 | 1.75 | 20 | 2.458 | -2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | aVL | 1 | 1 | 0.8576 | 0.8724 | 0.8779 | 5.278 | 5.376 | 0.0264 | 0.0291 | 0.7237 | 0.6526 | 0.6541 | 0.5 | 4 | 2.488 | 2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | V2 | 1 | 2 | 0.7016 | 0.2522 | 0.3497 | -1.238 | -1.2 | 0.3551 | 0.3388 | 0.2042 | 0.1581 | 0.1498 | 4.25 | 20 | 2.398 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | V5 | 1 | 3 | 0.8616 | 0.5715 | 0.6024 | 1.328 | 1.457 | 0.132 | 0.1397 | 0.3215 | 0.013 | 0.0186 | 2.25 | 36 | 2.444 | -8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | III | 2 | 0 | 0.9784 | 0.7461 | 0.7675 | 3.327 | 3.373 | 0.0837 | 0.061 | 0.6671 | 0.8555 | 0.8561 | 0.5 | 4 | 2.488 | 2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | aVF | 2 | 1 | 0.956 | 0.7569 | 0.7744 | 3.554 | 3.566 | 0.0797 | 0.0797 | 0.6459 | 0.7016 | 0.4901 | 4.25 | 36 | 2.398 | 2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | V3 | 2 | 2 | 0.9432 | 0.2329 | 0.3001 | -6.609 | -6.546 | 0.4481 | 0.4755 | 0.447 | 0.3003 | 0.5728 | 6.25 | 84 | 2.352 | -16 | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | V6 | 2 | 3 | 0.9352 | 0.4959 | 0.5581 | 1.147 | 1.196 | 0.095 | 0.0965 | 0.3218 | -0.0123 | 0.0072 | 7 | 52 | 2.336 | -6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | classical | eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115551.csv | II rhythm | 3 | 0 | 0.9988 | 0.6153 | 0.727 | 1.355 | 1.726 | 0.1211 | 0.1149 | 0.5908 | 0.7061 | 0.6938 | 0.5 | 24 | 9.95 | -8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | I | 0 | 0 | 0.9952 | 0.8835 | 0.931 | 5.547 | 7.701 | 0.0415 | 0.0276 | 0.9973 | 0.9179 | 0.9202 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | aVR | 0 | 1 | 0.992 | 0.9529 | 0.9762 | 10.352 | 11.564 | 0.0348 | 0.0258 | 0.929 | 0.9066 | 0.9054 | -0.25 | -6 | 2.5 | 2 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | V1 | 0 | 2 | 0.9976 | 0.9545 | 0.9894 | 10.403 | 16.493 | 0.0726 | 0.0691 | 0.9448 | 0.9798 | 0.9798 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | V4 | 0 | 3 | 0.9992 | 0.9413 | 0.9917 | 9.519 | 10.937 | 0.0683 | 0.0683 | 0.9001 | 0.9118 | 0.9039 | 1 | 8 | 2.476 | 4 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | II | 1 | 0 | 0.9992 | 0.9457 | 0.972 | 9.599 | 9.782 | 0.0487 | 0.0413 | 0.9246 | 0.853 | 0.844 | 0.25 | 2 | 2.494 | 2 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | aVL | 1 | 1 | 0.996 | 0.9461 | 0.955 | 9.431 | 9.868 | 0.0179 | 0.0181 | 0.8195 | 0.8721 | 0.8608 | -0.25 | -6 | 2.5 | 2 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | V2 | 1 | 2 | 0.9992 | 0.9565 | 0.9939 | 10.641 | 18.829 | 0.0877 | 0.0863 | 0.9366 | 0.9605 | 0.9605 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | V5 | 1 | 3 | 0.9976 | 0.9237 | 0.9933 | 8.285 | 12.08 | 0.0611 | 0.0627 | 0.8655 | 0.917 | 0.9118 | 0.75 | 2 | 2.482 | 8 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | III | 2 | 0 | 0.9952 | 0.9142 | 0.951 | 7.495 | 8.654 | 0.0666 | 0.0378 | 0.9439 | 0.9638 | 0.9617 | 0.25 | 0 | 2.494 | 2 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | aVF | 2 | 1 | 0.9976 | 0.9283 | 0.9769 | 8.572 | 9.406 | 0.0487 | 0.0448 | 0.8813 | 0.8624 | 0.8585 | -0.5 | -10 | 2.5 | 2 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | V3 | 2 | 2 | 0.9976 | 0.896 | 0.9719 | 7.027 | 12.083 | 0.0998 | 0.0989 | 0.8436 | 0.9383 | 0.9323 | 0.25 | 10 | 2.49 | -6 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | V6 | 2 | 3 | 0.9992 | 0.9347 | 0.9946 | 9.159 | 12.404 | 0.0391 | 0.0384 | 0.892 | 0.9136 | 0.9015 | 1 | 6 | 2.476 | 6 | ok | ok
B (ideal photos) | 38 | NORM | pair_20261013_115548 | learned | eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv | II rhythm | 3 | 0 | 0.9994 | 0.7638 | 0.9399 | 3.345 | 3.724 | 0.0923 | 0.0913 | 0.7609 | 0.8379 | 0.8201 | 0.5 | 26 | 9.95 | -6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | I | 0 | 0 | 0.9792 | 0.9221 | 0.9464 | 7.675 | 7.675 | 0.0853 | 0.0677 | 0.6963 | 0.5979 | 0.5945 | 1 | 12 | 2.476 | 0 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | aVR | 0 | 1 | 0.9624 | 0.9609 | 0.9714 | 11.101 | 11.913 | 0.0394 | 0.0372 | 0.8497 | 0.8326 | 0.8324 | 0.25 | 0 | 2.494 | 2 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | V1 | 0 | 2 | 0.9624 | 0.5973 | 0.6153 | 0.526 | 0.711 | 0.2581 | 0.2645 | 0.6206 | 0.4812 | 0.4735 | 0.25 | 10 | 2.49 | -6 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | V4 | 0 | 3 | 0.98 | 0.4661 | 0.474 | 1.312 | 1.34 | 0.1545 | 0.154 | 0.2768 | 0.0418 | 0.0398 | -1 | -18 | 2.5 | 8 | UNRELIABLE | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | II | 1 | 0 | 0.9792 | 0.9369 | 0.9797 | 9.162 | 12.439 | 0.0373 | 0.0373 | 0.9138 | 0.7425 | 0.7399 | -0.25 | -8 | 2.498 | 6 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | aVL | 1 | 1 | 0.9624 | 0.9856 | 0.9867 | 15.014 | 15.014 | 0.0274 | 0.0276 | 0.912 | 0.8062 | 0.8027 | 0.25 | 2 | 2.494 | 0 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | V2 | 1 | 2 | 0.9624 | 0.9019 | 0.9066 | 7.435 | 7.522 | 0.217 | 0.1681 | 0.756 | 0.1495 | 0.1465 | 0.5 | 8 | 2.488 | -2 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | V5 | 1 | 3 | 0.98 | 0.6005 | 0.6159 | 1.99 | 2.064 | 0.1338 | 0.1288 | 0.2885 | 0.0971 | 0.0971 | 0 | -8 | 2.492 | 8 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | III | 2 | 0 | 0.684 | 0.4214 | 0.4414 | -9.93 | -9.91 | 0.3616 | 0.3634 | 1.2558 | 0.7836 | 0.7027 | 3.75 | 56 | 2.41 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | aVF | 2 | 1 | 0.9616 | 0.9681 | 0.9681 | 12.123 | 12.123 | 0.0234 | 0.0203 | 0.9495 | 0.918 | 0.918 | 0 | 0 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | V3 | 2 | 2 | 0.9648 | 0.7579 | 0.7919 | 4.073 | 4.691 | 0.163 | 0.1506 | 0.6462 | 0.7953 | 0.7953 | 0 | 6 | 2.494 | -6 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | V6 | 2 | 3 | 0.9784 | 0.3963 | 0.4436 | -2.829 | -2.747 | 0.242 | 0.1878 | 0.5525 | 0.5423 | 0.5437 | 2.5 | 28 | 2.44 | -8 | UNRELIABLE | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_132949.csv | II rhythm | 3 | 0 | 0.9992 | 0.8695 | 0.893 | 5.853 | 5.893 | 0.0606 | 0.0536 | 0.9065 | 0.7167 | 0.7165 | 0.25 | 14 | 9.976 | -2 | ok | check | LEAD_COVERAGE/warn/III
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | I | 0 | 0 | 0.9952 | 0.9966 | 0.9966 | 21.703 | 21.703 | 0.0443 | 0.0135 | 0.9783 | 0.9479 | 0.9479 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | aVR | 0 | 1 | 0.9976 | 0.9634 | 0.9912 | 11.725 | 17.526 | 0.0391 | 0.0346 | 0.9187 | 0.9464 | 0.9464 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | V1 | 0 | 2 | 0.9976 | 0.9569 | 0.9957 | 10.722 | 20.349 | 0.09 | 0.0818 | 0.9332 | 0.9493 | 0.9492 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | V4 | 0 | 3 | 0.9976 | 0.8914 | 0.9961 | 7.059 | 10.378 | 0.086 | 0.0795 | 0.8765 | 0.9353 | 0.9202 | 1 | 6 | 2.476 | 6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | II | 1 | 0 | 0.9952 | 0.9455 | 0.9936 | 9.886 | 18.834 | 0.0353 | 0.0344 | 0.9114 | 0.887 | 0.8871 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | aVL | 1 | 1 | 0.9976 | 0.9732 | 0.9943 | 12.814 | 18.625 | 0.0355 | 0.0356 | 0.9225 | 0.8896 | 0.8895 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | V2 | 1 | 2 | 0.9976 | 0.985 | 0.9977 | 15.542 | 23.438 | 0.1371 | 0.0661 | 0.975 | 0.9382 | 0.9378 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | V5 | 1 | 3 | 0.9976 | 0.9338 | 0.9909 | 9.11 | 10.829 | 0.0808 | 0.0568 | 0.9361 | 0.9404 | 0.933 | 1 | 8 | 2.476 | 4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | III | 2 | 0 | 0.9928 | 0.9044 | 0.9927 | 7.46 | 12.544 | 0.0539 | 0.0491 | 0.8797 | 0.8928 | 0.8903 | -0.75 | -14 | 2.5 | 6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | aVF | 2 | 1 | 0.996 | 0.9862 | 0.987 | 15.368 | 15.61 | 0.0189 | 0.014 | 1.0024 | 0.9497 | 0.9499 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | V3 | 2 | 2 | 0.9992 | 0.9711 | 0.9927 | 12.911 | 18.732 | 0.0907 | 0.0544 | 0.9474 | 0.8862 | 0.8862 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | V6 | 2 | 3 | 0.996 | 0.9747 | 0.9955 | 13.162 | 13.557 | 0.0831 | 0.0298 | 0.9842 | 0.9816 | 0.9773 | 0.75 | 8 | 2.482 | 2 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_132946 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv | II rhythm | 3 | 0 | 0.9986 | 0.9289 | 0.9652 | 8.759 | 8.759 | 0.0445 | 0.0383 | 0.9137 | 0.8823 | 0.8715 | 0.25 | 14 | 9.976 | 0 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | I | 0 | 0 | 0.9776 | 0.9486 | 0.9539 | 9.885 | 9.885 | 0.0771 | 0.0525 | 0.8216 | 0.704 | 0.7016 | 0.5 | 8 | 2.488 | 0 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | aVR | 0 | 1 | 0.9616 | 0.9428 | 0.9481 | 9.505 | 9.842 | 0.0449 | 0.0447 | 0.8174 | 0.8497 | 0.8501 | 0.25 | 2 | 2.494 | 2 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | V1 | 0 | 2 | 0.96 | 0.8973 | 0.9209 | 6.859 | 7.719 | 0.1343 | 0.1276 | 0.7139 | 0.7187 | 0.7187 | 0 | 6 | 2.494 | -6 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | V4 | 0 | 3 | 0.7176 | 0.1227 | 0.1262 | -6.738 | -6.738 | 0.4118 | 0.3892 | 0.2401 | 0.4353 | 0.4355 | -0.25 | 0 | 2.5 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | II | 1 | 0 | 0.98 | 0.9598 | 0.9828 | 10.959 | 13.081 | 0.0297 | 0.0304 | 0.9276 | 0.7443 | 0.7428 | -0.25 | -6 | 2.5 | 4 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | aVL | 1 | 1 | 0.96 | 0.9864 | 0.9864 | 15.462 | 15.462 | 0.0262 | 0.0262 | 0.9286 | 0.82 | 0.82 | 0 | 0 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | V2 | 1 | 2 | 0.9592 | 0.9676 | 0.9865 | 12.144 | 15.272 | 0.1496 | 0.0978 | 0.9054 | 0.8524 | 0.8521 | 0.25 | 8 | 2.492 | -6 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | V5 | 1 | 3 | 0.7136 | 0.1478 | 0.1555 | -10.363 | -10.363 | 0.578 | 0.5342 | 0.4472 | 0.4557 | 0.4542 | 0 | 8 | 2.492 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | III | 2 | 0 | 0.9792 | 0.9107 | 0.9666 | 7.469 | 9.82 | 0.0544 | 0.049 | 0.9358 | 0.8009 | 0.7963 | -0.75 | -12 | 2.5 | 4 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | aVF | 2 | 1 | 0.96 | 0.9691 | 0.9691 | 11.927 | 11.927 | 0.0255 | 0.0208 | 0.9541 | 0.8937 | 0.8937 | 0 | 0 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | V3 | 2 | 2 | 0.8904 | 0.5367 | 0.5551 | 0.085 | 0.286 | 0.2604 | 0.2383 | 0.614 | 0.6972 | 0.6911 | 0.5 | 12 | 2.488 | -6 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | V6 | 2 | 3 | 0.98 | 0.8377 | 0.8699 | 5.405 | 5.82 | 0.0981 | 0.0728 | 0.6478 | 0.5568 | 0.5567 | 1.25 | 14 | 2.47 | -4 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv | II rhythm | 3 | 0 | 0.9992 | 0.9054 | 0.9753 | 7.55 | 8.079 | 0.0494 | 0.0441 | 0.883 | 0.7341 | 0.7282 | 0.25 | 16 | 9.976 | -4 | ok | check | LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | I | 0 | 0 | 0.9976 | 0.9927 | 0.9961 | 18.391 | 18.711 | 0.0491 | 0.0197 | 0.9796 | 0.9609 | 0.9568 | 0.25 | 2 | 2.494 | 2 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | aVR | 0 | 1 | 0.9992 | 0.9423 | 0.9884 | 9.761 | 16.425 | 0.0469 | 0.0434 | 0.919 | 0.9408 | 0.9366 | -0.25 | -8 | 2.498 | 6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | V1 | 0 | 2 | 0.9976 | 0.9459 | 0.9961 | 9.716 | 20.704 | 0.0981 | 0.0918 | 0.9281 | 0.9559 | 0.9557 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | V4 | 0 | 3 | 0.9976 | 0.8894 | 0.9932 | 6.916 | 10.117 | 0.0892 | 0.0808 | 0.877 | 0.9284 | 0.9091 | 1 | 6 | 2.476 | 6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | II | 1 | 0 | 0.9976 | 0.9734 | 0.9958 | 12.871 | 20.081 | 0.025 | 0.0244 | 0.9376 | 0.8664 | 0.8666 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | aVL | 1 | 1 | 0.9976 | 0.9828 | 0.9948 | 14.688 | 19.29 | 0.0291 | 0.0287 | 0.9362 | 0.9198 | 0.9198 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | V2 | 1 | 2 | 0.9968 | 0.983 | 0.9974 | 14.999 | 22.894 | 0.1404 | 0.0704 | 0.9683 | 0.9321 | 0.9318 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | V5 | 1 | 3 | 0.9984 | 0.9005 | 0.9879 | 7.34 | 10.693 | 0.0902 | 0.0696 | 0.9021 | 0.9299 | 0.9282 | 1 | 6 | 2.476 | 6 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | III | 2 | 0 | 0.996 | 0.9503 | 0.9926 | 10.285 | 17.892 | 0.0429 | 0.0355 | 0.9187 | 0.9094 | 0.9052 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | aVF | 2 | 1 | 0.9976 | 0.9566 | 0.9878 | 10.621 | 14.237 | 0.0275 | 0.0241 | 0.9658 | 0.939 | 0.9303 | -0.25 | -6 | 2.5 | 4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | V3 | 2 | 2 | 0.996 | 0.9466 | 0.9921 | 10.271 | 18.323 | 0.1033 | 0.0738 | 0.9219 | 0.901 | 0.9012 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | V6 | 2 | 3 | 0.9992 | 0.9676 | 0.9974 | 12.117 | 12.821 | 0.0833 | 0.0336 | 0.9725 | 0.9715 | 0.9692 | 1 | 10 | 2.476 | 2 | ok | ok
B (ideal photos) | 40 | NORM | pair_20261013_133135 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv | II rhythm | 3 | 0 | 0.9994 | 0.9045 | 0.9809 | 7.475 | 7.643 | 0.0506 | 0.0445 | 0.8926 | 0.9128 | 0.8723 | 0.25 | 14 | 9.976 | -2 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | I | 0 | 0 | 0.9792 | 0.943 | 0.9741 | 9.885 | 11.932 | 0.1149 | 0.0441 | 0.8602 | 0.7712 | 0.7671 | 0.25 | 0 | 2.494 | 4 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | aVR | 0 | 1 | 0.964 | 0.9263 | 0.93 | 7.978 | 8.131 | 0.0868 | 0.0565 | 0.731 | 0.5833 | 0.5833 | 0 | 2 | 2.498 | -2 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | V1 | 0 | 2 | 0.9648 | 0.7613 | 0.9341 | 3.75 | 8.002 | 0.0868 | 0.0851 | 0.6021 | 0.6555 | 0.6555 | 0.5 | 14 | 2.486 | -8 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | V4 | 0 | 3 | 0.7856 | 0.0934 | 0.1479 | -6.864 | -6.605 | 0.5698 | 0.5107 | 0.1886 | 0.3855 | 0.3864 | 3.5 | 32 | 2.416 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | II | 1 | 0 | 0.9816 | 0.7653 | 0.7873 | 3.565 | 3.973 | 0.112 | 0.1142 | 0.733 | 0.7567 | 0.7565 | 0 | -4 | 2.496 | 4 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | aVL | 1 | 1 | 0.9616 | 0.9559 | 0.9575 | 10.599 | 10.599 | 0.0205 | 0.0203 | 0.8843 | 0.6096 | 0.602 | -0.25 | -4 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | V2 | 1 | 2 | 0.9632 | 0.7726 | 0.8515 | 4.118 | 5.706 | 0.1748 | 0.1518 | 0.6324 | 0.3953 | 0.3947 | 0.25 | 10 | 2.49 | -8 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | V5 | 1 | 3 | 0.9584 | 0.1971 | 0.2465 | -5.54 | -5.473 | 0.5454 | 0.43 | 0.3685 | 0.6075 | 0.5938 | 1.75 | 20 | 2.458 | -4 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | III | 2 | 0 | 0.9808 | -0.1455 | -0.0117 | -7.392 | -6.945 | 0.1832 | 0.1563 | -0.2775 | 0.5787 | 0.5692 | 10 | 98 | 2.272 | -194 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | aVF | 2 | 1 | 0.9624 | 0.8809 | 0.8965 | 6.273 | 6.662 | 0.0833 | 0.0589 | 0.6364 | 0.4516 | 0.4522 | 0 | -4 | 2.496 | 4 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | V3 | 2 | 2 | 0.9648 | 0.2062 | 0.2962 | -7.384 | -7.199 | 0.4773 | 0.4897 | 0.4799 | 0.8363 | 0.8178 | 5.5 | 100 | 2.37 | -10 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | V6 | 2 | 3 | 0.9584 | 0.3173 | 0.4005 | -2.295 | -2.173 | 0.3035 | 0.2805 | 0.4075 | 0.2653 | 0.2603 | 1.75 | 14 | 2.458 | -4 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133705.csv | II rhythm | 3 | 0 | 0.9772 | 0.6025 | 0.6512 | 0.524 | 0.569 | 0.1815 | 0.1611 | 0.6685 | 0.7274 | 0.7225 | 0.25 | 14 | 9.976 | -2 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | I | 0 | 0 | 0.9976 | 0.9372 | 0.9837 | 9.53 | 14.997 | 0.1183 | 0.0459 | 0.8784 | 0.8597 | 0.8595 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | aVR | 0 | 1 | 0.9976 | 0.9271 | 0.9863 | 8.705 | 11.775 | 0.085 | 0.052 | 0.8604 | 0.8904 | 0.885 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | V1 | 0 | 2 | 0.9992 | 0.8378 | 0.9949 | 4.996 | 18.942 | 0.0763 | 0.0737 | 0.8155 | 0.9532 | 0.9519 | -0.25 | 4 | 2.496 | -8 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | V4 | 0 | 3 | 0.996 | 0.8275 | 0.9737 | 5.24 | 5.414 | 0.1643 | 0.1267 | 0.758 | 0.8392 | 0.8113 | 0.75 | 4 | 2.482 | 2 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | II | 1 | 0 | 0.9952 | 0.8844 | 0.9818 | 6.834 | 14.218 | 0.0819 | 0.0784 | 0.8252 | 0.8036 | 0.8034 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | aVL | 1 | 1 | 0.996 | 0.9546 | 0.9748 | 10.521 | 12.366 | 0.0205 | 0.0205 | 0.9195 | 0.8249 | 0.7871 | -0.25 | -6 | 2.5 | 2 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | V2 | 1 | 2 | 0.9976 | 0.9109 | 0.9983 | 7.789 | 24.431 | 0.1211 | 0.0995 | 0.8928 | 0.9159 | 0.916 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | V5 | 1 | 3 | 0.9976 | 0.7917 | 0.9558 | 4.537 | 4.592 | 0.157 | 0.1348 | 0.728 | 0.8692 | 0.8162 | 1 | 4 | 2.476 | 2 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | III | 2 | 0 | 0.992 | 0.5506 | 0.9596 | 0.881 | 9.111 | 0.1543 | 0.0603 | 0.4949 | 0.5863 | 0.5602 | -0.25 | -10 | 2.496 | 10 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | aVF | 2 | 1 | 0.9976 | 0.9152 | 0.979 | 8.145 | 11.107 | 0.0794 | 0.0475 | 0.8331 | 0.8911 | 0.8762 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | V3 | 2 | 2 | 0.9992 | 0.9213 | 0.9912 | 8.435 | 16.982 | 0.104 | 0.0793 | 0.8855 | 0.8554 | 0.8524 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | V6 | 2 | 3 | 0.996 | 0.8692 | 0.9794 | 6.394 | 6.397 | 0.1333 | 0.1032 | 0.8209 | 0.8546 | 0.8276 | 0.75 | 6 | 2.482 | 2 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133702 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv | II rhythm | 3 | 0 | 0.998 | 0.906 | 0.906 | 7.546 | 7.546 | 0.079 | 0.0718 | 0.8586 | 0.875 | 0.875 | 0 | 0 | 10 | 0 | ok | ok
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | I | 0 | 0 | 0.9544 | 0.7711 | 0.8594 | 4.028 | 4.038 | 0.1126 | 0.0865 | 0.563 | 0.4229 | 0.4228 | 1.25 | 16 | 2.47 | -2 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | aVR | 0 | 1 | 0.9192 | 0.9724 | 0.9747 | 12.557 | 12.893 | 0.0626 | 0.0333 | 0.8893 | 0.769 | 0.7691 | 0 | -2 | 2.498 | 2 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | V1 | 0 | 2 | 0.9544 | 0.4375 | 0.5524 | -2.576 | -1.845 | 0.1978 | 0.1763 | 0.5644 | 0.8733 | 0.8698 | -0.5 | 4 | 2.496 | -8 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | V4 | 0 | 3 | 0.9792 | 0.3926 | 0.4738 | 0.092 | 0.285 | 0.2324 | 0.2293 | 0.3189 | 0.0141 | 0.0422 | -4.75 | -60 | 2.5 | 6 | UNRELIABLE | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | II | 1 | 0 | 0.9568 | 0.7254 | 0.8242 | 2.02 | 2.401 | 0.1622 | 0.1364 | 0.8232 | 0.7254 | 0.7239 | 1 | 16 | 2.476 | -4 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | aVL | 1 | 1 | 0.8976 | 0.948 | 0.9516 | 9.624 | 9.624 | 0.0222 | 0.0227 | 0.847 | 0.6794 | 0.6562 | -0.25 | -4 | 2.5 | 0 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | V2 | 1 | 2 | 0.96 | 0.837 | 0.9918 | 5.299 | 17.119 | 0.1513 | 0.1325 | 0.7879 | 0.8173 | 0.8162 | -0.25 | 6 | 2.494 | -8 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | V5 | 1 | 3 | 0.7504 | 0.1752 | 0.23 | -3.261 | -3.043 | 0.3906 | 0.3307 | 0.2325 | 0.3991 | 0.3974 | 1 | 10 | 2.476 |  | UNRELIABLE | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | III | 2 | 0 | 0.9808 | 0.7811 | 0.8619 | 3.76 | 4.937 | 0.14 | 0.0433 | 0.7377 | 0.5997 | 0.5985 | 0.5 | 8 | 2.488 | -4 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | aVF | 2 | 1 | 0.948 | 0.8901 | 0.8935 | 6.892 | 6.919 | 0.0892 | 0.0549 | 0.7213 | 0.6832 | 0.6722 | -0.25 | -2 | 2.5 | -2 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | V3 | 2 | 2 | 0.9624 | 0.3639 | 0.4303 | -3.191 | -2.85 | 0.3085 | 0.3022 | 0.5062 | 0.7229 | 0.7217 | -0.75 | 2 | 2.498 | -10 | UNRELIABLE | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | V6 | 2 | 3 | 0.9784 | 0.6336 | 0.7248 | 2.496 | 2.538 | 0.176 | 0.1616 | 0.3385 | 0.2308 | 0.2311 | 1.5 | 6 | 2.464 | 2 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | classical | eacg_sheet_ptbxl57_kit_500hz_300dpi_classical_20261013_133918.csv | II rhythm | 3 | 0 | 0.999 | 0.2683 |  | -2.328 | -1.959 | 0.2469 | 0.2237 | 0.3152 | 0.6472 |  | -1.5 | 2774 | 7.226 | -8 | UNRELIABLE | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V5
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | I | 0 | 0 | 0.9984 | 0.9052 | 0.9735 | 7.764 | 8.839 | 0.1243 | 0.0563 | 0.8592 | 0.8707 | 0.8677 | 0.75 | 14 | 2.482 | -4 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | aVR | 0 | 1 | 0.9992 | 0.9405 | 0.9842 | 9.585 | 12.115 | 0.0842 | 0.047 | 0.8932 | 0.9079 | 0.9072 | -0.25 | -6 | 2.5 | 4 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | V1 | 0 | 2 | 0.9968 | 0.8811 | 0.9956 | 6.365 | 20.174 | 0.0658 | 0.063 | 0.8517 | 0.9611 | 0.9611 | 0 | 6 | 2.494 | -6 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | V4 | 0 | 3 | 0.9976 | 0.6476 | 0.9668 | 2.15 | 2.588 | 0.2087 | 0.1809 | 0.6005 | 0.8673 | 0.8205 | 1.5 | 4 | 2.464 | 4 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | II | 1 | 0 | 0.9976 | 0.9111 | 0.982 | 7.929 | 8.518 | 0.0732 | 0.0691 | 0.8599 | 0.906 | 0.878 | 0.5 | 8 | 2.488 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | aVL | 1 | 1 | 0.9992 | 0.9477 | 0.9729 | 9.892 | 11.55 | 0.0221 | 0.022 | 0.9244 | 0.8606 | 0.8058 | -0.25 | -6 | 2.5 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | V2 | 1 | 2 | 0.996 | 0.9167 | 0.9972 | 8.097 | 22.412 | 0.1189 | 0.096 | 0.8935 | 0.9056 | 0.9055 | 0 | 6 | 2.494 | -6 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | V5 | 1 | 3 | 0.9992 | 0.687 | 0.9493 | 2.768 | 2.863 | 0.185 | 0.1652 | 0.6362 | 0.8437 | 0.8259 | 1.5 | 6 | 2.464 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | III | 2 | 0 | 0.9992 | 0.853 | 0.8899 | 5.426 | 5.853 | 0.1411 | 0.0357 | 0.8277 | 0.629 | 0.608 | 0.25 | 4 | 2.494 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | aVF | 2 | 1 | 0.9976 | 0.9474 | 0.9848 | 10.182 | 12.671 | 0.0757 | 0.0376 | 0.8906 | 0.84 | 0.8231 | -0.25 | -6 | 2.5 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | V3 | 2 | 2 | 0.9976 | 0.9126 | 0.9929 | 7.971 | 17.393 | 0.1066 | 0.0836 | 0.8876 | 0.9204 | 0.9188 | -0.25 | 2 | 2.498 | -6 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | V6 | 2 | 3 | 0.9976 | 0.7407 | 0.9792 | 3.377 | 3.448 | 0.1706 | 0.146 | 0.7151 | 0.8405 | 0.8406 | 1.25 | 6 | 2.47 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 57 | NORM | pair_20261013_133915 | learned | eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv | II rhythm | 3 | 0 | 0.9988 | 0.8626 | 0.9544 | 5.911 | 5.995 | 0.0921 | 0.0866 | 0.8236 | 0.9043 | 0.8864 | 0.25 | 12 | 9.976 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | I | 0 | 0 | 0.9784 | 0.5477 | 0.6174 | 1.43 | 1.779 | 0.1018 | 0.1022 | 0.2215 | 0.1922 | 0.192 | 0 | 10 | 2.49 | -10 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | aVR | 0 | 1 | 0.96 | 0.7104 | 0.7207 | 2.635 | 2.665 | 0.0685 | 0.0668 | 0.3269 | 0.1916 | 0.1917 | 0.75 | 8 | 2.482 | 4 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | V1 | 0 | 2 | 0.9608 | 0.3903 | 0.4275 | -2.69 | -2.542 | 0.2319 | 0.2255 | 0.5296 | 0.3859 | 0.3909 | 3 | 40 | 2.428 | -16 | UNRELIABLE | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | V4 | 0 | 3 | 0.976 | 0.1478 | 0.4499 | -0.509 | 0.553 | 0.2203 | 0.2193 | 0.0773 | -0.0002 | -0.0078 | 3 | 28 | 2.428 | 14 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | II | 1 | 0 | 0.9784 | 0.5298 | 0.7072 | 0.555 | 1.72 | 0.0686 | 0.0672 | 0.5046 | 0.1291 | 0.1272 | -1.25 | -14 | 2.5 | 6 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | aVL | 1 | 1 | 0.96 | 0.4706 | 0.4729 | 0.934 | 0.94 | 0.0975 | 0.0964 | 0.1305 | 0.0693 | 0.0668 | -0.25 | -6 | 2.5 | 2 | UNRELIABLE | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | V2 | 1 | 2 | 0.9608 | 0.4611 | 0.4854 | -0.692 | -0.563 | 0.212 | 0.2163 | 0.4811 | 0.3035 | 0.2967 | 1.5 | 26 | 2.464 | -6 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | V5 | 1 | 3 | 0.976 | 0.2103 | 0.4377 | 0.128 | 0.203 | 0.1593 | 0.1595 | 0.0592 | 0.0091 | 0.0512 | 3 | 34 | 2.428 | 8 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | III | 2 | 0 | 0.9784 | 0.7946 | 0.8236 | 4.21 | 4.268 | 0.0554 | 0.054 | 0.6037 | 0.5083 | 0.5064 | 1 | 12 | 2.476 | 2 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | aVF | 2 | 1 | 0.9592 | 0.8578 | 0.8794 | 5.368 | 5.627 | 0.0506 | 0.0272 | 0.7304 | 0.4988 | 0.4962 | -0.75 | -14 | 2.5 | 4 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | V3 | 2 | 2 | 0.9616 | 0.2552 | 0.3068 | -2.611 | -2.354 | 0.2702 | 0.2724 | 0.2885 | 0.5089 | 0.5102 | -1 | 6 | 2.494 | -12 | UNRELIABLE | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | V6 | 2 | 3 | 0.5528 | -0.1609 | -0.1342 | -11.839 | -11.816 | 0.5245 | 0.5126 | -0.4885 | 0.8281 | 0.8179 | -10 | -96 | 2.5 |  | UNRELIABLE | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145329.csv | II rhythm | 3 | 0 | 0.999 | 0.7335 | 0.7756 | 3.472 | 3.472 | 0.0521 | 0.0493 | 0.594 | 0.2181 | 0.206 | 0.25 | 18 | 9.976 | 0 | ok | check | NO_PULSE/warn; LEAD_COVERAGE/warn/V6
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | I | 0 | 0 | 0.9984 | 0.9882 | 0.9937 | 16.31 | 16.524 | 0.0207 | 0.0184 | 0.961 | 0.9476 | 0.9384 | 0.25 | 4 | 2.494 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | aVR | 0 | 1 | 0.9984 | 0.9265 | 0.9893 | 8.64 | 16.285 | 0.0396 | 0.0335 | 0.8685 | 0.9292 | 0.9292 | 0 | -4 | 2.496 | 4 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | V1 | 0 | 2 | 0.9992 | 0.9598 | 0.9982 | 11.057 | 18.771 | 0.0477 | 0.0463 | 0.9438 | 0.968 | 0.9603 | -0.25 | 2 | 2.498 | -4 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | V4 | 0 | 3 | 0.996 | 0.8509 | 0.9916 | 5.508 | 8.478 | 0.1105 | 0.1097 | 0.7854 | 0.9164 | 0.9097 | 0.5 | 2 | 2.488 | 4 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | II | 1 | 0 | 0.996 | 0.9797 | 0.9877 | 13.572 | 15.273 | 0.0196 | 0.015 | 0.8946 | 0.842 | 0.8421 | 0 | -2 | 2.498 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | aVL | 1 | 1 | 0.9984 | 0.9397 | 0.9941 | 9.426 | 15.692 | 0.0368 | 0.0362 | 0.8644 | 0.9413 | 0.9309 | -0.25 | -8 | 2.498 | 6 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | V2 | 1 | 2 | 0.9992 | 0.9513 | 0.9958 | 10.214 | 18.096 | 0.0774 | 0.0616 | 0.903 | 0.9376 | 0.9324 | -0.25 | 2 | 2.498 | -4 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | V5 | 1 | 3 | 0.996 | 0.8531 | 0.992 | 5.602 | 7.727 | 0.085 | 0.0849 | 0.791 | 0.9474 | 0.9372 | 0.5 | 0 | 2.488 | 4 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | III | 2 | 0 | 0.9984 | 0.9907 | 0.9937 | 17.155 | 17.155 | 0.0283 | 0.0122 | 0.9533 | 0.876 | 0.8695 | 0.25 | 4 | 2.494 | 0 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | aVF | 2 | 1 | 0.9976 | 0.951 | 0.979 | 9.889 | 13.231 | 0.0419 | 0.0161 | 0.9478 | 0.9267 | 0.9265 | 0 | -4 | 2.496 | 4 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | V3 | 2 | 2 | 0.9992 | 0.9617 | 0.9917 | 11.242 | 15.722 | 0.0587 | 0.0553 | 0.9231 | 0.9414 | 0.9379 | -0.25 | 0 | 2.5 | -2 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | V6 | 2 | 3 | 0.996 | 0.9445 | 0.99 | 9.837 | 10.207 | 0.0431 | 0.0423 | 0.9152 | 0.9267 | 0.9222 | 0.5 | 4 | 2.488 | 2 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145325 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145325.csv | II rhythm | 3 | 0 | 0.9996 | 0.8773 | 0.9632 | 6.523 | 6.523 | 0.0427 | 0.0347 | 0.8131 | 0.8845 | 0.8753 | 0.25 | 12 | 9.976 | 0 | ok | check | NO_PULSE/warn
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | I | 0 | 0 | 0.78 | 0.3566 | 0.3747 | -4.259 | -4.194 | 0.2063 | 0.1967 | 0.558 | 0.6932 | 0.684 | 0.5 | 4 | 2.488 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | aVR | 0 | 1 | 0.9616 | 0.9407 | 0.9846 | 9.477 | 14.525 | 0.0358 | 0.0304 | 0.8638 | 0.7941 | 0.7941 | 0 | -4 | 2.496 | 4 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | V1 | 0 | 2 | 0.964 | 0.5436 | 0.6217 | -1.871 | -1.22 | 0.2086 | 0.2052 | 0.7576 | 0.7123 | 0.7117 | 0 | 12 | 2.488 | -12 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | V4 | 0 | 3 | 0.9784 | 0.0831 | 0.1549 | -6.482 | -6.452 | 0.4643 | 0.4363 | 0.1291 | -0.4833 | 0.1629 | 1.25 | 28 | 2.47 | -2 | UNRELIABLE | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | II | 1 | 0 | 0.9784 | 0.8831 | 0.9309 | 6.248 | 6.607 | 0.0364 | 0.0349 | 0.9044 | 0.7432 | 0.7397 | -0.75 | -12 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | aVL | 1 | 1 | 0.9616 | 0.8953 | 0.9104 | 6.285 | 6.464 | 0.0538 | 0.052 | 0.6166 | 0.5977 | 0.5937 | 0.5 | 4 | 2.488 | 2 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | V2 | 1 | 2 | 0.964 | 0.8532 | 0.9294 | 5.597 | 8.124 | 0.116 | 0.1049 | 0.6811 | 0.5778 | 0.5778 | 0 | 6 | 2.494 | -6 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | V5 | 1 | 3 | 0.9784 | 0.6351 | 0.873 | 2.216 | 2.545 | 0.1254 | 0.1254 | 0.4527 | 0.664 | 0.7319 | 2.25 | 28 | 2.444 | 4 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | III | 2 | 0 | 0.9792 | 0.9652 | 0.9791 | 11.632 | 13.298 | 0.0347 | 0.023 | 0.9429 | 0.822 | 0.8216 | 0.25 | 8 | 2.492 | -4 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | aVF | 2 | 1 | 0.9608 | 0.9719 | 0.9719 | 12.414 | 12.414 | 0.045 | 0.0121 | 0.953 | 0.8533 | 0.8533 | 0 | 0 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | V3 | 2 | 2 | 0.9624 | 0.2027 | 0.2426 | -6.001 | -5.834 | 0.537 | 0.4025 | 0.3733 | 0.4509 | 0.4461 | -0.75 | 4 | 2.496 | -8 | UNRELIABLE | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | V6 | 2 | 3 | 0.98 | 0.0868 | 0.16 | -4.935 | -4.891 | 0.2421 | 0.2315 | 0.1284 | 0.7222 | 0.7217 | 2.25 | 36 | 2.444 | 6 | UNRELIABLE | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | classical | eacg_sheet_ptbxl160_kit_500hz_300dpi_classical_20261013_145446.csv | II rhythm | 3 | 0 | 0.9988 | 0.8743 | 0.9564 | 6.443 | 6.661 | 0.0415 | 0.035 | 0.7953 | 0.754 | 0.7472 | 0.25 | 14 | 9.976 | -2 | ok | check | LEAD_COVERAGE/warn/I
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | I | 0 | 0 | 0.9968 | 0.9839 | 0.9954 | 15.022 | 18.163 | 0.0234 | 0.0214 | 0.9647 | 0.9586 | 0.9491 | 0.25 | 2 | 2.494 | 2 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | aVR | 0 | 1 | 0.9976 | 0.9014 | 0.9924 | 7.366 | 17.938 | 0.0445 | 0.0387 | 0.8609 | 0.9328 | 0.9328 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | V1 | 0 | 2 | 0.9984 | 0.8897 | 0.9975 | 6.68 | 13.898 | 0.0773 | 0.0767 | 0.8749 | 0.9592 | 0.9534 | -0.5 | 4 | 2.496 | -8 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | V4 | 0 | 3 | 0.9976 | 0.8392 | 0.9909 | 5.129 | 9.377 | 0.1153 | 0.1146 | 0.7931 | 0.9367 | 0.9171 | 0.5 | 10 | 2.488 | -4 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | II | 1 | 0 | 0.9952 | 0.9825 | 0.9927 | 14.443 | 17.633 | 0.0185 | 0.0136 | 0.9242 | 0.8631 | 0.8632 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | aVL | 1 | 1 | 0.9976 | 0.959 | 0.9956 | 11.064 | 19.604 | 0.0306 | 0.03 | 0.9078 | 0.9527 | 0.9527 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | V2 | 1 | 2 | 0.9984 | 0.9448 | 0.9966 | 9.685 | 18.673 | 0.0804 | 0.0655 | 0.9026 | 0.9576 | 0.9414 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | V5 | 1 | 3 | 0.9976 | 0.8377 | 0.9922 | 5.104 | 7.17 | 0.09 | 0.09 | 0.7998 | 0.9397 | 0.9333 | 0.5 | 0 | 2.488 | 4 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | III | 2 | 0 | 0.9976 | 0.9874 | 0.9951 | 16.011 | 18.614 | 0.0288 | 0.0139 | 0.9628 | 0.9188 | 0.9089 | 0.25 | 6 | 2.494 | -2 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | aVF | 2 | 1 | 0.996 | 0.9725 | 0.9811 | 12.26 | 13.76 | 0.04 | 0.0123 | 0.988 | 0.9443 | 0.9443 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | V3 | 2 | 2 | 0.9976 | 0.9473 | 0.9921 | 9.878 | 17.702 | 0.0681 | 0.0647 | 0.9069 | 0.9515 | 0.9515 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | V6 | 2 | 3 | 0.9984 | 0.8796 | 0.9882 | 6.454 | 8.415 | 0.0631 | 0.0624 | 0.8635 | 0.9572 | 0.9536 | 0.75 | 2 | 2.482 | 6 | ok | ok
B (ideal photos) | 160 | STTC | pair_20261013_145443 | learned | eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv | II rhythm | 3 | 0 | 0.9994 | 0.8745 | 0.9653 | 6.385 | 6.385 | 0.0442 | 0.0353 | 0.8234 | 0.909 | 0.8931 | 0.25 | 12 | 9.976 | 0 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | I | 0 | 0 | 0.98 | 0.885 | 0.9215 | 6.094 | 6.475 | 0.0428 | 0.0343 | 0.933 | 0.8261 | 0.8269 | -0.75 | -8 | 2.5 | -2 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | aVR | 0 | 1 | 0.9624 | 0.6296 | 0.6485 | 1.486 | 1.69 | 0.1548 | 0.1367 | 0.6488 | 0.7584 | 0.7423 | -0.25 | -8 | 2.498 | 4 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | V1 | 0 | 2 | 0.9616 | 0.4072 | 0.4279 | -5.426 | -5.354 | 0.3354 | 0.3646 | 0.6913 | 0.8454 | 0.8395 | -0.5 | 0 | 2.5 | -6 | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | V4 | 0 | 3 | 0.8344 | 0.3682 | 0.4394 | -0.806 | -0.777 | 0.2553 | 0.2336 | 0.3748 | 0.2367 | 0.2055 | 5 | 58 | 2.38 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | II | 1 | 0 | 0.82 | 0.2776 | 0.3502 | -2.705 | -2.567 | 0.4289 | 0.3254 | 0.3619 | 0.4114 | 0.4083 | 3.25 | 52 | 2.422 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | aVL | 1 | 1 | 0.9624 | 0.4258 | 0.4668 | -2.803 | -2.733 | 0.1695 | 0.1577 | 0.6752 | 0.8919 | 0.8928 | 1.5 | 20 | 2.464 | 2 | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | V2 | 1 | 2 | 0.9616 | 0.8726 | 0.874 | 6.53 | 6.53 | 0.2116 | 0.1265 | 0.7286 | 0.1545 | 0.1524 | 0.25 | 4 | 2.494 | 0 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | V5 | 1 | 3 | 0.7224 | 0.079 | 0.111 | -7.215 | -7.214 | 0.7035 | 0.5831 | 0.1632 | 0.4574 | 0.4511 | 2.5 | 24 | 2.44 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | III | 2 | 0 | 0.98 | 0.6948 | 0.7043 | 2.433 | 2.448 | 0.1698 | 0.1582 | 0.6806 | 0.6389 | 0.6389 | 1 | 16 | 2.476 | -2 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | aVF | 2 | 1 | 0.916 | 0.4649 | 0.4839 | 1.024 | 1.055 | 0.2452 | 0.229 | 0.3618 | 0.1327 | 0.1793 | 1.25 | 8 | 2.47 | 4 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | V3 | 2 | 2 | 0.8384 | 0.4317 | 0.4766 | 0.085 | 0.385 | 0.2386 | 0.2405 | 0.3956 | 0.5405 | 0.8473 | 0.5 | 16 | 2.484 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | V6 | 2 | 3 | 0.7264 | 0.1249 | 0.1583 | -4.274 | -4.273 | 0.4206 | 0.3788 | 0.1676 | 0.2561 | 0.2582 | 3.75 | 48 | 2.41 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_144830.csv | II rhythm | 3 | 0 | 0.959 | 0.1067 |  | -6.708 | -6.692 | 0.7541 | 0.5431 | 0.1839 | 0.1063 |  | 1.5 | -5434 | 4.418 | -6 | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | I | 0 | 0 | 0.996 | 0.9802 | 0.9802 | 14.024 | 14.024 | 0.0296 | 0.0138 | 0.9543 | 0.9176 | 0.9176 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | aVR | 0 | 1 | 0.9984 | 0.9611 | 0.9886 | 11.585 | 16.65 | 0.0623 | 0.0427 | 0.9217 | 0.9462 | 0.9461 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | V1 | 0 | 2 | 0.9976 | 0.9665 | 0.9966 | 12.017 | 21.782 | 0.0853 | 0.0489 | 0.949 | 0.9646 | 0.9653 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | V4 | 0 | 3 | 0.9984 | 0.8498 | 0.9919 | 5.663 | 7.175 | 0.1264 | 0.1109 | 0.8307 | 0.9609 | 0.9456 | 1 | 8 | 2.476 | 4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | II | 1 | 0 | 0.996 | 0.9957 | 0.9974 | 20.961 | 20.961 | 0.0661 | 0.0213 | 0.9855 | 0.9167 | 0.9116 | 0.25 | 2 | 2.494 | 0 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | aVL | 1 | 1 | 0.9984 | 0.956 | 0.986 | 11.265 | 15.96 | 0.0589 | 0.0312 | 0.9292 | 0.9395 | 0.9375 | 0.25 | -2 | 2.492 | 4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | V2 | 1 | 2 | 0.9984 | 0.9805 | 0.9954 | 14.54 | 20.424 | 0.1551 | 0.0503 | 0.9499 | 0.8739 | 0.8737 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | V5 | 1 | 3 | 0.9976 | 0.8802 | 0.9947 | 6.686 | 9.128 | 0.1418 | 0.1177 | 0.8327 | 0.9103 | 0.9092 | 1 | 6 | 2.476 | 6 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | III | 2 | 0 | 0.996 | 0.9974 | 0.9974 | 23.128 | 23.128 | 0.0348 | 0.0146 | 0.9899 | 0.9467 | 0.9467 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | aVF | 2 | 1 | 0.9984 | 0.9701 | 0.991 | 12.754 | 17.726 | 0.1022 | 0.0593 | 0.9366 | 0.9202 | 0.9201 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | V3 | 2 | 2 | 0.9992 | 0.943 | 0.994 | 10.096 | 19.534 | 0.1462 | 0.076 | 0.9137 | 0.9544 | 0.9538 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | V6 | 2 | 3 | 0.9968 | 0.9613 | 0.9977 | 11.421 | 11.683 | 0.0897 | 0.0622 | 0.9144 | 0.933 | 0.9165 | 0.75 | 8 | 2.482 | 2 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_144827 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv | II rhythm | 3 | 0 | 0.999 | 0.9365 | 0.975 | 9.495 | 9.568 | 0.1158 | 0.0841 | 0.9153 | 0.9349 | 0.9339 | 0.25 | 10 | 9.976 | 2 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | I | 0 | 0 | 0.9808 | 0.7458 | 0.7746 | 1.852 | 2.016 | 0.0646 | 0.056 | 0.8859 | 0.838 | 0.8357 | -0.75 | -8 | 2.5 | -2 | ok | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | aVR | 0 | 1 | 0.9616 | 0.2022 | 0.2146 | -6.844 | -6.813 | 0.4607 | 0.3566 | 0.4177 | 0.3945 | 0.3906 | 0.75 | 18 | 2.482 | -4 | UNRELIABLE | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | V1 | 0 | 2 | 0.8368 | 0.4396 | 0.5001 | -0.78 | -0.377 | 0.2051 | 0.2135 | 0.4552 | 0.8104 | 0.8053 | 0.5 | 14 | 2.486 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | V4 | 0 | 3 | 0.832 | 0.4722 | 0.5257 | -0.168 | -0.051 | 0.2435 | 0.217 | 0.4879 | 0.4069 | 0.4087 | 9.25 | 64 | 2.288 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | II | 1 | 0 | 0.98 | 0.9242 | 0.9364 | 8.521 | 8.783 | 0.1079 | 0.0894 | 0.7804 | 0.6318 | 0.6317 | 0.5 | 6 | 2.488 | -2 | ok | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | aVL | 1 | 1 | 0.9632 | 0.4017 | 0.4689 | -2.861 | -2.771 | 0.1747 | 0.1587 | 0.6283 | 0.898 | 0.8994 | 2.25 | 26 | 2.444 | 2 | UNRELIABLE | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | V2 | 1 | 2 | 0.964 | 0.878 | 0.8933 | 6.807 | 6.856 | 0.2021 | 0.1226 | 0.7543 | 0.3161 | 0.3123 | 1 | 12 | 2.476 | -2 | ok | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | V5 | 1 | 3 | 0.7208 | -0.0042 | 0.0161 | -9.228 | -9.207 | 0.794 | 0.7352 | -0.0105 | 0.3712 | 0.3765 | 4 | 44 | 2.404 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | III | 2 | 0 | 0.9808 | 0.9591 | 0.9817 | 11.334 | 14.8 | 0.0649 | 0.0568 | 0.9431 | 0.8853 | 0.8853 | 0 | 6 | 2.494 | -6 | ok | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | aVF | 2 | 1 | 0.8672 | 0.4112 | 0.4138 | -2.053 | -2.053 | 0.4285 | 0.3263 | 0.5058 | 0.6376 | 0.6332 | 0.25 | 2 | 2.494 | 0 | ok | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | V3 | 2 | 2 | 0.9624 | 0.8603 | 0.8901 | 6.367 | 6.991 | 0.1752 | 0.1167 | 0.7805 | 0.9121 | 0.9619 | 1.5 | 18 | 2.464 | -4 | ok | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | V6 | 2 | 3 | 0.9792 | 0.2659 | 0.2763 | -0.261 | -0.261 | 0.2661 | 0.2387 | 0.1824 | 0.157 | 0.1568 | 1.5 | 22 | 2.464 | 0 | UNRELIABLE | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | classical | eacg_sheet_ptbxl167_kit_500hz_300dpi_classical_20261013_145132.csv | II rhythm | 3 | 0 | 0.999 | 0.3175 |  | -4.512 | -4.488 | 0.4641 | 0.4217 | 0.532 | 0.8585 |  | 4.5 | 6718 | 3.282 | -4 | UNRELIABLE | check | LEAD_COVERAGE/warn/V5
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | I | 0 | 0 | 0.9992 | 0.9873 | 0.9873 | 15.968 | 15.968 | 0.028 | 0.011 | 0.9727 | 0.9536 | 0.9536 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | aVR | 0 | 1 | 0.9968 | 0.9762 | 0.9909 | 13.706 | 15.626 | 0.0582 | 0.0335 | 0.9468 | 0.945 | 0.9424 | -0.25 | -6 | 2.5 | 2 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | V1 | 0 | 2 | 0.9976 | 0.9503 | 0.9963 | 10.301 | 21.305 | 0.0907 | 0.0596 | 0.9345 | 0.9727 | 0.9735 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | V4 | 0 | 3 | 0.9976 | 0.9322 | 0.9929 | 9.16 | 10.61 | 0.096 | 0.0741 | 0.8867 | 0.9201 | 0.9028 | 0.75 | 6 | 2.482 | 2 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | II | 1 | 0 | 0.996 | 0.9968 | 0.9971 | 22.306 | 22.306 | 0.0666 | 0.0183 | 0.9838 | 0.9479 | 0.9355 | 0.25 | 4 | 2.494 | 0 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | aVL | 1 | 1 | 0.9984 | 0.9673 | 0.9896 | 12.553 | 17.424 | 0.0553 | 0.0269 | 0.9424 | 0.9515 | 0.9516 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | V2 | 1 | 2 | 0.9992 | 0.9647 | 0.9977 | 12.016 | 21.053 | 0.1611 | 0.0673 | 0.9419 | 0.9194 | 0.8927 | 0.25 | 8 | 2.492 | -4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | V5 | 1 | 3 | 0.996 | 0.8832 | 0.996 | 6.814 | 10.726 | 0.14 | 0.116 | 0.8269 | 0.9096 | 0.8955 | 0.75 | 4 | 2.482 | 6 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | III | 2 | 0 | 0.9992 | 0.9932 | 0.9981 | 19.061 | 24.558 | 0.0397 | 0.0233 | 0.9869 | 0.9516 | 0.9516 | 0 | 2 | 2.498 | -2 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | aVF | 2 | 1 | 0.996 | 0.9697 | 0.9911 | 12.703 | 17.887 | 0.1025 | 0.0597 | 0.9464 | 0.9487 | 0.9486 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | V3 | 2 | 2 | 0.9976 | 0.9642 | 0.9931 | 12.095 | 18.986 | 0.1395 | 0.0603 | 0.9349 | 0.9438 | 0.9432 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | V6 | 2 | 3 | 0.9976 | 0.9463 | 0.9983 | 10.034 | 11.744 | 0.0976 | 0.073 | 0.9024 | 0.9187 | 0.9134 | 0.75 | 6 | 2.482 | 4 | ok | ok
B (ideal photos) | 167 | STTC | pair_20261013_145129 | learned | eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv | II rhythm | 3 | 0 | 0.9996 | 0.876 | 0.9834 | 6.566 | 6.685 | 0.1409 | 0.1178 | 0.8631 | 0.9497 | 0.9408 | 0.25 | 14 | 9.976 | -2 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | I | 0 | 0 | 0.9784 | 0.271 | 0.2774 | -0.475 | -0.475 | 0.2389 | 0.2142 | 0.2124 | 0.2041 | 0.1995 | -1 | -10 | 2.5 | 0 | UNRELIABLE | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | aVR | 0 | 1 | 0.9608 | 0.6583 | 0.677 | 2.608 | 2.764 | 0.0764 | 0.0763 | 0.3674 | 0.1536 | 0.1536 | 0.25 | -2 | 2.492 | 4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | V1 | 0 | 2 | 0.9616 | 0.5382 | 0.574 | 1.223 | 1.481 | 0.1402 | 0.1389 | 0.4354 | 0.3669 | 0.3667 | 1.25 | 20 | 2.47 | -6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | V4 | 0 | 3 | 0.9784 | 0.1644 | 0.2116 | -3.063 | -2.592 | 0.2417 | 0.251 | 0.1906 | -0.0054 | -0.0055 | -1.5 | -58 | 2.48 | 34 | UNRELIABLE | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | II | 1 | 0 | 0.9776 | 0.555 | 0.6431 | 1.469 | 1.918 | 0.0484 | 0.0486 | 0.3879 | 0.0097 | 0.0092 | -0.25 | -6 | 2.5 | 4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | aVL | 1 | 1 | 0.9616 | 0.2181 | 0.2341 | -4.321 | -4.305 | 0.3401 | 0.2854 | 0.3382 | 0.297 | 0.2957 | 2.25 | 26 | 2.444 | 2 | UNRELIABLE | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | V2 | 1 | 2 | 0.8256 | 0.0654 | 0.1761 | -5.696 | -5.448 | 0.4147 | 0.4357 | 0.077 | -0.0083 | 0.2971 | 0.75 | 100 | 2.4 |  | UNRELIABLE | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | V5 | 1 | 3 | 0.9328 | 0.5219 | 0.748 | 0.807 | 0.933 | 0.1187 | 0.1186 | 0.1027 | 0.1348 | 0.1345 | 1 | 12 | 2.476 | -4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | III | 2 | 0 | 0.9784 | 0.1549 | 0.1653 | -5.824 | -5.822 | 0.3205 | 0.3624 | 0.2071 | 0 | 0 | -2.25 | -20 | 2.5 | 2 | UNRELIABLE | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | aVF | 2 | 1 | 0.96 | 0.9792 | 0.9855 | 13.414 | 14.445 | 0.0503 | 0.018 | 0.9413 | 0.8747 | 0.8734 | -0.25 | -6 | 2.5 | 2 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | V3 | 2 | 2 | 0.9616 | 0.3168 | 0.348 | -0.802 | -0.707 | 0.2472 | 0.2536 | 0.2803 | 0.358 | 0.3679 | 7.5 | 76 | 2.326 | -6 | UNRELIABLE | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | V6 | 2 | 3 | 0.9784 | 0.7501 | 0.7569 | 2.809 | 2.809 | 0.0631 | 0.0642 | 0.3461 | 0.1846 | 0.1846 | 0.5 | 6 | 2.488 | 0 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134150.csv | II rhythm | 3 | 0 | 0.9992 | 0.6428 | 0.6724 | 1.957 | 2.19 | 0.0463 | 0.0463 | 0.5682 | 0.2245 | 0.2227 | 0.25 | 16 | 9.976 | -4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | I | 0 | 0 | 0.9976 | 0.9909 | 0.9953 | 17.257 | 19.572 | 0.0687 | 0.0278 | 0.937 | 0.9251 | 0.9251 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | aVR | 0 | 1 | 0.9984 | 0.9295 | 0.9942 | 8.859 | 16.863 | 0.0396 | 0.0372 | 0.8888 | 0.9345 | 0.9332 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | V1 | 0 | 2 | 0.9992 | 0.9529 | 0.995 | 10.521 | 17.537 | 0.0528 | 0.0476 | 0.9223 | 0.9606 | 0.9535 | -0.25 | 2 | 2.498 | -6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | V4 | 0 | 3 | 0.9976 | 0.8742 | 0.9871 | 6.35 | 9.467 | 0.1034 | 0.0849 | 0.8355 | 0.9053 | 0.8984 | 0.5 | 2 | 2.488 | 4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | II | 1 | 0 | 0.9968 | 0.9366 | 0.9771 | 9.062 | 12.777 | 0.0203 | 0.0203 | 0.8318 | 0.8694 | 0.8694 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | aVL | 1 | 1 | 0.9984 | 0.9546 | 0.9974 | 10.996 | 21.063 | 0.0837 | 0.0489 | 0.9212 | 0.9377 | 0.935 | -0.25 | -8 | 2.498 | 6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | V2 | 1 | 2 | 0.9992 | 0.9696 | 0.9953 | 12.296 | 19.047 | 0.0633 | 0.0549 | 0.9487 | 0.9433 | 0.9416 | -0.25 | 2 | 2.498 | -6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | V5 | 1 | 3 | 0.9976 | 0.885 | 0.9853 | 6.82 | 8.13 | 0.0755 | 0.0593 | 0.7918 | 0.9068 | 0.9025 | 0.5 | 2 | 2.488 | 4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | III | 2 | 0 | 0.9944 | 0.9621 | 0.9942 | 11.659 | 15.105 | 0.0836 | 0.0484 | 0.9512 | 0.9706 | 0.9691 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | aVF | 2 | 1 | 0.9976 | 0.9782 | 0.9939 | 14.032 | 19.277 | 0.0515 | 0.0168 | 0.9681 | 0.9412 | 0.9389 | -0.25 | -6 | 2.5 | 4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | V3 | 2 | 2 | 0.9984 | 0.9455 | 0.9881 | 10.081 | 16.504 | 0.0891 | 0.0724 | 0.9166 | 0.9322 | 0.9321 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | V6 | 2 | 3 | 0.9984 | 0.9747 | 0.9922 | 12.911 | 15.79 | 0.0219 | 0.0201 | 0.9008 | 0.9014 | 0.8801 | 0.25 | 0 | 2.494 | 2 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134146 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv | II rhythm | 3 | 0 | 0.9984 | 0.8542 | 0.8542 | 5.611 | 5.611 | 0.0308 | 0.0304 | 0.809 | 0.9328 | 0.9328 | 0 | 0 | 10 | 0 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | I | 0 | 0 | 0.9792 | 0.3707 | 0.3853 | 0.827 | 0.883 | 0.1857 | 0.1844 | 0.1582 | 0.0195 | 0.0269 | 0.5 | -2 | 2.486 | 8 | UNRELIABLE | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | aVR | 0 | 1 | 0.9632 | 0.7565 | 0.7704 | 3.603 | 3.65 | 0.0688 | 0.0681 | 0.4418 | 0.2868 | 0.2858 | -0.75 | -8 | 2.5 | -2 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | V1 | 0 | 2 | 0.964 | 0.7949 | 0.9033 | 4.35 | 6.347 | 0.0972 | 0.0969 | 0.5533 | 0.5554 | 0.555 | 0.25 | 14 | 2.486 | -10 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | V4 | 0 | 3 | 0.9816 | 0.9692 | 0.9692 | 11.782 | 11.782 | 0.0784 | 0.0454 | 0.8535 | 0.749 | 0.749 | 0 | 0 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | II | 1 | 0 | 0.98 | 0.427 | 0.6833 | -0.496 | 1.675 | 0.0609 | 0.0609 | 0.4208 | 0.0318 | 0.0318 | -0.5 | -14 | 2.498 | 8 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | aVL | 1 | 1 | 0.9648 | 0.5248 | 0.5488 | 0.942 | 0.982 | 0.1827 | 0.1557 | 0.4913 | 0.3927 | 0.3888 | 1.5 | 16 | 2.464 | 4 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | V2 | 1 | 2 | 0.7552 | 0.2955 | 0.3059 | -2.019 | -1.964 | 0.3499 | 0.2854 | 0.333 | 0.5545 | 0.5553 | 0.25 | 8 | 2.492 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | V5 | 1 | 3 | 0.8904 | 0.4708 | 0.482 | -3.4 | -3.398 | 0.2134 | 0.1925 | 0.7609 | 0.7677 | 0.7639 | 0.25 | 4 | 2.494 | 0 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | III | 2 | 0 | 0.9816 | 0.6716 | 0.7194 | 1.852 | 2.335 | 0.1696 | 0.1498 | 0.7246 | 0.6804 | 0.6769 | -0.75 | -16 | 2.5 | 6 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | aVF | 2 | 1 | 0.9632 | 0.9864 | 0.9883 | 15.489 | 16.072 | 0.0519 | 0.0142 | 0.9569 | 0.8879 | 0.8879 | 0 | -2 | 2.498 | 2 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | V3 | 2 | 2 | 0.9592 | 0.6396 | 0.6981 | 1.533 | 2.293 | 0.1946 | 0.1938 | 0.6696 | 0.8071 | 0.8053 | -0.25 | 4 | 2.496 | -6 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | V6 | 2 | 3 | 0.98 | 0.8214 | 0.8881 | 4.447 | 6.417 | 0.0553 | 0.0532 | 0.8304 | 0.7823 | 0.7828 | 0 | 6 | 2.494 | -6 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261013_134240.csv | II rhythm | 3 | 0 | 0.9986 | 0.5819 |  | 0.048 | 0.447 | 0.0582 | 0.0577 | 0.6783 | 0.7815 |  | -3 | -6598 | 3.712 | -4 | ok | check | LEAD_COVERAGE/warn/V2
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | I | 0 | 0 | 0.9968 | 0.9289 | 0.9935 | 8.946 | 14.56 | 0.0961 | 0.0724 | 0.8604 | 0.9048 | 0.9023 | -0.25 | -8 | 2.498 | 6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | aVR | 0 | 1 | 0.9976 | 0.8987 | 0.992 | 7.316 | 16.181 | 0.0461 | 0.0444 | 0.8591 | 0.9364 | 0.9345 | -0.5 | -12 | 2.5 | 6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | V1 | 0 | 2 | 0.9976 | 0.9525 | 0.9937 | 10.484 | 19.072 | 0.0542 | 0.0478 | 0.9277 | 0.9555 | 0.9554 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | V4 | 0 | 3 | 0.9992 | 0.9839 | 0.9865 | 15.046 | 15.75 | 0.0681 | 0.0312 | 0.9395 | 0.9348 | 0.9349 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | II | 1 | 0 | 0.9944 | 0.7751 | 0.9818 | 3.933 | 11.208 | 0.0366 | 0.0366 | 0.6712 | 0.9081 | 0.9008 | -0.25 | -10 | 2.496 | 6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | aVL | 1 | 1 | 0.9984 | 0.9126 | 0.9978 | 8.167 | 20.994 | 0.0956 | 0.0678 | 0.8843 | 0.9398 | 0.9374 | -0.25 | -10 | 2.496 | 8 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | V2 | 1 | 2 | 0.9976 | 0.9762 | 0.9948 | 13.354 | 19.214 | 0.0593 | 0.0486 | 0.9579 | 0.9486 | 0.9474 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | V5 | 1 | 3 | 0.9992 | 0.9774 | 0.9882 | 13.646 | 16.286 | 0.0557 | 0.027 | 0.9393 | 0.9373 | 0.9249 | 0.25 | 2 | 2.494 | 2 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | III | 2 | 0 | 0.9928 | 0.9038 | 0.9925 | 7.629 | 9.849 | 0.103 | 0.077 | 0.8913 | 0.9601 | 0.9599 | -0.75 | -18 | 2.5 | 6 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | aVF | 2 | 1 | 0.9968 | 0.9777 | 0.9944 | 13.88 | 19.705 | 0.0506 | 0.0171 | 0.9657 | 0.9649 | 0.9647 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | V3 | 2 | 2 | 0.9984 | 0.9406 | 0.9898 | 9.714 | 14.291 | 0.0918 | 0.0756 | 0.9117 | 0.958 | 0.9555 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | V6 | 2 | 3 | 0.9936 | 0.9621 | 0.9937 | 11.423 | 17.196 | 0.0261 | 0.0238 | 0.9238 | 0.9429 | 0.9285 | -0.25 | 0 | 2.5 | -2 | ok | ok
B (ideal photos) | 1737 | MI | pair_20261013_134236 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv | II rhythm | 3 | 0 | 0.9958 | 0.8075 |  | 4.425 | 4.938 | 0.0351 | 0.0349 | 0.7611 | 0.9339 |  | -3 | -6598 | 3.712 | -2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | I | 0 | 0 | 0.9808 | 0.9246 | 0.9284 | 8.044 | 8.044 | 0.0511 | 0.0452 | 0.7425 | 0.6749 | 0.6727 | 0.25 | 2 | 2.494 | 0 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | aVR | 0 | 1 | 0.9632 | 0.9122 | 0.9882 | 7.931 | 16.018 | 0.0502 | 0.0479 | 0.8601 | 0.8916 | 0.8916 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | V1 | 0 | 2 | 0.964 | 0.9596 | 0.9925 | 11.003 | 18.088 | 0.0637 | 0.0613 | 0.931 | 0.8895 | 0.8895 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | V4 | 0 | 3 | 0.98 | 0.9864 | 0.9897 | 15.958 | 15.958 | 0.0606 | 0.0309 | 0.9519 | 0.9614 | 0.9619 | 0.25 | 4 | 2.494 | 0 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | II | 1 | 0 | 0.98 | 0.9699 | 0.978 | 12.456 | 12.54 | 0.0641 | 0.0321 | 0.9434 | 0.8258 | 0.8254 | -0.5 | -6 | 2.5 | -2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | aVL | 1 | 1 | 0.964 | 0.917 | 0.9838 | 7.929 | 13.762 | 0.0338 | 0.0339 | 0.8181 | 0.8427 | 0.8427 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | V2 | 1 | 2 | 0.9648 | 0.9638 | 0.992 | 11.517 | 17.853 | 0.0934 | 0.0764 | 0.9319 | 0.9198 | 0.9187 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | V5 | 1 | 3 | 0.98 | 0.9549 | 0.9903 | 10.615 | 12.596 | 0.0449 | 0.0449 | 0.8797 | 0.8352 | 0.8338 | 0.5 | 6 | 2.488 | 2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | III | 2 | 0 | 0.9808 | 0.7981 | 0.9585 | 3.665 | 5.825 | 0.0729 | 0.062 | 0.8635 | 0.8326 | 0.8311 | -1 | -16 | 2.5 | 6 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | aVF | 2 | 1 | 0.9632 | 0.971 | 0.9841 | 12.625 | 15.076 | 0.0253 | 0.0231 | 0.9225 | 0.8184 | 0.8184 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | V3 | 2 | 2 | 0.964 | 0.9805 | 0.9934 | 14.237 | 18.834 | 0.1003 | 0.0434 | 0.9588 | 0.8797 | 0.8797 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | V6 | 2 | 3 | 0.98 | 0.5964 | 0.6218 | 0.929 | 1.194 | 0.1721 | 0.1358 | 0.6136 | 0.5858 | 0.5838 | 0.25 | 8 | 2.492 | -4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134608.csv | II rhythm | 3 | 0 | 0.9968 | 0.7392 |  | 2.778 | 3.108 | 0.12 | 0.1048 | 0.7802 | 0.8075 |  | -3.5 | 6592 | 3.408 | -4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | I | 0 | 0 | 0.9992 | 0.9878 | 0.9935 | 16.074 | 18.539 | 0.0313 | 0.0179 | 0.9461 | 0.9444 | 0.9444 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | aVR | 0 | 1 | 0.9984 | 0.8643 | 0.9941 | 5.95 | 19.31 | 0.0624 | 0.0602 | 0.8509 | 0.9733 | 0.9732 | 0 | -8 | 2.492 | 8 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | V1 | 0 | 2 | 0.9992 | 0.974 | 0.9928 | 12.871 | 18.412 | 0.0525 | 0.0494 | 0.9588 | 0.9719 | 0.9718 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | V4 | 0 | 3 | 0.9984 | 0.988 | 0.9984 | 16.564 | 19.496 | 0.0571 | 0.0288 | 0.9751 | 0.974 | 0.9746 | 0.25 | 2 | 2.494 | 2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | II | 1 | 0 | 0.9984 | 0.9978 | 0.9978 | 23.133 | 23.133 | 0.0563 | 0.0094 | 0.97 | 0.9421 | 0.9421 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | aVL | 1 | 1 | 0.9992 | 0.8503 | 0.9917 | 5.372 | 17.619 | 0.0456 | 0.0455 | 0.821 | 0.9543 | 0.9543 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | V2 | 1 | 2 | 1 | 0.9765 | 0.9935 | 13.372 | 18.828 | 0.082 | 0.0617 | 0.9568 | 0.9639 | 0.9629 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | V5 | 1 | 3 | 0.9984 | 0.9176 | 0.9961 | 7.993 | 13.185 | 0.0607 | 0.0607 | 0.9076 | 0.9573 | 0.952 | 0.5 | 4 | 2.488 | 4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | III | 2 | 0 | 0.9944 | 0.8892 | 0.994 | 6.818 | 10.192 | 0.0557 | 0.0431 | 0.8598 | 0.9591 | 0.9516 | -0.75 | -14 | 2.5 | 4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | aVF | 2 | 1 | 0.9984 | 0.9469 | 0.9929 | 10.058 | 18.454 | 0.0332 | 0.0311 | 0.929 | 0.9412 | 0.9314 | 0.25 | -2 | 2.492 | 4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | V3 | 2 | 2 | 0.9992 | 0.9907 | 0.9955 | 17.392 | 20.478 | 0.0933 | 0.0302 | 0.9837 | 0.9714 | 0.9713 | 0 | 2 | 2.498 | -2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | V6 | 2 | 3 | 0.9984 | 0.9927 | 0.9978 | 18.475 | 23.137 | 0.0347 | 0.018 | 0.9637 | 0.9399 | 0.9399 | 0 | 2 | 2.498 | -2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134604 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv | II rhythm | 3 | 0 | 0.9948 | 0.913 | 0.9447 | 7.942 | 9.898 | 0.0685 | 0.0579 | 0.9022 | 0.9339 | 0.9339 | 0 | 4 | 9.996 | -4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | I | 0 | 0 | 0.7184 | 0.2357 | 0.2604 | -5.79 | -5.719 | 0.2395 | 0.2223 | 0.3984 | 0.5437 | 0.5423 | 1.5 | 16 | 2.464 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | aVR | 0 | 1 | 0.9616 | 0.7524 | 0.8108 | 2.95 | 3.036 | 0.084 | 0.085 | 0.3452 | 0.3521 | 0.3453 | 1 | 4 | 2.476 | 4 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | V1 | 0 | 2 | 0.9624 | 0.9564 | 0.9926 | 10.678 | 18.059 | 0.0652 | 0.0636 | 0.923 | 0.8697 | 0.8695 | 0 | 6 | 2.494 | -6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | V4 | 0 | 3 | 0.9792 | 0.9439 | 0.9875 | 9.934 | 12.444 | 0.0853 | 0.0618 | 0.911 | 0.9587 | 0.9561 | 0.75 | 6 | 2.482 | 4 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | II | 1 | 0 | 0.72 | 0.3657 | 0.3821 | -9.205 | -9.19 | 0.434 | 0.3882 | 1.0054 | 0.7861 | 0.7853 | 1.5 | 16 | 2.464 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | aVL | 1 | 1 | 0.9568 | 0.8393 | 0.9078 | 5.252 | 6.906 | 0.0462 | 0.0462 | 0.6676 | 0.6626 | 0.6575 | 0.5 | 2 | 2.488 | 4 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | V2 | 1 | 2 | 0.9624 | 0.9488 | 0.9699 | 9.954 | 12.014 | 0.1084 | 0.0914 | 0.8601 | 0.9139 | 0.9129 | 0 | 6 | 2.494 | -6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | V5 | 1 | 3 | 0.9792 | 0.825 | 0.9885 | 4.977 | 8.102 | 0.0861 | 0.086 | 0.7508 | 0.8163 | 0.8132 | 1 | 8 | 2.476 | 6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | III | 2 | 0 | 0.9792 | 0.7348 | 0.8252 | 3.402 | 3.581 | 0.0719 | 0.0639 | 0.6222 | 0.582 | 0.5657 | -1 | -14 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | aVF | 2 | 1 | 0.9464 | 0.8894 | 0.9164 | 7.035 | 7.419 | 0.047 | 0.044 | 0.791 | 0.6965 | 0.6935 | 0.5 | 2 | 2.488 | 2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | V3 | 2 | 2 | 0.9624 | 0.9207 | 0.9312 | 7.524 | 8.071 | 0.1163 | 0.0941 | 0.986 | 0.8508 | 0.8503 | 0.25 | 6 | 2.494 | -4 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | V6 | 2 | 3 | 0.9792 | 0.9774 | 0.9934 | 13.386 | 13.795 | 0.0414 | 0.0324 | 0.9024 | 0.83 | 0.8276 | 0.5 | 6 | 2.488 | 2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | classical | eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134727.csv | II rhythm | 3 | 0 | 0.999 | 0.6266 |  | -0.281 | -0.249 | 0.172 | 0.1491 | 0.8159 | 0.7847 |  | 1.25 | -3240 | 6.636 | -2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | I | 0 | 0 | 0.9992 | 0.9643 | 0.992 | 11.608 | 14.739 | 0.0382 | 0.03 | 0.918 | 0.9422 | 0.9381 | 0.25 | 6 | 2.494 | -2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | aVR | 0 | 1 | 0.9984 | 0.9085 | 0.9962 | 7.691 | 20.926 | 0.0518 | 0.0493 | 0.8857 | 0.9685 | 0.9685 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | V1 | 0 | 2 | 0.9992 | 0.9709 | 0.992 | 12.392 | 17.892 | 0.0547 | 0.0522 | 0.9506 | 0.955 | 0.955 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | V4 | 0 | 3 | 0.9984 | 0.9089 | 0.9972 | 7.795 | 12.508 | 0.0941 | 0.079 | 0.8957 | 0.9771 | 0.9749 | 0.75 | 4 | 2.482 | 8 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | II | 1 | 0 | 0.9992 | 0.984 | 0.9954 | 15.139 | 17.503 | 0.0609 | 0.0235 | 0.9514 | 0.9215 | 0.9129 | 0.25 | 4 | 2.494 | -2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | aVL | 1 | 1 | 0.9984 | 0.8152 | 0.9921 | 4.482 | 17.728 | 0.0505 | 0.0504 | 0.7824 | 0.9293 | 0.9293 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | V2 | 1 | 2 | 0.9992 | 0.9796 | 0.9939 | 13.967 | 19.163 | 0.0789 | 0.0576 | 0.9616 | 0.9766 | 0.9757 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | V5 | 1 | 3 | 0.9984 | 0.7542 | 0.9966 | 3.257 | 7.719 | 0.1048 | 0.1048 | 0.7443 | 0.9517 | 0.9435 | 1 | 6 | 2.476 | 8 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | III | 2 | 0 | 0.9952 | 0.9482 | 0.9931 | 10.096 | 14.031 | 0.0448 | 0.0296 | 0.9005 | 0.9436 | 0.9388 | -0.25 | -6 | 2.5 | 4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | aVF | 2 | 1 | 0.9992 | 0.9333 | 0.9934 | 9.103 | 18.836 | 0.0366 | 0.0347 | 0.9043 | 0.9173 | 0.917 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | V3 | 2 | 2 | 0.9992 | 0.9923 | 0.9947 | 18.224 | 19.874 | 0.0935 | 0.0275 | 0.9848 | 0.9782 | 0.9782 | 0 | 2 | 2.498 | -2 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | V6 | 2 | 3 | 0.9984 | 0.9608 | 0.9975 | 11.352 | 14.975 | 0.0493 | 0.0409 | 0.9322 | 0.923 | 0.9202 | 0.5 | 4 | 2.488 | 4 | ok | ok
B (ideal photos) | 3908 | MI | pair_20261013_134722 | learned | eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv | II rhythm | 3 | 0 | 0.9958 | 0.9591 | 0.9621 | 11.248 | 11.571 | 0.0548 | 0.0395 | 0.9372 | 0.9377 | 0.9377 | 0 | 2 | 9.998 | -2 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | I | 0 | 0 | 0.9784 | 0.9939 | 0.9961 | 18.958 | 18.958 | 0.0612 | 0.0241 | 0.964 | 0.9192 | 0.9177 | 0.25 | 2 | 2.494 | 2 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | aVR | 0 | 1 | 0.948 | 0.8343 | 0.8359 | 4.898 | 4.898 | 0.094 | 0.0959 | 0.5965 | 0.3898 | 0.3895 | 0.25 | 2 | 2.494 | 0 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | V1 | 0 | 2 | 0.9608 | 0.3964 | 0.401 | -3.575 | -3.57 | 0.6015 | 0.563 | 0.5437 | 0.8688 | 0.8688 | -0.75 | -6 | 2.5 | -2 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | V4 | 0 | 3 | 0.7368 | 0.5668 | 0.5893 | 1.312 | 1.328 | 0.2458 | 0.2341 | 0.4807 | 0.291 | 0.3103 | 2.5 | 16 | 2.44 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | II | 1 | 0 | 0.9784 | 0.8921 | 0.9818 | 6.557 | 14.468 | 0.1018 | 0.0583 | 0.9191 | 0.9134 | 0.9117 | 0 | -8 | 2.492 | 8 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | aVL | 1 | 1 | 0.9592 | 0.9497 | 0.9632 | 9.953 | 10.599 | 0.0761 | 0.0655 | 0.9756 | 0.8982 | 0.8973 | 1 | 10 | 2.476 | 4 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | V2 | 1 | 2 | 0.9608 | 0.4353 | 0.45 | -1.985 | -1.894 | 0.5978 | 0.5752 | 0.5008 | 0.9452 | 0.945 | 0.25 | 12 | 2.488 | -10 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | V5 | 1 | 3 | 0.9792 | -0.0546 | 0.1089 | -7.279 | -7.067 | 0.2721 | 0.2629 | -0.1005 | 0.8278 | 0.8093 | 6.75 | 100 | 2.342 | -128 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | III | 2 | 0 | 0.9704 | 0.8847 | 0.9805 | 6.808 | 9.284 | 0.0822 | 0.0811 | 0.8198 | 0.8288 | 0.8289 | -1.25 | -26 | 2.5 | 12 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | aVF | 2 | 1 | 0.9608 | 0.962 | 0.9624 | 11.347 | 11.347 | 0.0348 | 0.0308 | 0.9268 | 0.9086 | 0.9086 | 0.25 | 2 | 2.494 | 0 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | V3 | 2 | 2 | 0.9632 | 0.7296 | 0.7462 | 0.323 | 0.512 | 0.3771 | 0.3347 | 0.9547 | 0.9196 | 0.9154 | -0.25 | 8 | 2.492 | -10 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | V6 | 2 | 3 | 0.9768 | 0.8006 | 0.8067 | 4.432 | 4.546 | 0.0669 | 0.0556 | 0.6146 | 0.6401 | 0.6401 | 0 | 2 | 2.498 | -2 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv | II rhythm | 3 | 0 | 0.9978 | 0.483 |  | -0.974 | -0.955 | 0.1527 | 0.1501 | 0.5671 | 0.8214 |  | -1 | 4244 | 5.756 | -2 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | I | 0 | 0 | 0.9976 | 0.9911 | 0.9967 | 17.404 | 19.974 | 0.061 | 0.0289 | 0.963 | 0.9651 | 0.9614 | 0.25 | 0 | 2.494 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | aVR | 0 | 1 | 0.9976 | 0.9814 | 0.9955 | 14.388 | 20.186 | 0.0369 | 0.0321 | 0.9493 | 0.9763 | 0.9764 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | V1 | 0 | 2 | 0.9976 | 0.9846 | 0.9987 | 15.188 | 23.062 | 0.0736 | 0.0649 | 0.9654 | 0.9809 | 0.9788 | -0.25 | 4 | 2.496 | -6 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | V4 | 0 | 3 | 0.9976 | 0.9657 | 0.996 | 11.805 | 13.702 | 0.0829 | 0.0699 | 0.9521 | 0.994 | 0.9932 | 1 | 4 | 2.476 | 6 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | II | 1 | 0 | 0.9944 | 0.8863 | 0.9876 | 6.406 | 16.067 | 0.0969 | 0.0593 | 0.9004 | 0.972 | 0.9709 | 0 | -8 | 2.492 | 8 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | aVL | 1 | 1 | 0.9976 | 0.9882 | 0.9973 | 16.402 | 22.478 | 0.043 | 0.0312 | 0.9664 | 0.9523 | 0.9517 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | V2 | 1 | 2 | 0.9976 | 0.9904 | 0.9992 | 17.24 | 25.25 | 0.0721 | 0.0629 | 0.9741 | 0.9814 | 0.9807 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | V5 | 1 | 3 | 0.9976 | 0.9674 | 0.9961 | 11.986 | 12.859 | 0.029 | 0.0286 | 0.9624 | 0.9766 | 0.9765 | 0.75 | 4 | 2.482 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | III | 2 | 0 | 0.9832 | 0.8535 | 0.9302 | 3.13 | 5.282 | 0.1263 | 0.1238 | 1.1366 | 0.988 | 0.9837 | -2.25 | -42 | 2.5 | 12 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | aVF | 2 | 1 | 0.9976 | 0.9863 | 0.9907 | 15.739 | 17.424 | 0.0275 | 0.0186 | 0.9874 | 0.9529 | 0.9529 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | V3 | 2 | 2 | 0.9992 | 0.9957 | 0.9989 | 20.963 | 26.816 | 0.0311 | 0.0311 | 0.9895 | 0.9892 | 0.9892 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | V6 | 2 | 3 | 0.996 | 0.9916 | 0.9932 | 17.104 | 17.104 | 0.0338 | 0.0129 | 0.9341 | 0.9387 | 0.9037 | 0.25 | 2 | 2.494 | 0 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_143426 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv | II rhythm | 3 | 0 | 0.9964 | 0.9094 | 0.9094 | 7.422 | 7.422 | 0.0571 | 0.0571 | 0.9092 | 0.9649 | 0.9649 | 0 | 0 | 10 | -2 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | I | 0 | 0 | 0.9736 | 0.4257 | 0.6454 | 0.531 | 0.818 | 0.1952 | 0.2013 | 0.2456 | 0.0125 | 0.1767 | 4.75 | 68 | 2.386 | -6 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | aVR | 0 | 1 | 0.8712 | 0.6526 | 0.817 | 2.342 | 2.389 | 0.1267 | 0.1287 | 0.4639 | 0.2847 | 0.5827 | 4.25 | 72 | 2.398 | -4 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | V1 | 0 | 2 | 0.9624 | 0.9002 | 0.9137 | 7.252 | 7.832 | 0.1658 | 0.1618 | 0.8153 | 0.8917 | 0.8843 | 0.25 | 10 | 2.49 | -6 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | V4 | 0 | 3 | 0.9808 | 0.8854 | 0.9012 | 6.593 | 6.647 | 0.13 | 0.1274 | 0.7269 | 0.7894 | 0.7858 | 1.25 | 10 | 2.47 | -2 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | II | 1 | 0 | 0.98 | 0.6748 | 0.7982 | 2.369 | 2.733 | 0.1151 | 0.0944 | 0.5814 | 0.4101 | 0.3912 | 1.5 | 20 | 2.464 | -14 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | aVL | 1 | 1 | 0.964 | 0.7943 | 0.8152 | 4.228 | 4.345 | 0.1263 | 0.1265 | 0.7154 | 0.8849 | 0.8815 | 1 | 16 | 2.476 | 2 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | V2 | 1 | 2 | 0.7232 | 0.226 | 0.2307 | -1.726 | -1.707 | 0.6114 | 0.5583 | 0.1722 | 0.0805 | 0.0796 | -0.25 | 2 | 2.498 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | V5 | 1 | 3 | 0.976 | 0.7933 | 0.8928 | 4.232 | 4.353 | 0.0696 | 0.0699 | 0.5357 | 0.5676 | 0.5543 | 1.75 | 6 | 2.458 | 4 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | III | 2 | 0 | 0.98 | 0.9473 | 0.9513 | 10.082 | 10.082 | 0.059 | 0.0556 | 0.9142 | 0.8449 | 0.8427 | -0.25 | -2 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | aVF | 2 | 1 | 0.964 | 0.5659 | 0.6144 | -1.095 | -0.823 | 0.1332 | 0.1289 | 0.7662 | 0.7021 | 0.6987 | 1.75 | 20 | 2.458 | 4 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | V3 | 2 | 2 | 0.7272 | 0.6322 | 0.6587 | 1.763 | 1.824 | 0.2789 | 0.2835 | 0.5587 | 0.6096 | 0.6153 | 2.25 | 24 | 2.444 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | V6 | 2 | 3 | 0.98 | 0.8814 | 0.9028 | 6.477 | 6.526 | 0.0542 | 0.044 | 0.7367 | 0.8628 | 0.8719 | 0.75 | 6 | 2.482 | -2 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv | II rhythm | 3 | 0 | 0.9982 | 0.7194 | 0.8228 | 2.658 | 3.187 | 0.0985 | 0.0988 | 0.689 | 0.853 | 0.8495 | 0.25 | 18 | 9.976 | -6 | ok | check | LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V3
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | I | 0 | 0 | 0.9984 | 0.9831 | 0.9868 | 14.763 | 15.701 | 0.0679 | 0.0391 | 0.9544 | 0.9748 | 0.9723 | 0.25 | 0 | 2.494 | 2 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | aVR | 0 | 1 | 0.9984 | 0.9932 | 0.9936 | 18.566 | 18.566 | 0.0269 | 0.0199 | 0.9638 | 0.9867 | 0.9864 | 0.25 | 2 | 2.494 | 0 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | V1 | 0 | 2 | 0.9976 | 0.9901 | 0.9991 | 17.008 | 23.015 | 0.0655 | 0.0526 | 0.9805 | 0.9942 | 0.9939 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | V4 | 0 | 3 | 0.9984 | 0.985 | 0.9946 | 15.288 | 15.34 | 0.0683 | 0.0468 | 0.9879 | 1.0106 | 1.0105 | 0.75 | 6 | 2.482 | 2 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | II | 1 | 0 | 0.996 | 0.9478 | 0.9845 | 9.957 | 13.188 | 0.0858 | 0.0394 | 0.9176 | 0.9776 | 0.9728 | 0.25 | -2 | 2.492 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | aVL | 1 | 1 | 0.9992 | 0.9935 | 0.9973 | 18.949 | 22.645 | 0.0394 | 0.0232 | 0.9781 | 0.9629 | 0.9629 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | V2 | 1 | 2 | 0.9976 | 0.9905 | 0.9994 | 17.257 | 24.594 | 0.0709 | 0.0628 | 0.9743 | 0.9778 | 0.9766 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | V5 | 1 | 3 | 0.9968 | 0.951 | 0.9885 | 10.242 | 11.082 | 0.0355 | 0.035 | 0.9371 | 0.9876 | 0.9865 | 0.75 | 4 | 2.482 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | III | 2 | 0 | 0.9896 | 0.8312 | 0.9296 | 5.009 | 6.753 | 0.0996 | 0.0998 | 0.8113 | 0.9766 | 0.9762 | -1.25 | -20 | 2.5 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | aVF | 2 | 1 | 0.9992 | 0.9584 | 0.9857 | 10.992 | 15.359 | 0.0379 | 0.0321 | 0.9417 | 0.9597 | 0.9578 | 0.25 | -2 | 2.492 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | V3 | 2 | 2 | 0.9968 | 0.9928 | 0.9984 | 18.649 | 24.697 | 0.0406 | 0.0406 | 0.9737 | 0.9804 | 0.9802 | -0.25 | 2 | 2.498 | -4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | V6 | 2 | 3 | 0.9992 | 0.9397 | 0.994 | 9.284 | 10.28 | 0.0467 | 0.0318 | 0.9136 | 0.9163 | 0.917 | 0.75 | 4 | 2.482 | 4 | ok | ok
B (ideal photos) | 4016 | CD | pair_20261013_144140 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv | II rhythm | 3 | 0 | 0.9992 | 0.9055 | 0.9694 | 7.238 | 7.488 | 0.0583 | 0.0583 | 0.9018 | 0.9635 | 0.9633 | 0.25 | 14 | 9.976 | -2 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | I | 0 | 0 | 0.9656 | 0.4356 | 0.4675 | -0.214 | -0.17 | 0.0963 | 0.0927 | 0.4074 | 0.1529 | 0.1459 | 1 | 14 | 2.476 | -4 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | aVR | 0 | 1 | 0.9624 | 0.564 | 0.6376 | 1.948 | 2.314 | 0.1741 | 0.1766 | 0.2804 | 0.354 | 0.3534 | -0.75 | -22 | 2.496 | 10 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | V1 | 0 | 2 | 0.54 | 0.4629 | 0.4952 | -1.435 | -1.316 | 0.3331 | 0.3386 | 0.5549 | 0.3967 | 0.3748 | 2 | 18 | 2.45 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | V4 | 0 | 3 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | II | 1 | 0 | 0.9776 | 0.1507 | 0.1842 | -0.656 | -0.526 | 0.5609 | 0.5268 | 0.1021 | -0.0086 | -0.0097 | 1.75 | -2 | 2.456 | 0 | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | aVL | 1 | 1 | 0.9632 | 0.0086 | 0.0489 | -5.618 | -5.615 | 0.479 | 0.4434 | 0.011 | -0.0005 | -0.0004 | 6.25 | 100 | 2.352 | 4 | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | V2 | 1 | 2 | 0.2424 | 0.1482 | 0.1776 | -3.263 | -3.19 | 0.6674 | 0.6429 | 0.1559 | 0 | 0 | 10 | 62 | 2.272 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | V5 | 1 | 3 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | III | 2 | 0 | 0.98 | 0.0819 | 0.1377 | -0.094 | -0.026 | 0.4917 | 0.5048 | 0.0206 | -0.0422 | 0.0895 | 10 | -100 | 2.172 | -150 | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | aVF | 2 | 1 | 0.944 | 0.0314 | 0.1504 | -0.512 | -0.383 | 0.4628 | 0.4605 | 0.0143 | -0.0773 | 0.1344 | 9.5 | -100 | 2.184 | 200 | FAILED | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | V3 | 2 | 2 | 0.8016 | 0.1518 | 0.1803 | -11.482 | -11.374 | 1.0048 | 1.0107 | 0.4554 | -0.7434 | -0.9104 | -5.5 | -26 | 2.5 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | V6 | 2 | 3 | 0.9072 | 0.1948 | 0.205 | -6.314 | -6.276 | 0.4351 | 0.3914 | 0.4018 | 0.7574 | 0.7408 | -1.5 | -26 | 2.5 | 4 | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162733.csv | II rhythm | 3 | 0 | 0.343 | 0.0572 |  | -1.889 | -1.887 | 0.5725 | 0.5521 | 0.0317 | 0 |  | -8 | 6392 | 3.608 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | I | 0 | 0 | 0.9976 | 0.8921 | 0.9479 | 6.585 | 8.479 | 0.0466 | 0.0424 | 0.9027 | 0.857 | 0.8507 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | aVR | 0 | 1 | 0.9968 | 0.9824 | 0.9955 | 14.968 | 16.804 | 0.0509 | 0.0394 | 0.9699 | 0.9698 | 0.9644 | -0.25 | -6 | 2.5 | 2 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | V1 | 0 | 2 | 0.9984 | 0.9054 | 0.9947 | 7.306 | 19.76 | 0.1282 | 0.1238 | 0.9026 | 0.954 | 0.9541 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | V4 | 0 | 3 | 0.9592 | 0.9846 | 0.9887 | 14.946 | 15.048 | 0.1221 | 0.0456 | 0.9114 | 0.9524 | 0.9498 | 0.25 | 2 | 2.494 | 2 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | II | 1 | 0 | 0.996 | 0.9369 | 0.9512 | 9.165 | 9.165 | 0.2121 | 0.1701 | 0.8068 | 0.9476 | 0.9375 | -0.25 | -4 | 2.5 | 0 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | aVL | 1 | 1 | 0.9976 | 0.9249 | 0.9905 | 8.101 | 14.651 | 0.108 | 0.0914 | 0.9677 | 0.9976 | 0.9937 | -0.25 | -10 | 2.496 | 6 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | V2 | 1 | 2 | 0.9992 | 0.9205 | 0.9949 | 8.051 | 18.691 | 0.1773 | 0.1748 | 0.9043 | 0.9709 | 0.9615 | -0.25 | 2 | 2.498 | -6 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | V5 | 1 | 3 | 0.9976 | 0.9931 | 0.996 | 18.826 | 18.826 | 0.1024 | 0.0295 | 0.9635 | 0.9649 | 0.9637 | 0.25 | 4 | 2.494 | 0 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | III | 2 | 0 | 0.9976 | 0.8495 | 0.909 | 5.73 | 6.499 | 0.2737 | 0.2582 | 0.6978 | 0.9396 | 0.9384 | 0.5 | 2 | 2.488 | 4 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | aVF | 2 | 1 | 0.9968 | 0.7285 | 0.7377 | 3.384 | 3.384 | 0.2893 | 0.2941 | 0.4947 | 0.8207 | 0.7573 | -0.25 | -4 | 2.5 | 0 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | V3 | 2 | 2 | 0.9976 | 0.8548 | 0.9876 | 5.592 | 16.096 | 0.166 | 0.1415 | 0.8466 | 0.878 | 0.8779 | 0 | 8 | 2.492 | -8 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | V6 | 2 | 3 | 0.9992 | 0.9042 | 0.9117 | 7.549 | 7.625 | 0.1033 | 0.0793 | 0.8848 | 0.9748 | 0.9661 | 0.25 | 4 | 2.494 | -2 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162729 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv | II rhythm | 3 | 0 | 0.999 | 0.9049 | 0.9807 | 7.563 | 7.771 | 0.2137 | 0.1859 | 0.8913 | 0.9584 | 0.9569 | 0.25 | 14 | 9.976 | -2 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | I | 0 | 0 | 0.9744 | 0.5017 | 0.5159 | -2.251 | -2.07 | 0.1146 | 0.1172 | 0.7192 | 0.7778 | 0.7733 | -0.25 | -8 | 2.498 | 4 | ok | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | aVR | 0 | 1 | 0.64 | 0.6246 | 0.651 | 2.214 | 2.415 | 0.1716 | 0.1713 | 0.535 | 0.6905 | 0.6795 | 0.75 | 4 | 2.482 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | V1 | 0 | 2 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | V4 | 0 | 3 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | II | 1 | 0 | 0.9784 | 0.3575 | 0.4153 | 0.662 | 0.978 | 0.4635 | 0.4526 | 0.1991 | 0.0793 | 0.0788 | 0 | -14 | 2.486 | 6 | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | aVL | 1 | 1 | 0.9264 | 0.1729 | 0.1733 | -3.859 | -3.851 | 0.3672 | 0.3621 | 0.2207 | 0.2686 | 0.2686 | 0 | -2 | 2.498 | 0 | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | V2 | 1 | 2 | 0.3008 | 0.2193 | 0.2302 | -0.372 | -0.353 | 0.4551 | 0.4609 | 0.118 | 0 | 0 | 10 | 58 | 2.272 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | V5 | 1 | 3 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | III | 2 | 0 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | aVF | 2 | 1 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | V3 | 2 | 2 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | V6 | 2 | 3 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | classical | eacg_sheet_ptbxl5830_kit_500hz_300dpi_classical_20261013_162842.csv | II rhythm | 3 | 0 | 0 | [13 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | I | 0 | 0 | 0.9968 | 0.9589 | 0.9748 | 10.978 | 13.077 | 0.0358 | 0.0255 | 0.9289 | 0.916 | 0.9162 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | aVR | 0 | 1 | 0.9968 | 0.9617 | 0.9955 | 11.653 | 17.594 | 0.0657 | 0.0578 | 0.9382 | 0.9644 | 0.9617 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | V1 | 0 | 2 | 0.9952 | 0.9176 | 0.9921 | 7.947 | 17.885 | 0.1206 | 0.115 | 0.9015 | 0.9468 | 0.9468 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | V4 | 0 | 3 | 0.9952 | 0.9387 | 0.9819 | 9.616 | 11.596 | 0.144 | 0.0842 | 0.9067 | 0.9623 | 0.9561 | 0.5 | 2 | 2.488 | 4 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | II | 1 | 0 | 0.9968 | 0.9606 | 0.9606 | 11.372 | 11.372 | 0.1947 | 0.1319 | 0.8945 | 0.9546 | 0.9546 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | aVL | 1 | 1 | 0.996 | 0.958 | 0.9913 | 10.478 | 14.389 | 0.0873 | 0.0695 | 1.0147 | 1.0756 | 1.0536 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | V2 | 1 | 2 | 0.9984 | 0.9071 | 0.9961 | 7.39 | 20.699 | 0.1909 | 0.1886 | 0.888 | 0.9586 | 0.9586 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | V5 | 1 | 3 | 0.9992 | 0.9779 | 0.9969 | 13.953 | 13.953 | 0.1156 | 0.0517 | 0.949 | 0.9679 | 0.9664 | 0.5 | 4 | 2.488 | 0 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | III | 2 | 0 | 0.9952 | 0.9328 | 0.9342 | 8.78 | 8.856 | 0.2123 | 0.1817 | 0.7801 | 0.9505 | 0.9505 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | aVF | 2 | 1 | 0.9984 | 0.7623 | 0.8142 | 3.945 | 4.012 | 0.274 | 0.2757 | 0.5903 | 0.7915 | 0.7593 | -0.75 | -10 | 2.5 | -2 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | V3 | 2 | 2 | 0.9992 | 0.8747 | 0.991 | 6.248 | 17.556 | 0.1592 | 0.1313 | 0.8647 | 0.9358 | 0.9357 | 0 | 8 | 2.492 | -8 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | V6 | 2 | 3 | 0.996 | 0.8406 | 0.8907 | 5.606 | 6.951 | 0.119 | 0.0992 | 0.7428 | 0.9577 | 0.9497 | 0.25 | 8 | 2.492 | -6 | ok | ok
B (ideal photos) | 5830 | HYP | pair_20261013_162839 | learned | eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv | II rhythm | 3 | 0 | 0.9986 | 0.9424 | 0.9604 | 9.732 | 9.751 | 0.1788 | 0.1449 | 0.928 | 0.97 | 0.968 | 0.25 | 12 | 9.976 | -2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | I | 0 | 0 | 0.8384 | 0.2376 | 0.2426 | -3.37 | -3.353 | 0.4085 | 0.3782 | 0.3154 | 0.3325 | 0.332 | 0.5 | 4 | 2.488 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | aVR | 0 | 1 | 0.9624 | 0.3707 | 0.371 | -4.472 | -4.471 | 0.4379 | 0.3796 | 0.6136 | 0.8097 | 0.8096 | 0 | 2 | 2.498 | -2 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | V1 | 0 | 2 | 0.872 | 0.1963 | 0.2046 | -1.234 | -1.197 | 0.7383 | 0.5744 | 0.1215 | 0.0954 | 0.0954 | 0.25 | 0 | 2.494 | 4 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | V4 | 0 | 3 | 0.724 | 0.182 | 0.2608 | -0.008 | 0.009 | 0.3074 | 0.2966 | 0.078 | 0.0162 | 0.0162 | 5 | 22 | 2.38 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | II | 1 | 0 | 0.9792 | 0.5312 | 0.5326 | 0.994 | 1.003 | 0.1954 | 0.18 | 0.4749 | 0.429 | 0.4274 | 0.25 | 2 | 2.494 | 0 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | aVL | 1 | 1 | 0.8256 | 0.2479 | 0.2546 | -6.176 | -6.153 | 0.3594 | 0.3261 | 0.4292 | 0.731 | 0.731 | 0 | -4 | 2.496 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | V2 | 1 | 2 | 0.876 | 0.3629 | 0.3838 | -1.168 | -1.04 | 0.5242 | 0.5121 | 0.3609 | 0.3601 | 0.3589 | 0.5 | 10 | 2.488 | -6 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | V5 | 1 | 3 | 0.976 | 0.6148 | 0.6969 | 1.644 | 1.656 | 0.2469 | 0.2512 | 0.2046 | 0.1598 | 0.1774 | 1.5 | 10 | 2.464 | 2 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | III | 2 | 0 | 0.9384 | 0.4308 | 0.4897 | -3.464 | -3.148 | 0.1412 | 0.133 | 0.6892 | 0.6918 | 0.6872 | 0.5 | 12 | 2.488 | -6 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | aVF | 2 | 1 | 0.9624 | 0.93 | 0.9468 | 9.047 | 10.129 | 0.0516 | 0.0325 | 0.8305 | 0.8074 | 0.8071 | 0 | -4 | 2.496 | 4 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | V3 | 2 | 2 | 0.9616 | 0.3997 | 0.455 | -1.623 | -1.218 | 0.5777 | 0.4936 | 0.4596 | 0.568 | 0.5667 | 0 | 10 | 2.49 | -10 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | V6 | 2 | 3 | 0.98 | 0.8261 | 0.8456 | 5.222 | 5.403 | 0.1506 | 0.1258 | 0.6594 | 0.7109 | 0.709 | 0.75 | 10 | 2.482 | -4 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv | II rhythm | 3 | 0 | 0.9682 | 0.4583 | 0.5669 | 0.254 | 0.347 | 0.223 | 0.1988 | 0.424 | 0.5663 | 0.5656 | 0.5 | 30 | 9.95 | -6 | ok | check | LEAD_COVERAGE/warn/V4
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | I | 0 | 0 | 0.996 | 0.9691 | 0.9778 | 12.347 | 13.765 | 0.0882 | 0.0619 | 0.9497 | 0.9586 | 0.9588 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | aVR | 0 | 1 | 0.9984 | 0.9586 | 0.9835 | 11.226 | 14.181 | 0.0874 | 0.0623 | 0.9238 | 0.9273 | 0.92 | -0.25 | -6 | 2.5 | 4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | V1 | 0 | 2 | 0.9984 | 0.9156 | 0.9917 | 8.074 | 17.724 | 0.2247 | 0.1967 | 0.8763 | 0.9651 | 0.9652 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | V4 | 0 | 3 | 0.9992 | 0.9183 | 0.9947 | 8.248 | 9.786 | 0.1395 | 0.1147 | 0.8703 | 0.9238 | 0.9173 | 0.5 | 4 | 2.488 | 2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | II | 1 | 0 | 0.9976 | 0.9772 | 0.984 | 13.786 | 15.28 | 0.0539 | 0.0413 | 0.9528 | 0.9594 | 0.9594 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | aVL | 1 | 1 | 0.9968 | 0.9265 | 0.9749 | 8.647 | 12.993 | 0.0607 | 0.0592 | 0.862 | 0.8998 | 0.8997 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | V2 | 1 | 2 | 0.9984 | 0.9548 | 0.9898 | 10.554 | 16.747 | 0.1331 | 0.1328 | 0.916 | 0.9137 | 0.9131 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | V5 | 1 | 3 | 0.9992 | 0.9367 | 0.9961 | 9.375 | 12.334 | 0.1323 | 0.1032 | 0.9157 | 0.9686 | 0.9679 | 0.5 | 2 | 2.488 | 4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | III | 2 | 0 | 0.9976 | 0.9561 | 0.9707 | 10.365 | 11.766 | 0.0377 | 0.0271 | 0.843 | 0.8909 | 0.891 | 0 | 2 | 2.498 | -2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | aVF | 2 | 1 | 0.9968 | 0.8591 | 0.963 | 6.135 | 11.767 | 0.0619 | 0.0455 | 0.8056 | 0.9226 | 0.9228 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | V3 | 2 | 2 | 0.9976 | 0.964 | 0.9894 | 11.536 | 16.744 | 0.1209 | 0.1085 | 0.9437 | 0.9476 | 0.9474 | 0 | 4 | 2.496 | -4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | V6 | 2 | 3 | 0.9976 | 0.9891 | 0.9976 | 16.938 | 17.842 | 0.1015 | 0.0327 | 0.9782 | 0.9774 | 0.9765 | 0.25 | 2 | 2.494 | 2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162442 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162442.csv | II rhythm | 3 | 0 | 0.9994 | 0.8664 | 0.9627 | 6.102 | 6.163 | 0.122 | 0.1014 | 0.8595 | 0.9597 | 0.9574 | 0.25 | 14 | 9.976 | -2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | I | 0 | 0 | 0.9792 | 0.6112 | 0.7021 | 1.535 | 1.854 | 0.2123 | 0.215 | 0.1966 | 0.1503 | 0.1507 | 0 | -12 | 2.488 | 12 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | aVR | 0 | 1 | 0.9592 | 0.7902 | 0.7975 | 3.56 | 3.582 | 0.1525 | 0.1505 | 0.3992 | 0.362 | 0.3547 | 0.25 | 0 | 2.494 | 2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | V1 | 0 | 2 | 0.96 | 0.1872 | 0.3467 | 0.116 | 0.148 | 0.4793 | 0.4917 | 0.0222 | 0.0149 | 0.0259 | -2.75 | 100 | 2.4 | -150 | UNRELIABLE | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | V4 | 0 | 3 | 0.9784 | -0.0106 | 0.0821 | -4.478 | -4.111 | 0.4591 | 0.4963 | -0.0109 | -0.1435 | 0.0085 | 10 | -12 | 2.26 | 64 | UNRELIABLE | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | II | 1 | 0 | 0.9768 | 0.7266 | 0.8057 | 3.079 | 3.702 | 0.1398 | 0.1416 | 0.385 | 0.3063 | 0.3065 | 0 | -8 | 2.492 | 8 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | aVL | 1 | 1 | 0.9056 | 0.4318 | 0.4512 | -0.048 | -0.035 | 0.1646 | 0.161 | 0.3944 | 0.3569 | 0.3564 | 1.75 | 16 | 2.458 | 2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | V2 | 1 | 2 | 0.9552 | 0.4858 | 0.4982 | 0.918 | 0.967 | 0.4017 | 0.4027 | 0.1898 | 0.0143 | 0.0158 | 0.25 | 10 | 2.49 | -6 | UNRELIABLE | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | V5 | 1 | 3 | 0.9784 | 0.5159 | 0.581 | 0.939 | 1.004 | 0.266 | 0.2724 | 0.115 | 0.0671 | 0.0671 | 1.5 | 24 | 2.464 | -12 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | III | 2 | 0 | 0.9344 | 0.5417 | 0.5455 | -1.645 | -1.62 | 0.1172 | 0.1079 | 0.7529 | 0.6252 | 0.6236 | 0.25 | 4 | 2.494 | -2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | aVF | 2 | 1 | 0.9608 | 0.888 | 0.9092 | 7.059 | 7.831 | 0.0558 | 0.0409 | 0.7364 | 0.7235 | 0.7235 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | V3 | 2 | 2 | 0.9592 | 0.4298 | 0.4541 | 0.722 | 0.822 | 0.3682 | 0.3768 | 0.1659 | 0.0307 | 0.0374 | 0.25 | 18 | 2.482 | -14 | UNRELIABLE | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | V6 | 2 | 3 | 0.9784 | 0.6217 | 0.7032 | 1.543 | 1.6 | 0.1958 | 0.1922 | 0.1891 | 0.1849 | 0.1931 | 1.75 | 16 | 2.458 | 8 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | classical | eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv | II rhythm | 3 | 0 | 0.999 | 0.3075 | 0.3377 | -0.147 | -0.087 | 0.2214 | 0.2082 | 0.2361 | 0.1264 | 0.1236 | 0.5 | 24 | 9.95 | 8 | UNRELIABLE | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | I | 0 | 0 | 0.9984 | 0.9541 | 0.9725 | 10.664 | 12.497 | 0.098 | 0.0752 | 0.9189 | 0.9564 | 0.9524 | -0.25 | -6 | 2.5 | 4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | aVR | 0 | 1 | 0.9968 | 0.9501 | 0.9805 | 10.418 | 14.447 | 0.0912 | 0.0684 | 0.925 | 0.941 | 0.9408 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | V1 | 0 | 2 | 0.9984 | 0.9379 | 0.987 | 9.397 | 15.612 | 0.1996 | 0.1689 | 0.8831 | 0.9554 | 0.9541 | -0.25 | 2 | 2.498 | -6 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | V4 | 0 | 3 | 0.9992 | 0.8877 | 0.9909 | 6.953 | 9.851 | 0.1539 | 0.1331 | 0.8044 | 0.8963 | 0.8931 | 0.5 | 2 | 2.488 | 4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | II | 1 | 0 | 0.9952 | 0.9668 | 0.9798 | 12.185 | 14.252 | 0.0624 | 0.0496 | 0.9366 | 0.9413 | 0.9411 | 0.25 | 0 | 2.494 | 2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | aVL | 1 | 1 | 0.9992 | 0.9159 | 0.9712 | 8.082 | 12.447 | 0.0647 | 0.0632 | 0.8508 | 0.8802 | 0.8802 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | V2 | 1 | 2 | 0.9968 | 0.9305 | 0.9844 | 8.671 | 14.969 | 0.1654 | 0.165 | 0.9102 | 0.902 | 0.8937 | 0.25 | 8 | 2.492 | -6 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | V5 | 1 | 3 | 0.9992 | 0.9082 | 0.9932 | 7.848 | 10.222 | 0.1481 | 0.123 | 0.8592 | 0.9564 | 0.9524 | 0.75 | 4 | 2.482 | 6 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | III | 2 | 0 | 0.9944 | 0.9542 | 0.9652 | 10.196 | 11.162 | 0.0381 | 0.0276 | 0.8391 | 0.819 | 0.819 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | aVF | 2 | 1 | 0.9992 | 0.8736 | 0.9551 | 6.602 | 10.962 | 0.059 | 0.0431 | 0.8173 | 0.8947 | 0.8948 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | V3 | 2 | 2 | 0.9968 | 0.9066 | 0.9846 | 7.387 | 15.193 | 0.185 | 0.1749 | 0.894 | 0.9453 | 0.9449 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | V6 | 2 | 3 | 0.9992 | 0.9823 | 0.9951 | 14.86 | 15.152 | 0.1058 | 0.0415 | 0.9638 | 0.9694 | 0.9671 | 0.5 | 8 | 2.488 | -2 | ok | ok
B (ideal photos) | 7916 | HYP | pair_20261013_162603 | learned | eacg_sheet_ptbxl7916_kit_500hz_300dpi_model_20261013_162603.csv | II rhythm | 3 | 0 | 0.9988 | 0.9292 | 0.956 | 8.917 | 8.917 | 0.1017 | 0.0733 | 0.905 | 0.9468 | 0.9433 | 0.25 | 12 | 9.976 | 0 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | I | 0 | 0 | 0.9792 | 0.6932 | 0.7373 | 2.194 | 2.832 | 0.1084 | 0.1048 | 0.6846 | 0.8124 | 0.8127 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | aVR | 0 | 1 | 0.9624 | 0.9777 | 0.9828 | 13.339 | 13.657 | 0.0583 | 0.0312 | 0.9027 | 0.8368 | 0.8358 | 0.25 | 2 | 2.494 | 0 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | V1 | 0 | 2 | 0.9608 | 0.8557 | 0.9776 | 5.396 | 9.793 | 0.12 | 0.1177 | 0.8584 | 0.9253 | 0.9238 | 0.75 | 20 | 2.48 | -10 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | V4 | 0 | 3 | 0.9776 | 0.5763 | 0.6038 | 1.886 | 2.031 | 0.2982 | 0.2965 | 0.3061 | 0.0132 | 0.0142 | 0.75 | 2 | 2.482 | 6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | II | 1 | 0 | 0.9792 | 0.6637 | 0.716 | 1.759 | 2.515 | 0.1429 | 0.1242 | 0.6781 | 0.7362 | 0.7371 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | aVL | 1 | 1 | 0.9608 | 0.9324 | 0.9789 | 8.8 | 13.523 | 0.0318 | 0.0267 | 0.8788 | 0.7316 | 0.7316 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | V2 | 1 | 2 | 0.9608 | 0.7467 | 0.8941 | 3.854 | 6.826 | 0.113 | 0.1048 | 0.5322 | 0.4117 | 0.4091 | 0.5 | 20 | 2.48 | -12 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | V5 | 1 | 3 | 0.9776 | 0.6876 | 0.6937 | 2.782 | 2.782 | 0.2203 | 0.2114 | 0.3827 | 0.1229 | 0.1205 | 0.5 | 6 | 2.488 | 0 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | III | 2 | 0 | 0.9768 | 0.7821 | 0.8856 | 3.301 | 4.124 | 0.0886 | 0.0366 | 0.8373 | 0.6735 | 0.6718 | -0.75 | -12 | 2.5 | 2 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | aVF | 2 | 1 | 0.9608 | 0.9823 | 0.9927 | 14.467 | 17.533 | 0.0875 | 0.0269 | 0.9452 | 0.7109 | 0.7085 | 0.25 | 0 | 2.494 | 4 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | V3 | 2 | 2 | 0.96 | 0.5677 | 0.8255 | 1.815 | 4.419 | 0.2237 | 0.2268 | 0.3806 | 0.2303 | 0.3681 | 0.75 | 24 | 2.476 | -12 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | V6 | 2 | 3 | 0.9792 | 0.7582 | 0.7973 | 3.943 | 4.037 | 0.1292 | 0.108 | 0.5816 | 0.6642 | 0.6606 | 0.75 | 12 | 2.482 | -2 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv | II rhythm | 3 | 0 | 0.999 | 0.7641 | 0.8296 | 2.78 | 3.129 | 0.1339 | 0.1326 | 0.8152 | 0.7648 | 0.7645 | 0.25 | 18 | 9.976 | -6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | I | 0 | 0 | 0.9984 | 0.9416 | 0.9936 | 9.482 | 15.622 | 0.0495 | 0.0453 | 0.9147 | 0.9511 | 0.9503 | -0.25 | -8 | 2.498 | 4 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | aVR | 0 | 1 | 0.9992 | 0.9618 | 0.9893 | 11.321 | 16.6 | 0.0623 | 0.0393 | 0.9175 | 0.9223 | 0.9224 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | V1 | 0 | 2 | 0.9976 | 0.9341 | 0.983 | 8.881 | 14.605 | 0.0845 | 0.0788 | 0.9172 | 0.9477 | 0.9468 | 0.25 | 10 | 2.49 | -8 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | V4 | 0 | 3 | 0.9968 | 0.943 | 0.9956 | 9.893 | 12.016 | 0.1686 | 0.118 | 0.8935 | 0.9458 | 0.941 | 0.5 | 2 | 2.488 | 4 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | II | 1 | 0 | 0.9984 | 0.9083 | 0.9942 | 7.62 | 17.168 | 0.0853 | 0.0632 | 0.879 | 0.8735 | 0.8713 | -0.25 | -10 | 2.496 | 6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | aVL | 1 | 1 | 0.9984 | 0.9086 | 0.9916 | 7.481 | 16.472 | 0.0356 | 0.0311 | 0.876 | 0.8907 | 0.8856 | -0.25 | -10 | 2.496 | 8 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | V2 | 1 | 2 | 0.9984 | 0.9493 | 0.9889 | 10.432 | 15.83 | 0.075 | 0.0491 | 0.8865 | 0.8748 | 0.874 | -0.25 | 2 | 2.498 | -6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | V5 | 1 | 3 | 0.9976 | 0.9708 | 0.9914 | 12.678 | 12.678 | 0.1294 | 0.0677 | 0.9429 | 0.952 | 0.9483 | 0.5 | 6 | 2.488 | 0 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | III | 2 | 0 | 0.9936 | 0.912 | 0.9234 | 7.657 | 7.685 | 0.086 | 0.0222 | 0.8846 | 0.7895 | 0.7789 | -0.25 | -4 | 2.5 | 2 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | aVF | 2 | 1 | 0.9984 | 0.9774 | 0.9964 | 13.458 | 20.707 | 0.0591 | 0.0302 | 0.9385 | 0.808 | 0.8081 | 0 | -6 | 2.494 | 6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | V3 | 2 | 2 | 0.9968 | 0.9079 | 0.9786 | 7.8 | 13.726 | 0.1162 | 0.1138 | 0.8198 | 0.898 | 0.9089 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | V6 | 2 | 3 | 0.9984 | 0.9903 | 0.9936 | 17.215 | 17.215 | 0.0747 | 0.0234 | 0.968 | 0.9685 | 0.9597 | 0.25 | 4 | 2.494 | 0 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162024 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162024.csv | II rhythm | 3 | 0 | 0.9988 | 0.9185 | 0.9781 | 8.027 | 8.162 | 0.0724 | 0.0725 | 0.8737 | 0.9355 | 0.9215 | 0.25 | 14 | 9.976 | -2 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | I | 0 | 0 | 0.7296 | 0.2845 | 0.2918 | -4.837 | -4.825 | 0.2541 | 0.2355 | 0.4445 | 0.5519 | 0.5512 | 0.5 | 8 | 2.488 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | aVR | 0 | 1 | 0.9584 | 0.921 | 0.9257 | 7.918 | 8.025 | 0.07 | 0.0582 | 0.7596 | 0.4881 | 0.4867 | 0.25 | 2 | 2.494 | 0 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | V1 | 0 | 2 | 0.9616 | 0.8335 | 0.9336 | 5.002 | 6.895 | 0.1267 | 0.1231 | 0.7826 | 0.8615 | 0.8617 | 1 | 20 | 2.476 | -6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | V4 | 0 | 3 | 0.7752 | 0.4196 | 0.4463 | 0.978 | 1.04 | 0.3267 | 0.3292 | 0.1923 | 0.1167 | 0.1112 | 1.25 | 2 | 2.47 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | II | 1 | 0 | 0.976 | 0.8393 | 0.862 | 5.391 | 5.556 | 0.1028 | 0.0817 | 0.708 | 0.3676 | 0.365 | 0.75 | 8 | 2.482 | 2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | aVL | 1 | 1 | 0.9608 | 0.9487 | 0.9811 | 9.901 | 13.669 | 0.0298 | 0.0235 | 0.8592 | 0.7009 | 0.7009 | 0 | -4 | 2.496 | 4 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | V2 | 1 | 2 | 0.9592 | 0.7342 | 0.8802 | 3.683 | 5.693 | 0.1148 | 0.1068 | 0.5533 | 0.5135 | 0.5135 | 1.25 | 22 | 2.47 | -8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | V5 | 1 | 3 | 0.9784 | 0.6891 | 0.7508 | 2.879 | 3.039 | 0.2217 | 0.2091 | 0.402 | 0.2294 | 0.225 | 1 | 8 | 2.476 | 6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | III | 2 | 0 | 0.9768 | 0.8675 | 0.8877 | 5.895 | 6.566 | 0.0892 | 0.0272 | 0.8221 | 0.6559 | 0.6547 | 0.25 | 6 | 2.494 | -2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | aVF | 2 | 1 | 0.96 | 0.9683 | 0.9816 | 11.966 | 14.102 | 0.0628 | 0.0358 | 0.9139 | 0.578 | 0.5779 | 0.25 | -2 | 2.492 | 4 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | V3 | 2 | 2 | 0.9592 | 0.3984 | 0.5852 | -0.043 | 0.602 | 0.3008 | 0.2809 | 0.3511 | 0.3215 | 0.5091 | 1.25 | 26 | 2.47 | -8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | V6 | 2 | 3 | 0.9248 | 0.4727 | 0.4874 | -2.116 | -2.104 | 0.2458 | 0.217 | 0.6668 | 0.681 | 0.681 | 0.5 | 8 | 2.488 | 0 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | classical | eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv | II rhythm | 3 | 0 | 0.9986 | 0.7903 | 0.8477 | 3.896 | 3.896 | 0.1186 | 0.1166 | 0.7676 | 0.6985 | 0.6943 | 0.5 | 30 | 9.95 | 0 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/V4
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | I | 0 | 0 | 0.9976 | 0.9857 | 0.9929 | 15.523 | 18.497 | 0.0307 | 0.0226 | 0.9631 | 0.9488 | 0.9488 | 0 | -2 | 2.498 | 2 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | aVR | 0 | 1 | 0.996 | 0.9631 | 0.9928 | 11.432 | 17.798 | 0.0611 | 0.0388 | 0.9028 | 0.9381 | 0.9381 | 0 | -4 | 2.496 | 4 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | V1 | 0 | 2 | 0.9992 | 0.9472 | 0.9844 | 9.845 | 15.086 | 0.0773 | 0.0705 | 0.9279 | 0.9647 | 0.9646 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | V4 | 0 | 3 | 0.996 | 0.9261 | 0.992 | 8.795 | 12.839 | 0.1785 | 0.1339 | 0.8589 | 0.8866 | 0.8862 | 0.25 | -2 | 2.492 | 4 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | II | 1 | 0 | 0.996 | 0.9909 | 0.9957 | 17.181 | 17.181 | 0.0636 | 0.021 | 0.961 | 0.9156 | 0.8993 | 0.25 | 2 | 2.494 | 0 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | aVL | 1 | 1 | 0.9984 | 0.9092 | 0.992 | 7.56 | 15.396 | 0.0351 | 0.0308 | 0.8493 | 0.8856 | 0.8832 | -0.25 | -10 | 2.496 | 8 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | V2 | 1 | 2 | 0.9976 | 0.9407 | 0.9926 | 9.791 | 18.533 | 0.0778 | 0.0529 | 0.8767 | 0.9286 | 0.9306 | 0 | 6 | 2.494 | -6 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | V5 | 1 | 3 | 0.9976 | 0.9483 | 0.9937 | 10.223 | 13.755 | 0.142 | 0.0898 | 0.922 | 0.9741 | 0.9641 | 0.5 | 2 | 2.488 | 4 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | III | 2 | 0 | 0.9952 | 0.9664 | 0.9664 | 11.44 | 11.44 | 0.0854 | 0.0144 | 0.8803 | 0.8591 | 0.8591 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | aVF | 2 | 1 | 0.9984 | 0.9606 | 0.9964 | 11.104 | 19.868 | 0.0629 | 0.0396 | 0.9268 | 0.8849 | 0.8777 | -0.25 | -10 | 2.496 | 8 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | V3 | 2 | 2 | 0.9976 | 0.8724 | 0.9797 | 6.471 | 14.001 | 0.1345 | 0.1327 | 0.781 | 0.9383 | 0.9388 | 0 | 8 | 2.492 | -8 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | V6 | 2 | 3 | 0.9976 | 0.9934 | 0.9934 | 18.781 | 18.781 | 0.0749 | 0.0196 | 0.9647 | 0.9859 | 0.9859 | 0 | 0 | 2.5 | 0 | ok | ok
B (ideal photos) | 8504 | CD | pair_20261013_162240 | learned | eacg_sheet_ptbxl8504_kit_500hz_300dpi_model_20261013_162240.csv | II rhythm | 3 | 0 | 0.999 | 0.9161 | 0.9743 | 7.887 | 7.89 | 0.0736 | 0.0736 | 0.8763 | 0.9195 | 0.8976 | 0.25 | 14 | 9.976 | -2 | ok | ok

## Excluded (Set B)
file | reason
eacg_sheet_photo_classical_20261013_120746.csv | no benchmark target selected (not scorable)
eacg_sheet_photo_classical_20261013_123033.csv | no benchmark target selected (not scorable)
eacg_sheet_photo_classical_20261013_132824.csv | no benchmark target selected (not scorable)
eacg_sheet_photo_model_20261013_120742.csv | no benchmark target selected (not scorable)
eacg_sheet_photo_model_20261013_123029.csv | no benchmark target selected (not scorable)
eacg_sheet_photo_model_20261013_132821.csv | no benchmark target selected (not scorable)
eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145450.csv | same lead data as eacg_sheet_ptbxl160_kit_500hz_300dpi_model_20261013_145443.csv (a second export of one run)
eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144834.csv | same lead data as eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_144827.csv (a second export of one run)
eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145138.csv | same lead data as eacg_sheet_ptbxl167_kit_500hz_300dpi_model_20261013_145129.csv (a second export of one run)
eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134157.csv | same lead data as eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134146.csv (a second export of one run)
eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134244.csv | same lead data as eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261013_134236.csv (a second export of one run)
eacg_sheet_ptbxl38_kit_500hz_300dpi_classical_20261013_115329.csv | failed retake: quality unreliable (SCALE_DISAGREE/severe)
eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115325.csv | failed retake: quality unreliable (SCALE_DISAGREE/severe)
eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115509.csv | same lead data as eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115454.csv (a second export of one run)
eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115555.csv | same lead data as eacg_sheet_ptbxl38_kit_500hz_300dpi_model_20261013_115548.csv (a second export of one run)
eacg_sheet_ptbxl3908_kit_500hz_300dpi_classical_20261013_134358.csv | failed retake: quality unreliable (FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6)
eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134354.csv | failed retake: quality unreliable (FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4)
eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134612.csv | same lead data as eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134604.csv (a second export of one run)
eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134731.csv | same lead data as eacg_sheet_ptbxl3908_kit_500hz_300dpi_model_20261013_134722.csv (a second export of one run)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429_78f.csv | same lead data as eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143429.csv (a second export of one run)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143742.csv | failed retake: quality unreliable (FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V6)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_143918.csv | failed retake: quality unreliable (NO_TICK_SCALE/severe)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144027.csv | failed retake: quality unreliable (NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144_9bf.csv | same lead data as eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144144.csv (a second export of one run)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261013_144506.csv | wrong target: the sheet photographed is PTB-XL 167 (PCC 0.969 against 167's truth, 0.01 against 4016's)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143435.csv | same lead data as eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143426.csv (a second export of one run)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143737.csv | failed retake: quality unreliable (FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_143914.csv | failed retake: quality unreliable (NO_TICK_SCALE/severe)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144024.csv | failed retake: quality unreliable (NO_TICK_SCALE/severe)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144148.csv | same lead data as eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144140.csv (a second export of one run)
eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261013_144503.csv | wrong target: the sheet photographed is PTB-XL 167 (PCC 0.969 against 167's truth, 0.01 against 4016's)
eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138_b3a.csv | same lead data as eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261013_133138.csv (a second export of one run)
eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132959.csv | same lead data as eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_132946.csv (a second export of one run)
eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133144.csv | same lead data as eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261013_133135.csv (a second export of one run)
eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133725.csv | same lead data as eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133702.csv (a second export of one run)
eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133929.csv | same lead data as eacg_sheet_ptbxl57_kit_500hz_300dpi_model_20261013_133915.csv (a second export of one run)
eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162739.csv | same lead data as eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv (a second export of one run)
eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162741.csv | same lead data as eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162729.csv (a second export of one run)
eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162845.csv | same lead data as eacg_sheet_ptbxl5830_kit_500hz_300dpi_model_20261013_162839.csv (a second export of one run)
eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445_aa1.csv | same lead data as eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162445.csv (a second export of one run)
eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607_9a8.csv | same lead data as eacg_sheet_ptbxl7916_kit_500hz_300dpi_classical_20261013_162607.csv (a second export of one run)
eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027_b4c.csv | same lead data as eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162027.csv (a second export of one run)
eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243_ed9.csv | same lead data as eacg_sheet_ptbxl8504_kit_500hz_300dpi_classical_20261013_162243.csv (a second export of one run)
eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104_a07.csv | same lead data as eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261013_113104.csv (a second export of one run)
eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113101_ce1.csv | same lead data as eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113100.csv (a second export of one run)
eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113605.csv | same lead data as eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv (a second export of one run)
eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113606.csv | same lead data as eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261013_113548.csv (a second export of one run)