## README
item | value
What | Set C (robustness): photographs of the printed sheets under tilt, shadow and crumpling, digitised on the Galaxy A30s, both tracers, scored against the PTB-XL truth.
Conditions | tilt: phone tilted ~25-35 deg from parallel; shadow: hard shadow from a single warm lamp; crumple: crumpled sheets; crumple rescue: the failed crumple photographs re-run with the t0 guard (in-sample); tilt rescue: the failed tilt PTB-XL 9 photograph re-run with the t0 guard (in-sample)
Unit | The lead (12 per run; the rhythm strip separately in 'runs'). Medians [IQR] over leads.
Failed run | Median lead PCC below 0.5.
Rescue rows | IN-SAMPLE: the t0 guard was designed after seeing these failures (progress_log §41-42); not pooled with the conditions.
Sheets | summary = per condition and tracer, with Set B and Set A for reference; runs = one row per export; leads = every lead scored; rescue_before_after = the four crumple failures before and after the guard.
Sources | sheet_scores_setC_tilt.csv; sheet_scores_setC_shadow.csv; sheet_scores_setC_crumple.csv; sheet_scores_setC_crumple_rescue.csv; sheet_scores_setC_tilt_rescue.csv; sheet_scores.csv x test_set.csv (A, B); setC_crumple_rescue_before_after.csv
Made by | python/summarise_set_c.py (reads the committed scores; scores by score_device_sheets via score_device_set.py)
Commit | 87906d8

## summary
group | tracer | runs | leads | runs_failed | verdict_ok_check_unreliable | pcc_asis_median | pcc_asis_q1 | pcc_asis_q3 | pcc_aligned_median | pcc_aligned_q1 | pcc_aligned_q3 | snr_db_asis_median | snr_db_asis_q1 | snr_db_asis_q3 | shifted_snr_db_asis_median | shifted_snr_db_asis_q1 | shifted_snr_db_asis_q3 | rmse_mv_asis_median | rmse_mv_asis_q1 | rmse_mv_asis_q3 | gain_ratio_median | gain_ratio_q1 | gain_ratio_q3 | r_ratio_median | r_ratio_q1 | r_ratio_q3 | leads_pcc_ge_0.9 | fraction_pcc_ge_0.9
C tilt | classical | 4 | 48 | 1 | 1/3/0 | 0.563 | 0.437 | 0.728 | 0.579 | 0.503 | 0.77 | 1.359 | 0.268 | 3.092 | 1.485 | 0.468 | 3.537 | 0.141 | 0.083 | 0.198 | 0.391 | 0.218 | 0.588 | 0.176 | 0.039 | 0.325 | 0 | 0
C tilt | learned | 6 | 72 | 2 | 4/0/2 | 0.946 | 0.77 | 0.974 | 0.985 | 0.957 | 0.992 | 9.913 | 3.708 | 13.124 | 13.286 | 4.735 | 16.713 | 0.054 | 0.034 | 0.104 | 0.919 | 0.79 | 0.947 | 0.917 | 0.881 | 0.951 | 40 | 0.556
C shadow | classical | 3 | 36 | 0 | 0/2/1 | 0.556 | 0.384 | 0.744 | 0.59 | 0.444 | 0.816 | 0.98 | -0.271 | 3.514 | 1.125 | -0.218 | 4.27 | 0.12 | 0.096 | 0.27 | 0.477 | 0.304 | 0.732 | 0.394 | 0.124 | 0.847 | 3 | 0.083
C shadow | learned | 4 | 48 | 1 | 1/1/2 | 0.932 | 0.193 | 0.98 | 0.992 | 0.49 | 0.995 | 8.827 | -1.589 | 14.185 | 13.094 | -1.276 | 18.618 | 0.078 | 0.04 | 0.152 | 0.91 | 0.394 | 0.949 | 0.927 | 0.865 | 0.969 | 24 | 0.5
C crumple | classical | 7 | 84 | 6 | 0/1/6 | 0.183 | 0.026 | 0.495 | 0.353 | 0.228 | 0.739 | -0.887 | -2.997 | 1.038 | -0.654 | -2.573 | 1.41 | 0.188 | 0.12 | 0.297 | 0.135 | 0.021 | 0.35 | 0.035 | 0.001 | 0.366 | 4 | 0.048
C crumple | learned | 7 | 84 | 4 | 0/1/6 | 0.211 | 0 | 0.888 | 0.806 | 0.299 | 0.967 | -1.276 | -2.584 | 6.819 | -0.664 | -1.934 | 8.364 | 0.165 | 0.095 | 0.27 | 0.212 | 0 | 0.86 | 0.171 | 0.022 | 0.942 | 18 | 0.214
C crumple rescue | classical | 5 | 60 | 2 | 0/0/5 | 0.561 | 0.347 | 0.756 | 0.774 | 0.458 | 0.923 | 0.667 | 0.017 | 3.46 | 2.419 | 0.617 | 4.981 | 0.153 | 0.091 | 0.201 | 0.467 | 0.248 | 0.704 | 0.39 | 0.018 | 0.856 | 7 | 0.117
C crumple rescue | learned | 5 | 60 | 0 | 0/0/5 | 0.838 | 0.586 | 0.944 | 0.969 | 0.941 | 0.985 | 4.861 | 1.086 | 9.696 | 8.023 | 3.234 | 12.436 | 0.095 | 0.058 | 0.147 | 0.826 | 0.576 | 0.927 | 0.944 | 0.824 | 0.967 | 22 | 0.367
C tilt rescue | classical | 1 | 12 | 1 | 0/0/1 | 0.055 | -0.093 | 0.173 | 0.169 | 0.052 | 0.227 | -5.81 | -14.794 | -2.122 | -5.639 | -14.535 | -2.025 | 0.362 | 0.272 | 0.553 | 0.062 | -0.052 | 0.667 | 0.023 | -0.036 | 0.084 | 0 | 0
C tilt rescue | learned | 1 | 12 | 1 | 0/0/1 | -0.015 | -0.075 | 0.046 | 0.143 | 0.13 | 0.221 | -15.773 | -19.29 | -9.692 | -15.033 | -18.928 | -9.186 | 0.65 | 0.592 | 0.665 | -0.152 | -0.369 | 0.097 | 0.042 | 0.006 | 0.761 | 0 | 0
C all conditions (tilt + shadow + crumple) | classical | 14 | 168 | 7 | 1/6/7 | 0.448 | 0.139 | 0.693 | 0.537 | 0.303 | 0.775 | 0.373 | -2.019 | 2.476 | 0.566 | -1.516 | 3.109 | 0.164 | 0.107 | 0.273 | 0.31 | 0.097 | 0.604 | 0.103 | 0.013 | 0.522 | 7 | 0.042
C all conditions (tilt + shadow + crumple) | learned | 17 | 204 | 7 | 5/2/10 | 0.848 | 0.101 | 0.964 | 0.966 | 0.359 | 0.992 | 5.378 | -2.373 | 11.621 | 7.936 | -1.68 | 16.032 | 0.096 | 0.049 | 0.208 | 0.832 | 0.1 | 0.939 | 0.899 | 0.098 | 0.952 | 82 | 0.402
B ideal (reference) | classical | 24 | 288 |  |  | 0.646 | 0.393 | 0.884 | 0.72 | 0.436 | 0.925 | 2.005 | -1.218 | 6.522 | 2.408 | -1.022 | 7.708 | 0.128 | 0.057 | 0.253 | 0.589 | 0.345 | 0.809 | 0.598 | 0.258 | 0.807 | 60 | 0.208
B ideal (reference) | learned | 24 | 288 |  |  | 0.947 | 0.909 | 0.973 | 0.992 | 0.985 | 0.995 | 9.996 | 7.65 | 12.828 | 15.503 | 11.661 | 18.422 | 0.049 | 0.03 | 0.076 | 0.915 | 0.868 | 0.949 | 0.941 | 0.91 | 0.958 | 230 | 0.799
A clean (reference) | classical | 12 | 144 |  |  | 0.971 | 0.952 | 0.984 | 0.988 | 0.973 | 0.995 | 12.498 | 10.444 | 15.262 | 15.518 | 11.945 | 19.231 | 0.036 | 0.025 | 0.056 | 0.937 | 0.89 | 0.952 | 0.879 | 0.827 | 0.918 | 134 | 0.931
A clean (reference) | learned | 12 | 144 |  |  | 0.994 | 0.991 | 0.997 | 0.998 | 0.995 | 0.999 | 19.309 | 17.416 | 21.882 | 22.971 | 19.595 | 26.134 | 0.018 | 0.01 | 0.027 | 0.981 | 0.968 | 0.99 | 0.971 | 0.956 | 0.979 | 142 | 0.986

## runs
condition | pair_id | export | photo | record | tracer | header_target | photo_rotation | corners | quality | warnings | ticks_found | t0_source | t0_guard | tick_vs_grid_pct | grid_inliers | median_pcc | median_pcc_aligned | median_snr_db | median_shifted_snr_db | median_rmse_mv | median_gain_ratio | median_r_ratio | leads_pcc_ge_0.9 | rhythm_pcc | trace_ms | orient_ms
tilt | 20261014_124720 | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | eacg_20261014_124637.jpg | ptbxl1737 | classical | ptbxl1737 |  | automatic (refined=false) | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | 9 | pulses |  | 0.2 | 0.978 | 0.601 | 0.641 | 0.966 | 1.024 | 0.143 | 0.414 | 0.219 | 0 | 0.6483 | 184
tilt | 20261014_124720 | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | eacg_20261014_124637.jpg | ptbxl1737 | learned | ptbxl1737 |  | automatic (refined=false) | ok |  | 9 | pulses |  | 0.2 | 0.978 | 0.973 | 0.988 | 12.829 | 16.159 | 0.036 | 0.925 | 0.915 | 9 | 0.8937 | 19434
tilt | 20261014_125027 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | eacg_20261014_124946.jpg | ptbxl4016 | classical | ptbxl4016 |  | automatic (refined=false) | check | LEAD_COVERAGE/warn/V4 | 9 | pulses |  | 0.1 | 0.896 | 0.431 | 0.503 | 0.579 | 0.713 | 0.195 | 0.175 | 0.034 | 0 | 0.4221 | 191
tilt | 20261014_125027 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | eacg_20261014_124946.jpg | ptbxl4016 | learned | ptbxl4016 |  | automatic (refined=false) | ok |  | 9 | pulses |  | 0.1 | 0.896 | 0.968 | 0.988 | 12.067 | 13.373 | 0.053 | 0.919 | 0.957 | 10 | 0.8723 | 19615
tilt | 20261014_124927 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | eacg_20261014_124851.jpg | ptbxl40 | classical | ptbxl40 |  | automatic (refined=false) | ok |  | 8 | pulses |  | 0.4 | 0.965 | 0.586 | 0.62 | 1.748 | 1.877 | 0.133 | 0.49 | 0.26 | 0 | 0.7643 | 188
tilt | 20261014_124846 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | eacg_20261014_124801.jpg | ptbxl40 | learned | ptbxl40 |  | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | 3 | rhythm_locator |  | 21.9 | 0.584 | 0.084 | 0.198 | -4.963 | -4.465 | 0.245 | 0.109 | 0.123 | 0 | -0.1681 | 19428
tilt | 20261014_124927 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | eacg_20261014_124851.jpg | ptbxl40 | learned | ptbxl40 |  | automatic (refined=false) | ok |  | 8 | pulses |  | 0.4 | 0.965 | 0.961 | 0.99 | 11.206 | 16.148 | 0.037 | 0.946 | 0.913 | 12 | 0.8603 | 19422
tilt | 20261014_124026 | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | eacg_20261014_123942.jpg | ptbxl9 | classical | ptbxl9 |  | automatic (refined=false) | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | 9 | pulses |  | 0.5 | 0.955 | 0.592 | 0.642 | 1.663 | 1.692 | 0.115 | 0.395 | 0.176 | 0 | 0.6937 | 182
tilt | 20261014_122954 | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | eacg_20261014_122915.jpg | ptbxl9 | learned | ptbxl9 |  | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 3 | rhythm_locator |  | 29.3 | 0.538 | -0.058 | 0.099 | -10.593 | -10.224 | 0.571 | -0.108 | 0.302 | 0 | -0.3106 | 20345
tilt | 20261014_124026 | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | eacg_20261014_123942.jpg | ptbxl9 | learned | ptbxl9 |  | automatic (refined=false) | ok |  | 9 | pulses |  | 0.5 | 0.955 | 0.941 | 0.985 | 9.287 | 14.611 | 0.045 | 0.887 | 0.929 | 9 | 0.9057 | 19433
shadow | 20261014_231732 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | eacg_20261014_231610.jpg | ptbxl4016 | classical | ptbxl4016 |  | user | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 8 | pulses |  | 1 | 0.868 | 0.549 | 0.57 | 0.182 | 0.306 | 0.18 | 0.525 | 0.691 | 2 | 0.5824 | 225
shadow | 20261014_232541 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | eacg_20261014_232033.jpg | ptbxl4016 | classical | ptbxl4016 | 90 clockwise (user 90) | user | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | 9 | pulses |  | 0.6 | 0.891 | 0.55 | 0.59 | 1.318 | 1.318 | 0.221 | 0.387 | 0.174 | 0 | 0.3139 | 248
shadow | 20261014_231732 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | eacg_20261014_231610.jpg | ptbxl4016 | learned | ptbxl4016 |  | user | unreliable | SCALE_DISAGREE/severe | 8 | pulses |  | 1 | 0.868 | 0.974 | 0.993 | 12.89 | 18.119 | 0.045 | 0.944 | 0.97 | 10 | 0.6662 | 19474
shadow | 20261014_232541 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | eacg_20261014_232033.jpg | ptbxl4016 | learned | ptbxl4016 | 90 clockwise (user 90) | user | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | 9 | pulses |  | 0.6 | 0.891 | 0.926 | 0.983 | 8.192 | 13.351 | 0.085 | 0.902 | 0.957 | 7 | 0.5644 | 19653
shadow | 20261014_235356 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | eacg_20261014_235259_40.jpg | ptbxl40 | classical | ptbxl40 |  | user | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | 7 | pulses |  | 0.1 | 0.959 | 0.746 | 0.771 | 3.558 | 3.872 | 0.107 | 0.591 | 0.574 | 1 | 0.7455 | 337
shadow | 20261014_234838 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | eacg_20261014_234710_40.jpg | ptbxl40 | learned | ptbxl40 | 180 clockwise (auto 270, user 270) | user | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | 1 | rhythm_locator |  |  | 0.976 | 0.1 | 0.29 | -4.213 | -4.132 | 0.252 | 0.145 | 0.003 | 0 | 0.0477 | 19734
shadow | 20261014_235356 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | eacg_20261014_235259_40.jpg | ptbxl40 | learned | ptbxl40 |  | user | ok |  | 7 | pulses |  | 0.1 | 0.959 | 0.936 | 0.994 | 9.231 | 14.564 | 0.071 | 0.914 | 0.92 | 7 | 0.8204 | 19940
crumple | 20261015_125332 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | eacg_20261015_125242.jpg | ptbxl4016 | classical | ptbxl4016 |  | automatic (refined=false) | check | NO_PULSE/warn | 8 | rhythm_locator |  | 0.4 | 0.696 | 0.439 | 0.519 | 0.258 | 0.457 | 0.198 | 0.329 | 0.2 | 0 | 0.4088 | 369 | 3618
crumple | 20261015_125428 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | eacg_20261015_125342.jpg | ptbxl4016 | classical | ptbxl4016 | 90 clockwise (auto 90) | user | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | 2 | pulses |  |  | 0.68 | 0.025 | 0.427 | -2.563 | -2.155 | 0.298 | 0.019 | 0.027 | 0 | 0.0844 | 455 | 3701
crumple | 20261015_125638 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | eacg_20261015_125545.jpg | ptbxl4016 | classical | ptbxl4016 | 270 clockwise (auto 90, user 180) | automatic (refined=false) | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | 8 | pulses |  | 1.2 | 0.663 | 0.886 | 0.942 | 6.741 | 8.017 | 0.084 | 0.814 | 0.901 | 4 | 0.4966 | 393 | 4491
crumple | 20261015_125332 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | eacg_20261015_125242.jpg | ptbxl4016 | learned | ptbxl4016 |  | automatic (refined=false) | check | NO_PULSE/warn | 8 | rhythm_locator |  | 0.4 | 0.696 | 0.932 | 0.967 | 8.609 | 10.517 | 0.064 | 0.93 | 0.96 | 7 | 0.4935 | 21244 | 3618
crumple | 20261015_125428 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | eacg_20261015_125342.jpg | ptbxl4016 | learned | ptbxl4016 | 90 clockwise (auto 90) | user | unreliable | NO_TICK_SCALE/severe | 2 | pulses |  |  | 0.68 | 0.014 | 0.601 | -2.873 | -1.945 | 0.294 | 0.014 | 0.044 | 0 | 0.0774 | 19790 | 3701
crumple | 20261015_125638 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | eacg_20261015_125545.jpg | ptbxl4016 | learned | ptbxl4016 | 270 clockwise (auto 90, user 180) | automatic (refined=false) | unreliable | SCALE_DISAGREE/severe | 8 | pulses |  | 1.2 | 0.663 | 0.928 | 0.964 | 8.595 | 10.183 | 0.064 | 0.905 | 0.97 | 7 | 0.4656 | 19753 | 4491
crumple | 20261015_124711 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | eacg_20261015_124607.jpg | ptbxl40 | classical | ptbxl40 |  | automatic (refined=false) | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | 9 | pulses |  | 1.1 | 0.797 | 0.383 | 0.425 | 0.673 | 0.757 | 0.15 | 0.187 | 0.034 | 0 | 0.476 | 347 | 3330
crumple | 20261015_124807 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | eacg_20261015_124720.jpg | ptbxl40 | classical | ptbxl40 | 270 clockwise (auto 270) | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | 3 | pulses |  | 11.7 | 0.565 | 0.082 | 0.215 | -2.308 | -2.05 | 0.208 | 0.078 | 0.017 | 0 | 0.1401 | 374 | 3615
crumple | 20261015_124910 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | eacg_20261015_124821.jpg | ptbxl40 | classical | ptbxl40 | 270 clockwise (auto 270) | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 3 | pulses |  | 18.3 | 0.536 | 0.026 | 0.247 | -0.78 | -0.653 | 0.188 | 0.014 | 0.021 | 0 | -0.0303 | 427 | 3870
crumple | 20261015_125142 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | eacg_20261015_125056.jpg | ptbxl40 | classical | ptbxl40 |  | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | 3 | rhythm_locator |  | 20.1 | 0.665 | 0.04 | 0.151 | -4.522 | -4.006 | 0.467 | 0.102 | 0 | 0 |  | 444 | 3369
crumple | 20261015_124711 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | eacg_20261015_124607.jpg | ptbxl40 | learned | ptbxl40 |  | automatic (refined=false) | unreliable | SCALE_DISAGREE/severe | 9 | pulses |  | 1.1 | 0.797 | 0.828 | 0.979 | 5.384 | 9.216 | 0.098 | 0.727 | 0.892 | 4 | 0.7321 | 20295 | 3330
crumple | 20261015_124807 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | eacg_20261015_124720.jpg | ptbxl40 | learned | ptbxl40 | 270 clockwise (auto 270) | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 3 | pulses |  | 11.7 | 0.565 | 0.126 | 0.294 | -1.811 | -1.442 | 0.202 | 0.108 | 0.017 | 0 | 0.129 | 20003 | 3615
crumple | 20261015_124910 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | eacg_20261015_124821.jpg | ptbxl40 | learned | ptbxl40 | 270 clockwise (auto 270) | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 3 | pulses |  | 18.3 | 0.536 | -0.048 | 0.297 | -2.189 | -1.569 | 0.24 | -0.046 | 0.022 | 0 | 0.0712 | 19553 | 3870
crumple | 20261015_125142 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | eacg_20261015_125056.jpg | ptbxl40 | learned | ptbxl40 |  | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 3 | rhythm_locator |  | 20.1 | 0.665 | 0.022 | 0.292 | -2.849 | -2.121 | 0.278 | 0.019 | 0.032 | 0 | 0.0347 | 20110 | 3369
crumple rescue | 20261015_155451 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | eacg_20261015_125342.jpg | ptbxl4016 | classical | ptbxl4016 | 90 clockwise (auto 90) | automatic (refined=false) | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | 8 | ticks | fired, used: t0 x 91.2 -> 164.7, ticks 2 -> 8, tick vs grid nan% -> +0.0% (t0 from the ticks) | 0 | 0.694 | 0.687 | 0.867 | 2.557 | 4.084 | 0.127 | 0.593 | 0.839 | 4 | 0.4773 | 409 | 3583
crumple rescue | 20261015_155451 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | eacg_20261015_125342.jpg | ptbxl4016 | learned | ptbxl4016 | 90 clockwise (auto 90) | automatic (refined=false) | unreliable | T0_GUARD/severe | 8 | ticks | fired, used: t0 x 91.2 -> 164.7, ticks 2 -> 8, tick vs grid nan% -> +0.0% (t0 from the ticks) | 0 | 0.694 | 0.936 | 0.971 | 8.895 | 10.313 | 0.062 | 0.932 | 0.972 | 7 | 0.4693 | 21050 | 3583
crumple rescue | 20261015_153336 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | eacg_20261015_124720.jpg | ptbxl40 | classical | ptbxl40 | 270 clockwise (auto 270) | automatic (refined=false) | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | 8 | ticks | fired, used: t0 x 90.8 -> 156.9, ticks 3 -> 8, tick vs grid +11.7% -> -0.5% (t0 from the ticks) | -0.5 | 0.565 | 0.493 | 0.815 | 0.289 | 2.263 | 0.145 | 0.475 | 0.39 | 0 | 0.391 | 353 | 4263
crumple rescue | 20261015_154119 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | eacg_20261015_124821.jpg | ptbxl40 | classical | ptbxl40 | 270 clockwise (auto 270) | automatic (refined=false) | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | 8 | ticks | fired, used: t0 x 68.8 -> 178.7, ticks 3 -> 8, tick vs grid +18.3% -> -0.2% (t0 from the ticks) | -0.2 | 0.536 | 0.349 | 0.438 | 0.501 | 0.627 | 0.148 | 0.235 | -0.002 | 0 | 0.4204 | 430 | 3923
crumple rescue | 20261015_154211 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | eacg_20261015_125056.jpg | ptbxl40 | classical | ptbxl40 |  | automatic (refined=false) | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | 9 | ticks | fired, used: t0 x 30.0 -> 158.3, ticks 3 -> 9, tick vs grid +20.1% -> -0.2% (t0 from the ticks) | -0.2 | 0.665 | 0.626 | 0.775 | 1.88 | 2.097 | 0.138 | 0.55 | 0.695 | 1 |  | 481 | 3356
crumple rescue | 20261015_154533 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | eacg_20261015_125342.jpg | ptbxl4016 | classical | ptbxl40 | 90 clockwise (auto 90) | user | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | 8 | pulses |  | 1.2 | 0.675 | 0.625 | 0.789 | 0.911 | 3.171 | 0.184 | 0.598 | 0.82 | 2 | 0.3653 | 459 | 3548
crumple rescue | 20261015_153336 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | eacg_20261015_124720.jpg | ptbxl40 | learned | ptbxl40 | 270 clockwise (auto 270) | automatic (refined=false) | unreliable | T0_GUARD/severe | 8 | ticks | fired, used: t0 x 90.8 -> 156.9, ticks 3 -> 8, tick vs grid +11.7% -> -0.5% (t0 from the ticks) | -0.5 | 0.565 | 0.803 | 0.974 | 3.83 | 6.368 | 0.131 | 0.786 | 0.894 | 2 | 0.3705 | 19817 | 4263
crumple rescue | 20261015_154119 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | eacg_20261015_124821.jpg | ptbxl40 | learned | ptbxl40 | 270 clockwise (auto 270) | automatic (refined=false) | unreliable | T0_GUARD/severe | 8 | ticks | fired, used: t0 x 68.8 -> 178.7, ticks 3 -> 8, tick vs grid +18.3% -> -0.2% (t0 from the ticks) | -0.2 | 0.536 | 0.8 | 0.973 | 3.838 | 5.832 | 0.135 | 0.733 | 0.878 | 2 | 0.368 | 19885 | 3923
crumple rescue | 20261015_154211 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | eacg_20261015_125056.jpg | ptbxl40 | learned | ptbxl40 |  | automatic (refined=false) | unreliable | T0_GUARD/severe | 9 | ticks | fired, used: t0 x 30.0 -> 158.3, ticks 3 -> 9, tick vs grid +20.1% -> -0.2% (t0 from the ticks) | -0.2 | 0.665 | 0.787 | 0.957 | 4.119 | 5.139 | 0.119 | 0.792 | 0.948 | 2 | 0.5931 | 20388 | 3356
crumple rescue | 20261015_154533 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | eacg_20261015_125342.jpg | ptbxl4016 | learned | ptbxl40 | 90 clockwise (auto 90) | user | unreliable | SCALE_DISAGREE/severe | 8 | pulses |  | 1.2 | 0.675 | 0.935 | 0.967 | 9.04 | 10.216 | 0.061 | 0.909 | 0.957 | 9 | 0.44 | 20070 | 3548
tilt rescue | 20261015_153108 | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | eacg_20261014_122915.jpg | ptbxl9 | classical | ptbxl9 |  | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | 3 | ticks | fired, used: t0 x -1.0 -> 255.7, ticks 3 -> 3, tick vs grid +29.3% -> +4.0% (t0 from the ticks) | 4 | 0.538 | 0.055 | 0.169 | -5.81 | -5.639 | 0.362 | 0.062 | 0.023 | 0 | -0.0596 | 455 | 1855
tilt rescue | 20261015_153108 | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | eacg_20261014_122915.jpg | ptbxl9 | learned | ptbxl9 |  | automatic (refined=false) | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | 3 | ticks | fired, used: t0 x -1.0 -> 255.7, ticks 3 -> 3, tick vs grid +29.3% -> +4.0% (t0 from the ticks) | 4 | 0.538 | -0.015 | 0.143 | -15.773 | -15.033 | 0.649 | -0.152 | 0.042 | 0 | -0.2718 | 20038 | 1855

## leads
condition | export | local_time | target | ecg_id | tracer | corners | rect | calibration | timings_ms | memory_peak_mb | pss_end_mb | quality | warnings | lead | rhythm | row | col | coverage | pcc_asis | snr_db_asis | shifted_snr_db_asis | rmse_mv_asis | gain_ratio | r_ratio | align_lag_at_bound | align_overlap_s | align_scale_pct | align_lag_ms | pcc_aligned | r_ratio_aligned | shift_ms | shift_status | geometry | ticks_found | t0_source
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | I | FALSE | 0 | 0 | 0.9464 | 0.5148 | 1.327 | 1.364 | 0.1741 | 0.3857 | 0.3254 | FALSE | 2.444 | 2.25 | 24 | 0.5785 | 0.3262 | 0 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | aVR | FALSE | 0 | 1 | 0.9448 | 0.7419 | 3.648 | 4.202 | 0.0677 | 0.6076 | 0.4886 | FALSE | 2.492 | 0.25 | -2 | 0.7756 | 0.4883 | 4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V1 | FALSE | 0 | 2 | 0.048 | -0.2094 | -0.692 | -0.269 | 0.1731 | -0.0501 | 0 | TRUE | 2.4 | -10 | 100 | 0.0924 | 0 |  | UNRELIABLE | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V4 | FALSE | 0 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | II | FALSE | 1 | 0 | 0.9784 | 0.7008 | 2.957 | 3.501 | 0.0409 | 0.5175 | 0.0465 | FALSE | 2.488 | 0.5 | 6 | 0.7587 | 0.0448 | 4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | aVL | FALSE | 1 | 1 | 0.96 | 0.7717 | 3.718 | 3.732 | 0.1131 | 0.442 | 0.2609 | FALSE | 2.498 | 0 | 2 | 0.7733 | 0.2609 | -2 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V2 | FALSE | 1 | 2 | 0.1304 | 0.3157 | 0.167 | 0.195 | 0.2219 | 0.1438 | 0 | FALSE | 2.272 | 10 | 30 | 0.3359 | 0 |  | UNRELIABLE | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V5 | FALSE | 1 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | III | FALSE | 2 | 0 | 0.9472 | 0.4138 | 0.605 | 0.677 | 0.1729 | 0.3186 | 0.178 | FALSE | 2.45 | 2 | 26 | 0.4399 | 0.1781 | 0 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | aVF | FALSE | 2 | 1 | 0.744 | 0.6869 | 0.567 | 0.683 | 0.079 | 0.8983 | 0.6706 | FALSE | 2.5 | -0.75 | -16 | 0.7031 | 0.6687 |  | UNRELIABLE | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V3 | FALSE | 2 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V6 | FALSE | 2 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | 20261014_124723 | ptbxl1737_kit_500hz_300dpi | 1737 | classical | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=110 grid=948 locate=120 calibrate+segment=2 trace=184 cut=4 quality=0 score*=4 classify=58 | baseline=70.7/43.2 corners=70.7/43.2 rectify=110.8/43.4 grid=115.2/62.6 locate=107.2/45.1 calibrate+segment=78.8/45.2 trace=80.7/60.4 cut=79.3/25.5 quality=79.3/25.6 score*=79.3/26.6 classify=87.7/27.8 | 185 | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | II | TRUE | 3 | 0 | 0.9456 | 0.6483 | 1.742 | 1.782 | 0.0475 | 0.6063 | 0.5703 |  | 9.976 | 0.25 | 18 | 0.6958 | 0.5654 | -2 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | I | FALSE | 0 | 0 | 0.996 | 0.9877 | 16.035 | 16.302 | 0.032 | 0.9312 | 0.9119 | FALSE | 2.498 | 0 | -2 | 0.9885 | 0.9119 | 2 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | aVR | FALSE | 0 | 1 | 0.9992 | 0.9721 | 12.815 | 16.016 | 0.0236 | 0.9418 | 0.899 | FALSE | 2.498 | 0 | -2 | 0.987 | 0.899 | 2 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | V1 | FALSE | 0 | 2 | 0.9872 | 0.9743 | 13.063 | 19.388 | 0.0355 | 0.9236 | 0.9403 | FALSE | 2.498 | -0.25 | 2 | 0.9965 | 0.9397 | -4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | V4 | FALSE | 0 | 3 | 0.9984 | 0.7566 | 3.471 | 4.298 | 0.1183 | 0.7202 | 0.9177 | FALSE | 2.476 | 1 | 6 | 0.9702 | 0.9146 | 4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | II | FALSE | 1 | 0 | 0.9984 | 0.8811 | 6.472 | 10.116 | 0.0273 | 0.8014 | 0.8632 | FALSE | 2.494 | 0.25 | 0 | 0.9541 | 0.8568 | 4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | aVL | FALSE | 1 | 1 | 0.9976 | 0.9734 | 13.183 | 18.241 | 0.038 | 0.9387 | 0.8947 | FALSE | 2.5 | -0.25 | -6 | 0.9934 | 0.8926 | 4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | V2 | FALSE | 1 | 2 | 0.9912 | 0.9871 | 15.671 | 19.808 | 0.0372 | 0.9639 | 0.9542 | FALSE | 2.496 | 0 | 4 | 0.9957 | 0.9543 | -4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | V5 | FALSE | 1 | 3 | 0.9936 | 0.8298 | 5.052 | 6.535 | 0.0727 | 0.7863 | 0.8637 | FALSE | 2.482 | 0.75 | 4 | 0.98 | 0.8582 | 4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | III | FALSE | 2 | 0 | 0.9936 | 0.9579 | 11.245 | 16.958 | 0.0508 | 0.926 | 0.9517 | FALSE | 2.498 | -0.25 | -8 | 0.9964 | 0.9485 | 4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | aVF | FALSE | 2 | 1 | 0.9992 | 0.9842 | 15.187 | 15.93 | 0.0147 | 0.983 | 0.9645 | FALSE | 2.5 | -0.25 | -4 | 0.9869 | 0.9612 | 2 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | V3 | FALSE | 2 | 2 | 0.996 | 0.954 | 10.806 | 16.754 | 0.0666 | 0.9199 | 0.9245 | FALSE | 2.496 | 0 | 4 | 0.989 | 0.9244 | -4 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | V6 | FALSE | 2 | 3 | 0.9968 | 0.9744 | 12.842 | 12.842 | 0.0202 | 0.8958 | 0.8893 | FALSE | 2.494 | 0.25 | 4 | 0.9878 | 0.872 | 0 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | 20261014_124720 | ptbxl1737_kit_500hz_300dpi | 1737 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1846x1252 | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=88 grid=974 locate=122 calibrate+segment=2 trace=19434 cut=8 quality=0 score*=7 classify=57 | baseline=63.3/43.2 corners=63.3/43.2 rectify=103.3/43.3 grid=109.7/71.4 locate=101.4/40.2 calibrate+segment=73.3/40.3 trace=495.0/58.5 cut=433.0/40.8 quality=433.1/40.8 score*=433.1/42.3 classify=442.2/43.7 | 529.8 | ok |  | II | TRUE | 3 | 0 | 0.9976 | 0.8937 | 6.916 | 6.916 | 0.0262 | 0.8626 | 0.885 |  | 10 | 0 | 0 | 0.8937 | 0.885 | 0 | ok | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | I | FALSE | 0 | 0 | 0.976 | 0.1025 | -2.146 | -1.992 | 0.274 | 0.0924 | -0.0323 | FALSE | 2.336 | 7 | 52 | 0.1527 | 0.1971 | -4 | UNRELIABLE | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | aVR | FALSE | 0 | 1 | 0.9448 | 0.4733 | 1.02 | 1.038 | 0.1498 | 0.2231 | -0.0182 | FALSE | 2.5 | -2.5 | -32 | 0.4916 | -0.0182 | 2 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | V1 | FALSE | 0 | 2 | 0.9608 | 0.3488 | 0.369 | 0.525 | 0.3575 | 0.1012 | -0.0027 | TRUE | 2.376 | 5.25 | 100 | 0.4634 | 0.0314 | -28 | UNRELIABLE | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | V4 | FALSE | 0 | 3 | 0.4176 | 0.0373 | -5.57 | -5.371 | 0.5169 | 0.0448 | -0.0129 | FALSE | 2.364 | 5.75 | 0 | 0.0718 | 0 |  | UNRELIABLE | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | II | FALSE | 1 | 0 | 0.9808 | 0.7725 | 3.606 | 3.65 | 0.0819 | 0.4482 | 0.3251 | FALSE | 2.5 | -0.5 | -6 | 0.7868 | 0.325 | 4 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | aVL | FALSE | 1 | 1 | 0.9264 | 0.2839 | -0.016 | 0.412 | 0.2063 | 0.107 | 0.0289 | TRUE | 2.4 | -2.5 | 100 | 0.5471 | 0.1684 | -134 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | V2 | FALSE | 1 | 2 | 0.9616 | 0.4719 | 0.547 | 0.597 | 0.4297 | 0.1485 | 0.0118 | FALSE | 2.444 | 2.25 | 38 | 0.5033 | 0.0122 | -18 | UNRELIABLE | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | V5 | FALSE | 1 | 3 | 0.98 | 0.4303 | -4.126 | -3.739 | 0.1829 | 0.6966 | 0.907 | FALSE | 2.47 | 1 | -6 | 0.5036 | 0.9054 | 22 | UNRELIABLE | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | III | FALSE | 2 | 0 | 0.9776 | 0.4104 | 0.611 | 0.829 | 0.1655 | 0.1321 | 0.0386 | FALSE | 2.5 | -5.5 | -30 | 0.4738 | 0.0603 | -62 | UNRELIABLE | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | aVF | FALSE | 2 | 1 | 0.964 | 0.4308 | 0.828 | 0.962 | 0.1033 | 0.2135 | 0.0388 | FALSE | 2.422 | 3.25 | 74 | 0.5599 | 0.2518 | -40 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | V3 | FALSE | 2 | 2 | 0.9616 | 0.453 | 0.901 | 1.103 | 0.3131 | 0.2008 | 0.0608 | TRUE | 2.4 | 0.25 | 100 | 0.5208 | 0.0875 | -108 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | V6 | FALSE | 2 | 3 | 0.9384 | 0.7413 | 2.451 | 3.185 | 0.0699 | 0.7996 | 0.7931 | FALSE | 2.45 | 2 | 10 | 0.9518 | 0.9248 | 8 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | 20261014_125029 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=112 grid=966 locate=160 calibrate+segment=2 trace=191 cut=7 quality=0 score*=7 classify=56 | baseline=73.2/45.5 corners=73.2/45.5 rectify=113.2/45.7 grid=119.8/63.8 locate=111.7/36.3 calibrate+segment=83.3/36.4 trace=86.0/48.4 cut=83.5/43.3 quality=83.5/43.3 score*=83.5/44.7 classify=91.6/46.0 | 182.4 | check | LEAD_COVERAGE/warn/V4 | II | TRUE | 3 | 0 | 0.8366 | 0.4221 | 0.033 | 0.033 | 0.1337 | 0.3512 | -0.0697 |  | 9.094 | 1.75 | -734 | 0.6171 | 0.1236 |  | UNRELIABLE | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | I | FALSE | 0 | 0 | 0.9992 | 0.9635 | 11.427 | 11.427 | 0.0574 | 0.9062 | 0.9464 | FALSE | 2.482 | 0.75 | 8 | 0.9755 | 0.9425 | 2 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | aVR | FALSE | 0 | 1 | 0.9976 | 0.9645 | 11.646 | 13.376 | 0.0441 | 0.9416 | 0.9375 | FALSE | 2.496 | 0 | -4 | 0.9762 | 0.9376 | 4 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | V1 | FALSE | 0 | 2 | 0.9976 | 0.988 | 16.199 | 19.655 | 0.0578 | 0.9713 | 0.9857 | FALSE | 2.498 | -0.25 | 2 | 0.9958 | 0.9857 | -4 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | V4 | FALSE | 0 | 3 | 0.996 | 0.9191 | 8.03 | 10.619 | 0.108 | 0.9194 | 1.0081 | FALSE | 2.464 | 1.5 | 4 | 0.9939 | 1.0073 | 10 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | II | FALSE | 1 | 0 | 0.9992 | 0.9333 | 8.879 | 8.879 | 0.0446 | 0.8667 | 0.9344 | FALSE | 2.482 | 0.75 | 8 | 0.9697 | 0.9246 | 0 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | aVL | FALSE | 1 | 1 | 0.9976 | 0.9894 | 15.71 | 16.588 | 0.0337 | 0.9179 | 0.9172 | FALSE | 2.498 | 0 | -2 | 0.9921 | 0.9172 | 2 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | V2 | FALSE | 1 | 2 | 0.9968 | 0.9863 | 15.7 | 20.62 | 0.0751 | 0.9705 | 0.977 | FALSE | 2.498 | -0.25 | 2 | 0.9974 | 0.9765 | -6 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | V5 | FALSE | 1 | 3 | 0.9992 | 0.8696 | 5.917 | 7.069 | 0.0575 | 0.8738 | 1.0004 | FALSE | 2.464 | 1.5 | 6 | 0.9834 | 0.9946 | 8 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | III | FALSE | 2 | 0 | 0.9912 | 0.9894 | 16.862 | 19.294 | 0.0255 | 0.9551 | 0.9668 | FALSE | 2.5 | -0.25 | -6 | 0.9952 | 0.9634 | 2 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | aVF | FALSE | 2 | 1 | 0.9992 | 0.9713 | 12.488 | 13.37 | 0.027 | 0.9065 | 0.932 | FALSE | 2.498 | 0 | -2 | 0.9771 | 0.932 | 2 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | V3 | FALSE | 2 | 2 | 0.9968 | 0.9896 | 17.05 | 21.334 | 0.0488 | 0.9845 | 0.9854 | FALSE | 2.498 | -0.25 | 2 | 0.9972 | 0.985 | -4 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | V6 | FALSE | 2 | 3 | 0.9992 | 0.716 | 2.237 | 3.757 | 0.0716 | 0.7454 | 0.8788 | FALSE | 2.458 | 1.75 | 6 | 0.9768 | 0.9278 | 10 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | 20261014_125027 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1980x1439 | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=94 grid=983 locate=160 calibrate+segment=2 trace=19615 cut=8 quality=5 score*=9 classify=87 | baseline=68.7/45.4 corners=68.7/45.4 rectify=108.7/45.4 grid=116.8/63.5 locate=108.1/36.5 calibrate+segment=79.7/36.6 trace=501.6/65.5 cut=438.6/42.4 quality=438.7/42.4 score*=438.8/44.0 classify=447.5/45.4 | 540.4 | ok |  | II | TRUE | 3 | 0 | 0.9998 | 0.8723 | 5.938 | 5.938 | 0.0677 | 0.8703 | 0.9661 |  | 9.156 | -0.25 | 844 | 0.9113 | 0.9862 | 0 | ok | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | 9 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | I | FALSE | 0 | 0 | 0.98 | 0.5836 | 1.871 | 2.01 | 0.1322 | 0.391 | 0.2491 | FALSE | 2.434 | 2.75 | 36 | 0.6422 | 0.2489 | 4 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | aVR | FALSE | 0 | 1 | 0.9624 | 0.5875 | 1.775 | 2.064 | 0.1087 | 0.4851 | 0.211 | FALSE | 2.49 | 0.25 | -4 | 0.6165 | 0.2075 | 6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | V1 | FALSE | 0 | 2 | 0.9624 | 0.4424 | -0.25 | -0.132 | 0.2892 | 0.406 | 0.0592 | FALSE | 2.482 | 0.75 | 18 | 0.4626 | 0.0533 | -10 | UNRELIABLE | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | V4 | FALSE | 0 | 3 | 0.9792 | 0.525 | 1.691 | 1.743 | 0.1475 | 0.3155 | 0.0432 | FALSE | 2.492 | 0 | 8 | 0.5335 | 0.0481 | -8 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | II | FALSE | 1 | 0 | 0.98 | 0.7679 | 3.92 | 4.054 | 0.0683 | 0.6693 | 0.2957 | FALSE | 2.5 | -1.5 | -16 | 0.7956 | 0.2928 | -4 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | aVL | FALSE | 1 | 1 | 0.9392 | 0.5057 | 1.348 | 1.485 | 0.1332 | 0.3062 | 0.2712 | FALSE | 2.458 | 1.75 | 18 | 0.5788 | 0.266 | 6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | V2 | FALSE | 1 | 2 | 0.9624 | 0.8529 | 5.669 | 5.8 | 0.2061 | 0.6256 | 0.0714 | FALSE | 2.482 | 0.75 | 12 | 0.8643 | 0.0713 | -6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | V5 | FALSE | 1 | 3 | 0.9792 | 0.5627 | 1.721 | 1.731 | 0.1329 | 0.2608 | 0.0685 | FALSE | 2.496 | 0 | -4 | 0.566 | 0.0685 | 6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | III | FALSE | 2 | 0 | 0.9544 | 0.8352 | 4.341 | 4.59 | 0.0703 | 0.4946 | 0.333 | FALSE | 2.482 | 0.75 | 14 | 0.8585 | 0.3326 | -6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | aVF | FALSE | 2 | 1 | 0.9624 | 0.6193 | -3.953 | -3.631 | 0.1293 | 1.1989 | 0.9319 | FALSE | 2.494 | 0.25 | 6 | 0.6239 | 0.9319 | -2 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | V3 | FALSE | 2 | 2 | 0.9624 | 0.7661 | 4.238 | 4.768 | 0.1477 | 0.5829 | 0.2771 | FALSE | 2.486 | 0.5 | 14 | 0.8065 | 0.2769 | -6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | V6 | FALSE | 2 | 3 | 0.9792 | 0.4895 | -2.107 | -2.092 | 0.1728 | 0.6681 | 0.6785 | FALSE | 2.45 | 2 | 24 | 0.5279 | 0.6733 | -4 | UNRELIABLE | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | 20261014_124930 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=105 grid=910 locate=156 calibrate+segment=2 trace=188 cut=7 quality=0 score*=7 classify=59 | baseline=71.0/45.1 corners=71.0/45.1 rectify=111.1/45.2 grid=118.6/64.8 locate=110.3/44.9 calibrate+segment=81.8/45.0 trace=84.8/63.1 cut=83.1/30.2 quality=83.1/30.2 score*=83.3/31.6 classify=91.7/33.0 | 197.3 | ok |  | II | TRUE | 3 | 0 | 0.9312 | 0.7643 | 3.725 | 3.725 | 0.0685 | 0.7001 | 0.6475 |  | 9.95 | 0.5 | 30 | 0.841 | 0.6435 | 0 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | I | FALSE | 0 | 0 | 0.9976 | 0.0891 | -4.283 | -4.256 | 0.2684 | 0.1275 | 0.2033 | FALSE | 2.392 | 4.5 | 58 | 0.1192 | 0.1559 | 200 | FAILED | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVR | FALSE | 0 | 1 | 0.8256 | 0.0789 | -3.205 | -3.133 | 0.1929 | 0.0898 | 0.1122 | FALSE | 2.49 | -1.5 | 10 | 0.0951 | 0.1104 |  | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V1 | FALSE | 0 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V4 | FALSE | 0 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | FALSE | 1 | 0 | 0.9984 | 0.2482 | -8.726 | -8.509 | 0.2929 | 0.6991 | 0.635 | FALSE | 2.5 | -5.25 | -78 | 0.2988 | 0.8788 | 28 | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVL | FALSE | 1 | 1 | 0.8296 | -0.1518 | -2.765 | -2.666 | 0.2138 | -0.127 | -0.0644 | TRUE | 2.272 | 10 | 100 | -0.0411 | 0.1033 |  | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V2 | FALSE | 1 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V5 | FALSE | 1 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | III | FALSE | 2 | 0 | 0.9992 | -0.0441 | -5.642 | -4.673 | 0.2218 | -0.0671 | -0.0493 | FALSE | 2.5 | -10 | -44 | 0.2775 | 0.2514 | -124 | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVF | FALSE | 2 | 1 | 0.972 | 0.2626 | -12.387 | -12.213 | 0.3414 | 1.0438 | 0.1344 | FALSE | 2.356 | 5.25 | -20 | 0.3989 | 0.1361 | 94 | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V3 | FALSE | 2 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V6 | FALSE | 2 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | 20261014_124846 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1913x1439 | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=69 grid=866 locate=124 calibrate+segment=2 trace=19428 cut=5 quality=1 score*=4 classify=56 | baseline=76.5/45.2 corners=76.5/45.3 rectify=116.5/45.3 grid=124.4/69.1 locate=115.9/36.6 calibrate+segment=87.6/36.7 trace=509.6/60.9 cut=447.1/41.7 quality=447.1/41.7 score*=447.2/42.6 classify=455.9/43.9 | 543.3 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | TRUE | 3 | 0 | 0.4574 | -0.1681 | -5.717 | -5.524 | 0.203 | -0.2023 | 0 |  | 1.822 | 2.25 | 8178 |  |  |  | UNRELIABLE | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | I | FALSE | 0 | 0 | 0.9936 | 0.945 | 9.621 | 16.054 | 0.0541 | 0.937 | 0.9208 | FALSE | 2.498 | -0.25 | -8 | 0.9899 | 0.9181 | 6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | aVR | FALSE | 0 | 1 | 0.9984 | 0.9805 | 14.259 | 17.705 | 0.0258 | 0.9222 | 0.8989 | FALSE | 2.498 | 0 | -2 | 0.9922 | 0.8989 | 2 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | V1 | FALSE | 0 | 2 | 0.9984 | 0.9626 | 11.296 | 20.631 | 0.0765 | 0.939 | 0.9361 | FALSE | 2.494 | 0 | 6 | 0.9964 | 0.936 | -6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | V4 | FALSE | 0 | 3 | 0.9984 | 0.9778 | 13.718 | 13.718 | 0.0369 | 0.9557 | 0.9144 | FALSE | 2.488 | 0.5 | 6 | 0.9906 | 0.9012 | 0 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | II | FALSE | 1 | 0 | 0.996 | 0.9419 | 9.618 | 16.242 | 0.0354 | 0.8913 | 0.8782 | FALSE | 2.498 | -0.25 | -8 | 0.9894 | 0.8746 | 6 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | aVL | FALSE | 1 | 1 | 0.9984 | 0.9884 | 16.193 | 18.441 | 0.0241 | 0.9371 | 0.8995 | FALSE | 2.498 | 0 | -2 | 0.9937 | 0.8995 | 2 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | V2 | FALSE | 1 | 2 | 0.9984 | 0.9803 | 14.375 | 21.705 | 0.0756 | 0.966 | 0.9109 | FALSE | 2.496 | 0 | 4 | 0.9965 | 0.9103 | -4 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | V5 | FALSE | 1 | 3 | 0.9984 | 0.9587 | 11.115 | 12.129 | 0.0451 | 0.9543 | 0.9462 | FALSE | 2.476 | 1 | 10 | 0.9883 | 0.9342 | 2 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | III | FALSE | 2 | 0 | 0.9984 | 0.9462 | 9.788 | 10.207 | 0.0375 | 0.9534 | 0.8883 | FALSE | 2.5 | -0.5 | -4 | 0.9851 | 0.8829 | 2 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | aVF | FALSE | 2 | 1 | 0.9984 | 0.932 | 7.802 | 8.578 | 0.0334 | 1.0088 | 0.9433 | FALSE | 2.494 | 0.25 | 6 | 0.9585 | 0.9338 | -2 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | V3 | FALSE | 2 | 2 | 0.9984 | 0.946 | 10.207 | 17.032 | 0.0743 | 0.912 | 0.9107 | FALSE | 2.496 | 0 | 4 | 0.9899 | 0.9107 | -4 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | V6 | FALSE | 2 | 3 | 0.9984 | 0.9741 | 13.144 | 13.975 | 0.0299 | 0.9618 | 0.9526 | FALSE | 2.482 | 0.75 | 8 | 0.9971 | 0.9493 | 2 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | 20261014_124927 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1974x1336 | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=73 grid=944 locate=152 calibrate+segment=2 trace=19422 cut=8 quality=0 score*=7 classify=52 | baseline=66.9/44.2 corners=66.9/44.2 rectify=106.9/44.2 grid=115.2/72.8 locate=106.7/42.1 calibrate+segment=78.3/42.2 trace=500.2/65.2 cut=436.6/42.1 quality=436.4/42.2 score*=436.4/43.6 classify=445.2/45.1 | 586.4 | ok |  | II | TRUE | 3 | 0 | 0.9994 | 0.8603 | 5.861 | 6 | 0.0535 | 0.84 | 0.8681 |  | 9.976 | 0.25 | 12 | 0.9561 | 0.8544 | -2 | ok | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | 8 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | I | FALSE | 0 | 0 | 0.976 | 0.8325 | 5.025 | 5.399 | 0.046 | 0.7945 | 0.9243 | FALSE | 2.492 | 0.25 | -2 | 0.8467 | 0.9243 | 6 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | aVR | FALSE | 0 | 1 | 0.9624 | 0.8292 | 4.99 | 5.2 | 0.0541 | 0.7228 | 0.7818 | FALSE | 2.5 | -0.5 | -10 | 0.8395 | 0.7795 | 4 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | V1 | FALSE | 0 | 2 | 0.956 | 0.5898 | 1.359 | 1.387 | 0.1352 | 0.19 | 0.1467 | FALSE | 2.488 | 0.5 | 2 | 0.6009 | 0.1454 | 4 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | V4 | FALSE | 0 | 3 | 0.9696 | 0.5571 | 1.663 | 1.692 | 0.1892 | 0.3796 | -0.0537 | FALSE | 2.47 | 1.25 | 14 | 0.642 | -0.0548 | 2 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | II | FALSE | 1 | 0 | 0.504 | 0.4913 | -0.765 | -0.729 | 0.1154 | 0.4257 | 0.097 | FALSE | 2.376 | 5.25 | 84 | 0.5093 | 0.1 |  | UNRELIABLE | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | aVL | FALSE | 1 | 1 | 0.9584 | 0.782 | 4.012 | 4.012 | 0.0358 | 0.5931 | 0.1759 | FALSE | 2.5 | -0.75 | -14 | 0.8049 | 0.1756 | 0 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | V2 | FALSE | 1 | 2 | 0.8992 | 0.5919 | 1.645 | 1.666 | 0.2662 | 0.2375 | 0.1898 | FALSE | 2.494 | 0.25 | 6 | 0.5964 | 0.189 | -2 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | V5 | FALSE | 1 | 3 | 0.8704 | 0.573 | 1.663 | 1.663 | 0.1411 | 0.3436 | 0.039 | FALSE | 2.47 | 1.25 | 14 | 0.5769 | 0.0383 | 0 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | III | FALSE | 2 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | aVF | FALSE | 2 | 1 | 0.728 | 0.6348 | 2.486 | 2.699 | 0.0654 | 0.3949 | 0.222 | FALSE | 2.49 | 0.25 | 10 | 0.6585 | 0.2158 |  | UNRELIABLE | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | V3 | FALSE | 2 | 2 | 0.9176 | 0.1893 | -5.528 | -5.433 | 0.4858 | 0.3472 | -0.0261 | FALSE | 2.47 | 1.25 | 28 | 0.2162 | -0.0227 | -10 | UNRELIABLE | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | V6 | FALSE | 2 | 3 | 0.96 | 0.7152 | 3.226 | 3.573 | 0.0846 | 0.5763 | 0.4893 | FALSE | 2.482 | 0.75 | 12 | 0.7665 | 0.4886 | -6 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | 20261014_124029 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=134 grid=1104 locate=114 calibrate+segment=2 trace=182 cut=6 quality=0 score*=9 classify=58 | baseline=72.2/36.3 corners=72.2/36.3 rectify=114.3/36.4 grid=119.9/62.6 locate=109.7/46.4 calibrate+segment=80.7/23.4 trace=83.0/44.8 cut=81.2/45.9 quality=81.2/45.9 score*=81.3/45.9 classify=89.9/27.1 | 193.8 | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | II | TRUE | 3 | 0 | 0.7834 | 0.6937 | 2.893 | 2.893 | 0.0853 | 0.5965 | 0.1036 |  | 10 | 0 | 0 | 0.6937 | 0.1036 |  | UNRELIABLE | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | I | FALSE | 0 | 0 | 0.1696 | 0.1883 | -0.321 | -0.176 | 0.0851 | 0.1027 | 0 | FALSE | 2.42 | 0 | 80 | 0.2807 | 0 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVR | FALSE | 0 | 1 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V1 | FALSE | 0 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V4 | FALSE | 0 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | FALSE | 1 | 0 | 0.908 | -0.213 | -15.873 | -15.577 | 0.657 | -1.2731 | 1.5724 | TRUE | 2.172 | 10 | -100 | 0.066 | 1.1742 | 184 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVL | FALSE | 1 | 1 | 0.9984 | 0.1511 | -20.017 | -19.774 | 0.5695 | 1.523 | 0.2802 | FALSE | 2.266 | 10 | -6 | 0.2385 | 0.1535 | 200 | FAILED | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V2 | FALSE | 1 | 2 | 0.6704 | -0.0784 | -6.957 | -6.786 | 0.7167 | -0.1408 | 0.3243 | FALSE | 2.232 | 10 | -40 | -0.0195 | 0.397 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V5 | FALSE | 1 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | III | FALSE | 2 | 0 | 0.9928 | 0.0295 | -20.376 | -19.732 | 0.6831 | 0.3094 | 1.1097 | TRUE | 2.5 | -9.5 | -100 | 0.2335 | 1.5429 | -76 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVF | FALSE | 2 | 1 | 0.9984 | -0.0859 | -12.442 | -12.178 | 0.3648 | -0.3533 | 1.1296 | FALSE | 2.278 | 9.75 | 98 | 0.1323 | 1.2085 | -194 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V3 | FALSE | 2 | 2 | 0.9984 | -0.0375 | -6.948 | -6.85 | 0.5721 | -0.0747 | 0.1015 | TRUE | 2.286 | 4.75 | -100 | 0.0493 | 0.3513 | 140 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V6 | FALSE | 2 | 3 | 0.3616 | -0.1811 | -8.744 | -8.27 | 0.3356 | -0.4066 | 0 | TRUE | 2.172 | 10 | -100 | -0.0053 | 0.2523 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | 20261014_122954 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=73 grid=875 locate=409 calibrate+segment=24 trace=20345 cut=29 quality=9 score*=19 classify=101 | baseline=61.8/43.1 corners=61.8/43.1 rectify=100.9/43.2 grid=116.6/70.6 locate=106.1/39.5 calibrate+segment=79.0/21.5 trace=501.0/59.0 cut=439.4/41.6 quality=439.0/41.6 score*=439.0/42.7 classify=447.8/44.0 | 534.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | TRUE | 3 | 0 | 0.8354 | -0.3106 | -16.358 | -16.328 | 0.7823 | -2.0123 | 0.118 |  | 2.936 | -10 | 7064 |  |  |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | rhythm_locator
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | I | FALSE | 0 | 0 | 0.996 | 0.9234 | 8.193 | 13.203 | 0.0319 | 0.9415 | 0.9366 | FALSE | 2.494 | 0 | -6 | 0.9762 | 0.9379 | 6 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | aVR | FALSE | 0 | 1 | 0.9992 | 0.9758 | 13.28 | 14.73 | 0.0208 | 0.9469 | 0.9092 | FALSE | 2.5 | -0.25 | -6 | 0.9856 | 0.9027 | 2 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | V1 | FALSE | 0 | 2 | 0.996 | 0.908 | 7.55 | 20.314 | 0.0663 | 0.8581 | 0.9493 | FALSE | 2.494 | 0 | 6 | 0.9969 | 0.9489 | -6 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | V4 | FALSE | 0 | 3 | 0.9976 | 0.8674 | 6.046 | 6.046 | 0.1142 | 0.8427 | 0.9041 | FALSE | 2.482 | 0.75 | 8 | 0.9614 | 0.9006 | 0 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | II | FALSE | 1 | 0 | 0.9928 | 0.8722 | 6.241 | 16.37 | 0.0515 | 0.8492 | 0.9399 | FALSE | 2.49 | 0 | -10 | 0.9879 | 0.9403 | 10 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | aVL | FALSE | 1 | 1 | 0.996 | 0.9333 | 8.536 | 9.817 | 0.0213 | 0.8557 | 0.9158 | FALSE | 2.5 | -0.25 | -6 | 0.9564 | 0.9015 | 2 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | V2 | FALSE | 1 | 2 | 0.9976 | 0.9571 | 10.702 | 14.491 | 0.0938 | 0.8805 | 0.8818 | FALSE | 2.498 | 0 | 2 | 0.9841 | 0.8818 | -2 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | V5 | FALSE | 1 | 3 | 0.9976 | 0.9483 | 10.037 | 10.284 | 0.0538 | 0.9416 | 0.957 | FALSE | 2.488 | 0.5 | 4 | 0.9853 | 0.9514 | 2 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | III | FALSE | 2 | 0 | 0.9784 | 0.8099 | 4.42 | 10.315 | 0.0393 | 0.7774 | 0.9214 | FALSE | 2.498 | -0.75 | -20 | 0.9766 | 0.9122 | 10 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | aVF | FALSE | 2 | 1 | 0.996 | 0.9847 | 15.379 | 15.379 | 0.0148 | 0.9906 | 0.9671 | FALSE | 2.5 | 0 | 0 | 0.9847 | 0.9671 | 0 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | V3 | FALSE | 2 | 2 | 0.9976 | 0.9679 | 11.999 | 16.032 | 0.0646 | 0.8941 | 0.8804 | FALSE | 2.498 | 0 | 2 | 0.9892 | 0.8804 | -2 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | V6 | FALSE | 2 | 3 | 0.9928 | 0.9593 | 11.049 | 18.755 | 0.0344 | 0.946 | 0.9524 | FALSE | 2.496 | 0 | 4 | 0.9937 | 0.9524 | -4 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
tilt | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | 20261014_124026 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 1901x1253 | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=109 grid=1071 locate=116 calibrate+segment=2 trace=19433 cut=8 quality=1 score*=8 classify=57 | baseline=69.2/43.1 corners=69.2/43.1 rectify=111.2/43.2 grid=121.1/64.0 locate=107.0/40.9 calibrate+segment=78.1/41.0 trace=500.0/65.2 cut=436.6/34.6 quality=436.6/34.6 score*=436.7/36.1 classify=445.9/37.5 | 531.7 | ok |  | II | TRUE | 3 | 0 | 0.9964 | 0.9057 | 7.353 | 7.461 | 0.051 | 0.9526 | 0.9397 |  | 9.998 | 0 | 2 | 0.908 | 0.9397 | -2 | ok | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | I | FALSE | 0 | 0 | 0.98 | 0.5556 | -1.56 | -1.433 | 0.2561 | 0.7105 | 0.8588 | FALSE | 2.476 | 1 | 6 | 0.5767 | 0.8586 | 6 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVR | FALSE | 0 | 1 | 0.6928 | 0.282 | -3.981 | -3.953 | 0.2664 | 0.3772 | 0.369 | FALSE | 2.476 | 1 | 2 | 0.2986 | 0.3663 |  | UNRELIABLE | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V1 | FALSE | 0 | 2 | 0.964 | 0.3826 | -1.729 | -1.593 | 0.4552 | 0.3674 | 0.1389 | FALSE | 2.5 | -2.25 | -26 | 0.4085 | 0.1389 | -12 | UNRELIABLE | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V4 | FALSE | 0 | 3 | 0.4184 | 0.4388 | -0.371 | -0.366 | 0.2841 | 0.3017 | 0 | FALSE | 2.32 | 7.75 | 18 | 0.4794 | 0 |  | UNRELIABLE | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | FALSE | 1 | 0 | 0.9776 | 0.6581 | 1.577 | 9.364 | 0.1034 | 0.6721 | 0.8571 | FALSE | 2.48 | 0.25 | -14 | 0.9438 | 0.8536 | 16 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVL | FALSE | 1 | 1 | 0.9624 | 0.5419 | -3.29 | -3.253 | 0.3007 | 0.8101 | 0.8733 | FALSE | 2.458 | 1.75 | 30 | 0.5637 | 0.8704 | -4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V2 | FALSE | 1 | 2 | 0.7216 | 0.3346 | -2.284 | -2.274 | 0.5953 | 0.3661 | 0.5335 | FALSE | 2.488 | 0.5 | 6 | 0.3372 | 0.5323 |  | UNRELIABLE | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V5 | FALSE | 1 | 3 | 0.48 | 0.5802 | 1.705 | 1.875 | 0.0935 | 0.3065 | 0.0679 | FALSE | 2.496 | -1.5 | 4 | 0.6068 | 0.0703 |  | UNRELIABLE | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | III | FALSE | 2 | 0 | 0.98 | 0.9569 | 10.811 | 11.557 | 0.0511 | 0.8897 | 0.8597 | FALSE | 2.5 | -0.5 | -2 | 0.9678 | 0.8582 | -4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVF | FALSE | 2 | 1 | 0.9624 | 0.6948 | 0.98 | 1.125 | 0.1015 | 0.8578 | 0.8491 | FALSE | 2.47 | 1.25 | 16 | 0.7204 | 0.8483 | 4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V3 | FALSE | 2 | 2 | 0.964 | 0.9844 | 15.304 | 19.683 | 0.0596 | 0.949 | 0.9697 | FALSE | 2.492 | 0.25 | 8 | 0.9953 | 0.9697 | -6 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V6 | FALSE | 2 | 3 | 0.48 | 0.4374 | 0.735 | 0.979 | 0.0851 | 0.2676 | 0 | FALSE | 2.464 | 1.5 | 14 | 0.4951 | 0.6016 |  | UNRELIABLE | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | 20261014_231734 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1334 locate=232 calibrate+segment=3 trace=225 cut=7 quality=0 score*=5 classify=57 | baseline=74.1/46.5 corners=74.1/46.5 rectify=120.2/46.7 grid=133.9/77.8 locate=115.5/73.0 calibrate+segment=87.1/73.0 trace=91.0/54.6 cut=87.9/27.5 quality=88.0/27.5 score*=88.0/28.7 classify=96.3/30.1 | 203.8 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | TRUE | 3 | 0 | 0.8988 | 0.5824 | 1.202 | 1.325 | 0.1168 | 0.5138 | 0.8221 |  | 9.9 | 1 | 56 | 0.8313 | 0.835 | -8 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | I | FALSE | 0 | 0 | 0.0416 | 0.1021 | 0.047 | 0.049 | 0.2128 | 0.008 | 0 | FALSE | 2.444 | 2.25 | 48 | 0.1134 | 0 |  | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | aVR | FALSE | 0 | 1 | 0.8464 | 0.7169 | 3.111 | 3.13 | 0.1178 | 0.4881 | 0.4089 | FALSE | 2.398 | 4.25 | 68 | 0.8267 | 0.6667 |  | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | V1 | FALSE | 0 | 2 | 0.96 | 0.7662 | 3.794 | 4.066 | 0.241 | 0.6528 | 0.9085 | FALSE | 2.49 | 0.25 | 10 | 0.7816 | 0.907 | -8 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | V4 | FALSE | 0 | 3 | 0.9784 | 0.3207 | 0.377 | 0.442 | 0.2607 | 0.1632 | 0.2785 | FALSE | 2.482 | 0.75 | 0 | 0.3425 | 0.2763 | 8 | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | II | FALSE | 1 | 0 | 0.12 | 0.2764 | 0.314 | 0.319 | 0.1196 | 0.0467 | 0 | FALSE | 2.434 | 2.75 | 60 | 0.294 | 0 |  | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | aVL | FALSE | 1 | 1 | 0.9552 | 0.5009 | -3.496 | -3.477 | 0.3079 | 0.7538 | 0.8622 | FALSE | 2.416 | 3.5 | 58 | 0.5254 | 0.8611 | -4 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | V2 | FALSE | 1 | 2 | 0.9592 | 0.6525 | 2.176 | 2.267 | 0.3562 | 0.3871 | 0.0282 | FALSE | 2.498 | -2.5 | 2 | 0.6813 | 0.0528 | -8 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | V5 | FALSE | 1 | 3 | 0.98 | 0.5681 | 1.484 | 1.522 | 0.0959 | 0.2517 | 0.1206 | FALSE | 2.458 | 1.75 | 10 | 0.5896 | 0.1441 | 12 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | III | FALSE | 2 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | aVF | FALSE | 2 | 1 | 0.6096 | 0.3851 | -5.756 | -5.749 | 0.2205 | 0.6995 | 0.8045 | FALSE | 2.47 | 1.25 | 22 | 0.3955 | 0.8031 |  | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | V3 | FALSE | 2 | 2 | 0.9568 | 0.5498 | 1.644 | 1.949 | 0.2874 | 0.3473 | 0.1739 | FALSE | 2.484 | -1 | 16 | 0.5899 | 0.2035 | -20 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | V6 | FALSE | 2 | 3 | 0.9784 | 0.5576 | 1.318 | 1.318 | 0.0796 | 0.4349 | 0.1463 | FALSE | 2.422 | 3.25 | 44 | 0.6093 | 0.1354 | 0 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | 20261014_232544 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=159 grid=1358 locate=222 calibrate+segment=3 trace=248 cut=8 quality=0 score*=5 classify=56 | baseline=75.3/43.9 corners=75.3/43.9 rectify=121.3/44.1 grid=130.5/76.4 locate=114.9/65.3 calibrate+segment=86.6/65.4 trace=91.0/72.9 cut=88.3/48.8 quality=88.3/24.5 score*=88.3/25.7 classify=96.8/27.0 | 203.7 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | II | TRUE | 3 | 0 | 0.849 | 0.3139 | -0.602 | -0.485 | 0.1438 | 0.2534 | 0.2084 |  | 8.236 | 1.5 | -1616 |  |  |  | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | I | FALSE | 0 | 0 | 0.9984 | 0.9321 | 8.827 | 14.513 | 0.0775 | 0.8974 | 0.9709 | FALSE | 2.484 | 0.5 | -4 | 0.9864 | 0.9425 | 12 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | aVR | FALSE | 0 | 1 | 0.9976 | 0.9838 | 14.966 | 21.434 | 0.0301 | 0.9473 | 0.9621 | FALSE | 2.496 | 0 | -4 | 0.9969 | 0.9622 | 4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | V1 | FALSE | 0 | 2 | 0.9984 | 0.9863 | 15.582 | 17.62 | 0.062 | 0.9688 | 0.9772 | FALSE | 2.498 | -0.25 | 2 | 0.9921 | 0.9761 | -4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | V4 | FALSE | 0 | 3 | 0.9928 | 0.9746 | 13.12 | 19.163 | 0.0601 | 0.9539 | 0.9931 | FALSE | 2.496 | -0.25 | 4 | 0.9965 | 0.9923 | -6 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | II | FALSE | 1 | 0 | 0.9976 | 0.6152 | -0.525 | 1.507 | 0.1317 | 0.8216 | 0.9513 | FALSE | 2.5 | -2.25 | -52 | 0.8652 | 0.9543 | 10 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | aVL | FALSE | 1 | 1 | 0.9976 | 0.9913 | 17.581 | 22.4 | 0.0272 | 0.9685 | 0.9479 | FALSE | 2.496 | 0 | -4 | 0.9975 | 0.9476 | 4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | V2 | FALSE | 1 | 2 | 0.9984 | 0.9941 | 19.222 | 24.463 | 0.05 | 0.9707 | 0.9693 | FALSE | 2.498 | -0.25 | 2 | 0.9988 | 0.9687 | -4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | V5 | FALSE | 1 | 3 | 0.9992 | 0.9691 | 12.267 | 18.618 | 0.0277 | 0.9406 | 0.9801 | FALSE | 2.494 | 0 | 6 | 0.9931 | 0.9799 | -6 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | III | FALSE | 2 | 0 | 0.9984 | 0.8712 | 5.734 | 5.883 | 0.0918 | 0.9329 | 0.9868 | FALSE | 2.5 | -1 | -10 | 0.9246 | 0.9831 | 2 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | aVF | FALSE | 2 | 1 | 0.9976 | 0.9586 | 10.991 | 14.135 | 0.0321 | 0.9392 | 0.9426 | FALSE | 2.496 | 0 | -4 | 0.9801 | 0.9426 | 4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | V3 | FALSE | 2 | 2 | 0.9984 | 0.9932 | 18.865 | 26.671 | 0.0396 | 0.9735 | 0.9781 | FALSE | 2.496 | 0 | 4 | 0.999 | 0.9778 | -4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | V6 | FALSE | 2 | 3 | 0.9992 | 0.9738 | 12.66 | 16.776 | 0.0216 | 0.9193 | 0.8878 | FALSE | 2.494 | 0.25 | 6 | 0.9929 | 0.8853 | -4 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | 20261014_231732 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2188x1543 | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=142 grid=1377 locate=237 calibrate+segment=3 trace=19474 cut=9 quality=0 score*=7 classify=70 | baseline=71.0/48.9 corners=71.0/49.0 rectify=117.1/49.0 grid=126.6/73.9 locate=112.3/50.9 calibrate+segment=83.9/23.5 trace=505.5/65.9 cut=442.3/43.1 quality=442.3/43.1 score*=442.3/44.7 classify=451.0/46.1 | 549.6 | unreliable | SCALE_DISAGREE/severe | II | TRUE | 3 | 0 | 0.9996 | 0.6662 | 1.788 | 1.827 | 0.1092 | 0.6608 | 0.9532 |  | 9.926 | 0.75 | 48 | 0.9378 | 0.9586 | -2 | ok | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | 8 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | I | FALSE | 0 | 0 | 0.9984 | 0.9426 | 9.534 | 17.483 | 0.0714 | 0.9215 | 0.9709 | FALSE | 2.486 | 0.25 | -8 | 0.9921 | 0.9665 | 10 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | aVR | FALSE | 0 | 1 | 0.9968 | 0.9904 | 17.085 | 21.101 | 0.0236 | 0.9544 | 0.9694 | FALSE | 2.498 | 0 | -2 | 0.9967 | 0.9694 | 2 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V1 | FALSE | 0 | 2 | 0.9984 | 0.9892 | 16.631 | 21.976 | 0.055 | 0.9671 | 0.9803 | FALSE | 2.498 | -0.25 | 2 | 0.9987 | 0.9801 | -4 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V4 | FALSE | 0 | 3 | 0.9904 | 0.9805 | 14.247 | 16.473 | 0.0528 | 0.9602 | 0.9913 | FALSE | 2.488 | 0.5 | 0 | 0.9954 | 0.9902 | 4 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | II | FALSE | 1 | 0 | 0.9928 | 0.5153 | -0.698 | 8.541 | 0.1344 | 0.6108 | 0.9253 | FALSE | 2.486 | -0.5 | -26 | 0.9547 | 0.9658 | 20 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | aVL | FALSE | 1 | 1 | 0.9968 | 0.9948 | 19.402 | 22.156 | 0.0221 | 0.9736 | 0.9491 | FALSE | 2.494 | 0.25 | 0 | 0.998 | 0.9426 | 2 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V2 | FALSE | 1 | 2 | 0.9912 | 0.9854 | 14.185 | 15.11 | 0.0894 | 0.8819 | 0.9203 | FALSE | 2.498 | -0.25 | 2 | 0.9897 | 0.9195 | -4 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V5 | FALSE | 1 | 3 | 0.5544 | 0.6748 | 1.344 | 1.534 | 0.0974 | 0.6258 | 0.801 | FALSE | 2.482 | 0.75 | 12 | 0.6977 | 0.7997 |  | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | III | FALSE | 2 | 0 | 0.9776 | 0.91 | 6.849 | 11.592 | 0.0807 | 1.0147 | 0.9647 | FALSE | 2.496 | -0.5 | -16 | 0.9761 | 0.959 | 10 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | aVF | FALSE | 2 | 1 | 0.912 | -0.0047 | -13.874 | -13.71 | 0.5615 | -0.0224 | 1.4207 | FALSE | 2.5 | -4.75 | -6 | 0.1901 | 1.5961 | -44 | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V3 | FALSE | 2 | 2 | 0.7688 | 0.1927 | -3.224 | -3.208 | 0.5035 | 0.2453 | 0.8152 | FALSE | 2.5 | -1.5 | -18 | 0.2008 | 0.8136 |  | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | V6 | FALSE | 2 | 3 | 0.7656 | 0.1917 | -6.49 | -6.194 | 0.1956 | 0.3944 | 0.6144 | FALSE | 2.488 | 0 | 12 | 0.2585 | 0.6221 |  | UNRELIABLE | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | 20261014_232541 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2250x1536 | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=139 grid=1380 locate=232 calibrate+segment=3 trace=19653 cut=7 quality=0 score*=6 classify=52 | baseline=106.1/46.2 corners=106.1/46.2 rectify=152.1/46.3 grid=158.5/72.1 locate=111.9/54.5 calibrate+segment=83.4/15.8 trace=505.3/66.0 cut=443.7/42.9 quality=443.8/42.9 score*=443.8/44.3 classify=451.8/45.7 | 549.6 | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | II | TRUE | 3 | 0 | 0.942 | 0.5644 | -0.369 | -0.029 | 0.14 | 0.6789 | 0.6584 |  | 9.992 | 0 | 8 | 0.5991 | 0.6584 | -8 | ok | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | 9 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | I | FALSE | 0 | 0 | 0.3008 | 0.0616 | -0.001 | 0.06 | 0.1639 | 0.0123 | 0 | FALSE | 2.184 | 10 | -88 | 0.3267 | 0.0672 |  | UNRELIABLE | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | aVR | FALSE | 0 | 1 | 0.8744 | 0.1989 | -6.248 | -6.243 | 0.2739 | 0.3791 | 0.3936 | FALSE | 2.5 | -3.25 | -56 | 0.227 | 0.3861 | -4 | UNRELIABLE | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | V1 | FALSE | 0 | 2 | 0.9624 | 0.9164 | 7.909 | 12.069 | 0.113 | 0.8021 | 0.8703 | FALSE | 2.492 | 0 | 8 | 0.9738 | 0.8703 | -8 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | V4 | FALSE | 0 | 3 | 0.9784 | 0.8611 | 6.13 | 9.657 | 0.0885 | 0.8104 | 0.7534 | FALSE | 2.47 | 1.25 | 6 | 0.9799 | 0.7531 | 8 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | II | FALSE | 1 | 0 | 0.444 | 0.472 | 0.712 | 0.749 | 0.0988 | 0.2299 | 0.0538 | FALSE | 2.37 | 5.5 | 98 | 0.4916 | 0.0538 |  | UNRELIABLE | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | aVL | FALSE | 1 | 1 | 0.9624 | 0.7217 | 3.234 | 3.27 | 0.1072 | 0.4771 | 0.2815 | FALSE | 2.476 | 1 | 20 | 0.7372 | 0.2818 | -2 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | V2 | FALSE | 1 | 2 | 0.9008 | 0.5117 | -0.172 | -0.069 | 0.4037 | 0.5513 | 0.7623 | FALSE | 2.404 | 4 | 30 | 0.6107 | 0.6922 | -6 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | V5 | FALSE | 1 | 3 | 0.9784 | 0.8031 | 4.604 | 7.38 | 0.0954 | 0.756 | 0.8405 | FALSE | 2.458 | 1.75 | 12 | 0.989 | 0.8374 | 10 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | III | FALSE | 2 | 0 | 0.9792 | 0.3653 | 0.682 | 0.831 | 0.1071 | 0.1558 | 0.1273 | TRUE | 2.38 | 5 | 100 | 0.4957 | 0.1647 | -12 | UNRELIABLE | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | aVF | FALSE | 2 | 1 | 0.8944 | 0.7702 | 3.881 | 4.473 | 0.0525 | 0.6844 | 0.8172 | FALSE | 2.486 | 0.5 | 14 | 0.8053 | 0.8139 | -6 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | V3 | FALSE | 2 | 2 | 0.9608 | 0.7999 | 4.873 | 5.913 | 0.1373 | 0.6312 | 0.3224 | FALSE | 2.486 | 0.5 | 14 | 0.8532 | 0.3595 | -8 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | V6 | FALSE | 2 | 3 | 0.9784 | 0.8994 | 7.355 | 9.893 | 0.0581 | 0.8673 | 0.8456 | FALSE | 2.47 | 1.25 | 8 | 0.9848 | 0.8438 | 6 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | 20261014_235400 | ptbxl40_kit_500hz_300dpi | 40 | classical | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=217 grid=1451 locate=269 calibrate+segment=4 trace=337 cut=7 quality=1 score*=6 classify=55 | baseline=69.1/46.1 corners=69.1/46.1 rectify=115.3/46.3 grid=124.0/80.4 locate=109.5/62.9 calibrate+segment=81.1/16.5 trace=85.4/38.8 cut=82.5/28.6 quality=82.5/28.6 score*=82.5/29.9 classify=91.4/31.3 | 203.4 | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | II | TRUE | 3 | 0 | 0.9486 | 0.7455 | 2.796 | 2.829 | 0.0762 | 0.7705 | 0.5973 |  | 9.95 | 0.5 | 24 | 0.8076 | 0.5879 | 2 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | I | FALSE | 0 | 0 | 0.0432 | 0.1036 | -4.464 | -4.419 | 0.2741 | 0.1454 | 0 | FALSE | 2.426 | -10 | 74 | 0.1234 | 0 |  | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | aVR | FALSE | 0 | 1 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | V1 | FALSE | 0 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | V4 | FALSE | 0 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | II | FALSE | 1 | 0 | 0.988 | 0.141 | -4.843 | -4.821 | 0.1873 | 0.2283 | -0.0337 | FALSE | 2.486 | -10 | 14 | 0.29 | 0.099 | -132 | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | aVL | FALSE | 1 | 1 | 0.9992 | -0.0316 | -1.804 | -1.276 | 0.1914 | -0.0213 | -0.0073 | FALSE | 2.424 | -7.75 | 76 | 0.2591 | 0.0744 | -164 | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | V2 | FALSE | 1 | 2 | 0.9976 | -0.0151 | -1.589 | -1.579 | 0.4752 | -0.0097 | -0.0102 | FALSE | 2.456 | -10 | 44 | 0.4302 | 0.0755 | -200 | FAILED | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | V5 | FALSE | 1 | 3 | 0.9976 | 0.1717 | -3.853 | -3.837 | 0.2524 | 0.2506 | 0.3607 | TRUE | 2.4 | -8.25 | 100 | 0.4896 | 0.5941 | 6 | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | III | FALSE | 2 | 0 | 0.9808 | 0.0655 | -3.483 | -2.952 | 0.173 | 0.0753 | 0.0942 | TRUE | 2.334 | 2.75 | -100 | 0.1418 | 0.0661 | 152 | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | aVF | FALSE | 2 | 1 | 0.9992 | -0.0746 | -5.384 | -5.263 | 0.1524 | -0.1093 | 0.0032 | TRUE | 2.5 | -9.75 | -100 | 0.0349 | 0.0609 | 200 | FAILED | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | V3 | FALSE | 2 | 2 | 0.9976 | 0.1023 | -4.213 | -4.132 | 0.3909 | 0.1535 | 0.8649 | FALSE | 2.438 | -10 | 62 | 0.3496 | 0.0815 | -196 | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | V6 | FALSE | 2 | 3 | 0.9976 | 0.0997 | -6.272 | -5.713 | 0.2792 | 0.1966 | 0.7365 | FALSE | 2.32 | 7.75 | 98 | 0.3743 | 0.7415 | -44 | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | 20261014_234838 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2632x1831 | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | corners=0 rectify=188 grid=1724 locate=208 calibrate+segment=5 trace=19734 cut=7 quality=0 score*=5 classify=57 | baseline=104.9/48.6 corners=104.9/48.6 rectify=156.9/48.7 grid=176.4/85.4 locate=115.9/78.3 calibrate+segment=87.5/78.5 trace=509.3/78.6 cut=446.9/44.1 quality=446.9/44.1 score*=447.1/45.3 classify=456.1/46.7 | 547.8 | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | II | TRUE | 3 | 0 | 0.9424 | 0.0477 | -4.62 | -4.601 | 0.1789 | 0.0707 | 0.1787 |  | 9.256 | 2.25 | -524 | 0.5926 | 0.3124 | -32 | UNRELIABLE | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | 1 | rhythm_locator
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | I | FALSE | 0 | 0 | 0.9968 | 0.9784 | 13.864 | 16.256 | 0.0332 | 0.9486 | 0.9422 | FALSE | 2.488 | 0.5 | 4 | 0.9973 | 0.932 | 2 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | aVR | FALSE | 0 | 1 | 0.9984 | 0.8532 | 5.781 | 11.743 | 0.0686 | 0.817 | 0.92 | FALSE | 2.498 | -0.5 | -14 | 0.9918 | 0.9171 | 8 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | V1 | FALSE | 0 | 2 | 0.9992 | 0.9342 | 8.874 | 19.399 | 0.1011 | 0.9179 | 0.9543 | FALSE | 2.494 | 0 | 6 | 0.9945 | 0.9543 | -6 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | V4 | FALSE | 0 | 3 | 0.9976 | 0.8225 | 4.851 | 9.931 | 0.1025 | 0.8249 | 0.9273 | FALSE | 2.47 | 1.25 | 6 | 0.996 | 0.9152 | 10 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | II | FALSE | 1 | 0 | 0.9984 | 0.9521 | 10.431 | 16.033 | 0.0323 | 0.9221 | 0.8728 | FALSE | 2.492 | 0.25 | -2 | 0.9933 | 0.8697 | 6 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | aVL | FALSE | 1 | 1 | 0.9992 | 0.8812 | 6.561 | 13.094 | 0.0731 | 0.8358 | 0.8969 | FALSE | 2.498 | -0.5 | -14 | 0.995 | 0.8823 | 8 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | V2 | FALSE | 1 | 2 | 0.9976 | 0.9812 | 14.563 | 22.073 | 0.074 | 0.9668 | 0.9202 | FALSE | 2.496 | 0 | 4 | 0.9968 | 0.9198 | -4 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | V5 | FALSE | 1 | 3 | 0.9976 | 0.793 | 3.941 | 8.24 | 0.1029 | 0.8318 | 0.9406 | FALSE | 2.464 | 1.5 | 8 | 0.9779 | 0.9358 | 10 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | III | FALSE | 2 | 0 | 0.9968 | 0.9894 | 16.713 | 19.116 | 0.0169 | 0.9481 | 0.8847 | FALSE | 2.498 | 0 | -2 | 0.9943 | 0.8842 | 2 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | aVF | FALSE | 2 | 1 | 0.9992 | 0.9636 | 11.596 | 12.893 | 0.0216 | 0.9327 | 0.9114 | FALSE | 2.5 | -0.25 | -4 | 0.9854 | 0.9117 | 2 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | V3 | FALSE | 2 | 2 | 0.9864 | 0.9373 | 9.588 | 18.253 | 0.0798 | 0.9101 | 0.9028 | FALSE | 2.494 | 0 | 6 | 0.992 | 0.903 | -6 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | V6 | FALSE | 2 | 3 | 0.9976 | 0.8729 | 6.083 | 9.942 | 0.0673 | 0.8959 | 0.9695 | FALSE | 2.47 | 1.25 | 6 | 0.9933 | 0.9685 | 8 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
shadow | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | 20261014_235356 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | user | 2306x1573 | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | corners=0 rectify=157 grid=1431 locate=255 calibrate+segment=3 trace=19940 cut=35 quality=10 score*=28 classify=94 | baseline=58.9/29.8 corners=58.9/29.8 rectify=105.1/29.8 grid=118.7/61.2 locate=104.6/64.9 calibrate+segment=76.1/18.7 trace=498.0/58.9 cut=435.5/42.1 quality=435.7/42.5 score*=436.9/44.0 classify=444.2/45.1 | 534.7 | ok |  | II | TRUE | 3 | 0 | 0.9988 | 0.8204 | 4.158 | 4.203 | 0.0651 | 0.8949 | 0.8779 |  | 9.976 | 0.25 | 8 | 0.8634 | 0.8616 | 2 | ok | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | 7 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | I | FALSE | 0 | 0 | 0.8664 | 0.2489 | 0.099 | 0.38 | 0.2116 | 0.0999 | -0.0381 | FALSE | 2.434 | -5.5 | 66 | 0.3995 | 0.1306 | -184 | UNRELIABLE | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | aVR | FALSE | 0 | 1 | 0.8664 | 0.069 | -1.507 | -1.342 | 0.2004 | 0.0505 | -0.0109 | TRUE | 2.398 | 4.25 | 100 | 0.147 | -0.0385 | -42 | UNRELIABLE | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | V1 | FALSE | 0 | 2 | 0.9632 | 0.3215 | -1.343 | -1.325 | 0.4354 | 0.2965 | 0.4751 | FALSE | 2.5 | -2 | -14 | 0.3335 | 0.4734 | -6 | UNRELIABLE | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | V4 | FALSE | 0 | 3 | 0.9248 | 0.073 | -5.25 | -5.06 | 0.4983 | 0.0883 | 0.3596 | FALSE | 2.5 | -2.75 | -58 | 0.1229 | 0.3584 | 24 | UNRELIABLE | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | II | FALSE | 1 | 0 | 0.9744 | 0.2934 | -5.682 | -5.578 | 0.2385 | 0.515 | -0.2541 | TRUE | 2.464 | -2.5 | -100 | 0.306 | 0.0367 | -6 | UNRELIABLE | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | aVL | FALSE | 1 | 1 | 0.9264 | 0.3928 | 0.417 | 0.534 | 0.1962 | 0.1368 | 0.0409 | TRUE | 2.37 | 5.5 | 100 | 0.5183 | 0.0594 | -22 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | V2 | FALSE | 1 | 2 | 0.952 | 0.4977 | 1.024 | 1.178 | 0.4068 | 0.3274 | 0.0144 | FALSE | 2.478 | 0.25 | 22 | 0.5192 | 0.0252 | -6 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | V5 | FALSE | 1 | 3 | 0.98 | 0.8894 | 6.769 | 8.223 | 0.0522 | 0.8642 | 0.9397 | FALSE | 2.49 | 0.25 | -4 | 0.9213 | 0.9533 | 8 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | III | FALSE | 2 | 0 | 0.9392 | 0.4856 | 1.089 | 1.112 | 0.1566 | 0.3311 | -0.0191 | FALSE | 2.326 | 7.5 | 60 | 0.5868 | 0.3479 | 2 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | aVF | FALSE | 2 | 1 | 0.964 | 0.5216 | -1.09 | -0.938 | 0.1289 | 0.6733 | 0.868 | FALSE | 2.404 | 4 | 72 | 0.6302 | 0.8696 | -6 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | V3 | FALSE | 2 | 2 | 0.96 | 0.8874 | 6.741 | 7.694 | 0.1599 | 0.7223 | 0.8856 | FALSE | 2.49 | 0 | 10 | 0.9146 | 0.8854 | -10 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | V6 | FALSE | 2 | 3 | 0.972 | 0.7645 | 2.485 | 2.792 | 0.0696 | 0.8697 | 0.8821 | FALSE | 2.498 | -0.5 | 2 | 0.7858 | 0.8773 | -4 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | 20261015_125336 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=202 grid=2541 locate=335 calibrate+segment=4 trace=369 cut=11 quality=0 score*=6 classify=54 | baseline=73.2/51.8 corners=73.3/51.8 rectify=129.4/52.0 grid=151.8/94.1 locate=123.6/55.3 calibrate+segment=95.3/45.8 trace=100.7/70.2 cut=95.9/57.8 quality=95.9/32.4 score*=95.9/33.9 classify=104.0/35.7 | 222.2 | check | NO_PULSE/warn | II | TRUE | 3 | 0 | 0.9954 | 0.4088 | -0.386 | 0.382 | 0.1403 | 0.3758 | 0.5099 |  | 5.728 | -0.5 | 4272 |  |  | -20 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | I | FALSE | 0 | 0 | 0.8456 | 0.2617 | -0.809 | 1.353 | 0.2349 | 0.209 | 0.0322 | FALSE | 2.382 | 2.75 | -52 | 0.7337 | 0.2838 |  | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | aVR | FALSE | 0 | 1 | 0.5328 | 0.0449 | -5.288 | -3.172 | 0.3097 | 0.0683 | 0 | FALSE | 2.426 | -1.25 | 74 | 0.4784 | 0.4671 |  | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | V1 | FALSE | 0 | 2 | 0.9336 | -0.2346 | -2.736 | -2.538 | 0.5111 | -0.1735 | -0.1738 | TRUE | 2.4 | -10 | 100 | 0.328 | 0.0601 | -200 | FAILED | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | V4 | FALSE | 0 | 3 | 0.9808 | 0.0252 | -0.866 | -0.824 | 0.3008 | 0.0128 | 0.0211 | FALSE | 2.5 | -6 | -74 | 0.0579 | 0.038 | 26 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | II | FALSE | 1 | 0 | 0.9816 | -0.1602 | -7.54 | -6.054 | 0.2954 | -0.3183 | 0.1631 | TRUE | 2.316 | 3.5 | -100 | 0.3019 | 0.3479 | 136 | ok | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | aVL | FALSE | 1 | 1 | 0.9616 | 0.1021 | -5.446 | -4.391 | 0.3854 | 0.1751 | 0.7251 | FALSE | 2.5 | -7 | -34 | 0.546 | 0.805 | -38 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | V2 | FALSE | 1 | 2 | 0.8624 | -0.1511 | -2.171 | -2.032 | 0.5876 | -0.0974 | -0.0896 | TRUE | 2.4 | -10 | 100 | 0.3765 | 0.0104 | -200 | FAILED | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | V5 | FALSE | 1 | 3 | 0.9808 | 0.2212 | -7.465 | -7.366 | 0.2686 | 0.458 | 0 | FALSE | 2.408 | 2 | -42 | 0.2487 | 0.005 | 60 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | III | FALSE | 2 | 0 | 0.8824 | 0.0257 | -2.13 | -0.842 | 0.2269 | 0.0224 | 0.1497 | TRUE | 2.394 | 0.25 | -100 | 0.571 | 0.6687 | 158 | ok | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | aVF | FALSE | 2 | 1 | 0.9648 | 0.0108 | -4.885 | -1.227 | 0.1995 | 0.016 | 0.3843 | FALSE | 2.5 | -7.5 | -36 | 0.9551 | 0.8405 | -94 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | V3 | FALSE | 2 | 2 | 0.9664 | -0.0588 | -2.39 | -2.277 | 0.4574 | -0.0503 | -0.0701 | TRUE | 2.4 | -10 | 100 | 0.7298 | 0.9131 | -200 | FAILED | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | V6 | FALSE | 2 | 3 | 0.9792 | 0.2513 | -1.23 | -1.156 | 0.1068 | 0.219 | 0.0628 | FALSE | 2.5 | -6.5 | -32 | 0.2776 | 0.044 | 50 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | 20261015_125433 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=229 grid=2586 locate=354 calibrate+segment=4 trace=455 cut=9 quality=0 score*=8 classify=63 | baseline=73.6/51.2 corners=73.6/51.2 rectify=129.8/51.4 grid=148.6/100.3 locate=124.1/70.8 calibrate+segment=95.9/47.3 trace=102.4/75.9 cut=96.8/31.7 quality=96.8/31.7 score*=96.7/33.1 classify=106.1/34.6 | 226.3 | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | II | TRUE | 3 | 0 | 0.9836 | 0.0844 | -2.308 | -2.308 | 0.175 | 0.0782 | 0.1507 |  | 8.474 | -6.25 | 1526 |  |  | 0 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | I | FALSE | 0 | 0 | 0.8424 | 0.8368 | 5.088 | 5.506 | 0.1191 | 0.7912 | 0.8823 | FALSE | 2.428 | 3 | 42 | 0.9625 | 0.9226 |  | UNRELIABLE | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | aVR | FALSE | 0 | 1 | 0.6712 | 0.7185 | 2.919 | 3.046 | 0.1204 | 0.4524 | 0.4559 | FALSE | 2.398 | 4.25 | 60 | 0.7534 | 0.6562 |  | UNRELIABLE | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | V1 | FALSE | 0 | 2 | 0.9632 | 0.982 | 14.417 | 17.751 | 0.0709 | 0.9426 | 0.972 | FALSE | 2.5 | -0.5 | 0 | 0.995 | 0.9708 | -6 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | V4 | FALSE | 0 | 3 | 0.9808 | 0.8747 | 6.354 | 6.354 | 0.131 | 0.7363 | 0.9687 | FALSE | 2.5 | -1 | -14 | 0.883 | 0.9687 | 0 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | II | FALSE | 1 | 0 | 0.9784 | 0.3228 | -1.387 | 8.879 | 0.1455 | 0.3312 | 0.5565 | FALSE | 2.466 | 0 | -34 | 0.9326 | 0.8729 | 34 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | aVL | FALSE | 1 | 1 | 0.9624 | 0.9545 | 10.283 | 10.427 | 0.063 | 0.9904 | 0.906 | FALSE | 2.476 | 1 | 12 | 0.9694 | 0.9047 | 0 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | V2 | FALSE | 1 | 2 | 0.9632 | 0.9912 | 17.336 | 17.336 | 0.0622 | 0.953 | 0.9654 | FALSE | 2.488 | 0.5 | 4 | 0.9957 | 0.9634 | 0 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | V5 | FALSE | 1 | 3 | 0.9808 | 0.8977 | 7.154 | 7.154 | 0.0499 | 0.8574 | 0.9515 | FALSE | 2.5 | -0.75 | -8 | 0.9178 | 0.9515 | 0 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | III | FALSE | 2 | 0 | 0.9304 | 0.4571 | 0.329 | 4.235 | 0.171 | 0.4051 | 0.7811 | FALSE | 2.494 | -3.5 | -96 | 0.9358 | 0.9261 | 68 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | aVF | FALSE | 2 | 1 | 0.964 | 0.8974 | 7.128 | 9.887 | 0.05 | 0.8377 | 0.8966 | FALSE | 2.494 | 0 | 6 | 0.9476 | 0.8966 | -6 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | V3 | FALSE | 2 | 2 | 0.9576 | 0.958 | 11.047 | 17.928 | 0.0974 | 0.9484 | 0.9729 | FALSE | 2.492 | -0.25 | 8 | 0.9932 | 0.9729 | -10 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | V6 | FALSE | 2 | 3 | 0.9656 | 0.7517 | 3.243 | 5.164 | 0.0638 | 0.7094 | 0.8839 | FALSE | 2.496 | -1 | 4 | 0.8601 | 0.8792 | -10 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | 20261015_125643 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=222 grid=2723 locate=384 calibrate+segment=4 trace=393 cut=11 quality=1 score*=7 classify=57 | baseline=71.8/52.2 corners=71.8/52.2 rectify=127.8/52.4 grid=146.7/102.1 locate=121.8/65.9 calibrate+segment=93.6/49.2 trace=99.7/81.6 cut=94.7/56.7 quality=94.7/56.7 score*=94.7/58.1 classify=103.6/58.2 | 232.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | II | TRUE | 3 | 0 | 0.9914 | 0.4966 | 0.043 | 0.054 | 0.1335 | 0.4884 | 0.976 |  | 3.224 | -4 | 6776 |  |  | 2 | UNRELIABLE | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | I | FALSE | 0 | 0 | 0.9976 | 0.7113 | 2.6 | 3.94 | 0.1586 | 0.6827 | 0.8382 | FALSE | 2.392 | 4.5 | 72 | 0.9755 | 0.9825 | -20 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | aVR | FALSE | 0 | 1 | 0.9976 | 0.9449 | 9.69 | 10.604 | 0.0552 | 0.9422 | 0.943 | FALSE | 2.482 | 0.75 | 4 | 0.9676 | 0.9427 | 4 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | V1 | FALSE | 0 | 2 | 0.9992 | 0.9893 | 16.68 | 18.163 | 0.0547 | 0.9642 | 0.983 | FALSE | 2.5 | -0.25 | 0 | 0.9943 | 0.9818 | -4 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | V4 | FALSE | 0 | 3 | 0.9992 | 0.8914 | 6.77 | 10.43 | 0.1249 | 0.8873 | 0.9914 | FALSE | 2.492 | -0.5 | -20 | 0.959 | 0.9906 | 14 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | II | FALSE | 1 | 0 | 0.9984 | 0.6871 | 2.214 | 6.192 | 0.0961 | 0.6654 | 0.9555 | FALSE | 2.47 | 1.25 | 0 | 0.9425 | 0.948 | 16 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | aVL | FALSE | 1 | 1 | 0.9992 | 0.9879 | 16.062 | 19.614 | 0.0324 | 0.9468 | 0.9609 | FALSE | 2.496 | 0 | -4 | 0.9954 | 0.9603 | 4 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | V2 | FALSE | 1 | 2 | 0.9984 | 0.9873 | 16.023 | 17.451 | 0.0723 | 0.9761 | 0.9858 | FALSE | 2.488 | 0.5 | 2 | 0.9971 | 0.9849 | 4 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | V5 | FALSE | 1 | 3 | 0.9968 | 0.8769 | 6.23 | 8.211 | 0.0555 | 0.8703 | 0.9986 | FALSE | 2.492 | 0 | -8 | 0.9219 | 0.9984 | 8 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | III | FALSE | 2 | 0 | 0.9912 | 0.5255 | 0.363 | 6.034 | 0.1703 | 0.5351 | 0.9423 | FALSE | 2.498 | -2.25 | -60 | 0.95 | 0.9519 | 40 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | aVF | FALSE | 2 | 1 | 1 | 0.9543 | 10.554 | 11.857 | 0.0337 | 0.9174 | 0.9554 | FALSE | 2.498 | 0 | 2 | 0.9665 | 0.9553 | -2 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | V3 | FALSE | 2 | 2 | 0.9984 | 0.9744 | 13.135 | 18.595 | 0.0766 | 0.9801 | 0.9889 | FALSE | 2.496 | -0.25 | 4 | 0.9933 | 0.9879 | -8 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | V6 | FALSE | 2 | 3 | 0.9912 | 0.9187 | 7.528 | 7.936 | 0.0389 | 0.9743 | 0.9595 | FALSE | 2.5 | -0.25 | 0 | 0.9292 | 0.9428 | -2 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | 20261015_125332 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2815x1971 | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | corners=0 rectify=192 grid=2562 locate=340 calibrate+segment=4 trace=21244 cut=15 quality=0 score*=15 classify=79 | baseline=67.9/35.5 corners=67.9/35.5 rectify=124.0/35.5 grid=144.1/79.0 locate=119.3/63.0 calibrate+segment=91.3/53.8 trace=513.3/68.0 cut=450.4/45.2 quality=450.1/45.2 score*=450.4/46.8 classify=459.1/48.3 | 555.9 | check | NO_PULSE/warn | II | TRUE | 3 | 0 | 0.995 | 0.4935 | -0.132 | 0.146 | 0.1362 | 0.502 | 1.0331 |  | 5.736 | -0.5 | 4264 |  |  | -10 | ok | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | 8 | rhythm_locator
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | I | FALSE | 0 | 0 | 0.9552 | 0.0328 | -2.453 | 1.595 | 0.2838 | 0.0301 | 0.0315 | TRUE | 2.438 | -1.5 | -100 | 0.9519 | 0.8836 | 80 | ok | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | aVR | FALSE | 0 | 1 | 0.9976 | 0.302 | -2.885 | -2.17 | 0.2349 | 0.4021 | 0.7083 | FALSE | 2.5 | -6.75 | -32 | 0.9693 | 0.8935 | -82 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | V1 | FALSE | 0 | 2 | 0.9992 | -0.1944 | -3 | -2.83 | 0.5269 | -0.1613 | -0.1679 | TRUE | 2.4 | -10 | 100 | 0.512 | 0.8008 | -200 | FAILED | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | V4 | FALSE | 0 | 3 | 0.812 | 0.015 | -1.776 | -1.72 | 0.334 | 0.0112 | 0.0257 | FALSE | 2.5 | -5.25 | -82 | 0.0385 | 0.0354 |  | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | II | FALSE | 1 | 0 | 0.9536 | -0.1599 | -7.816 | -6.556 | 0.305 | -0.3312 | 0.1743 | TRUE | 2.406 | -0.25 | -100 | 0.2793 | 0.7138 | 138 | ok | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | aVL | FALSE | 1 | 1 | 0.9976 | 0.2177 | -3.523 | -1.047 | 0.3089 | 0.2981 | 0.6408 | FALSE | 2.5 | -7.25 | -36 | 0.9972 | 0.8912 | -80 | ok | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | V2 | FALSE | 1 | 2 | 0.9992 | -0.2393 | -2.862 | -2.662 | 0.6363 | -0.1831 | -0.139 | TRUE | 2.4 | -10 | 100 | 0.6897 | 0.8973 | -200 | FAILED | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | V5 | FALSE | 1 | 3 | 0.9192 | 0.1225 | -1.419 | -1.259 | 0.1339 | 0.096 | 0.0227 | FALSE | 2.5 | -6.5 | -88 | 0.1593 | 0.0416 | 42 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | III | FALSE | 2 | 0 | 0.916 | -0.003 | -2.832 | -1.683 | 0.246 | -0.0031 | 0.1672 | TRUE | 2.394 | 0.25 | -100 | 0.5075 | 0.7977 | 166 | ok | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | aVF | FALSE | 2 | 1 | 0.9976 | 0.0123 | -4.524 | -1.254 | 0.1913 | 0.0171 | 0.409 | FALSE | 2.5 | -7.5 | -38 | 0.9587 | 0.8856 | -96 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | V3 | FALSE | 2 | 2 | 0.9992 | -0.0623 | -2.402 | -2.296 | 0.458 | -0.0534 | -0.0719 | TRUE | 2.4 | -10 | 100 | 0.7513 | 0.9262 | -200 | FAILED | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | V6 | FALSE | 2 | 3 | 0.9952 | 0.3281 | -6.58 | -6.047 | 0.1977 | 0.6739 | 0.0569 | FALSE | 2.5 | -6.75 | -36 | 0.3689 | 0.0448 | -188 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | 20261015_125428 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2863x2009 | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2591 locate=349 calibrate+segment=4 trace=19790 cut=10 quality=0 score*=7 classify=54 | baseline=66.7/27.1 corners=66.7/27.1 rectify=122.6/27.1 grid=142.8/96.7 locate=118.1/62.1 calibrate+segment=89.9/55.4 trace=512.0/79.7 cut=449.5/45.3 quality=449.5/45.3 score*=449.5/46.7 classify=460.4/48.3 | 560.8 | unreliable | NO_TICK_SCALE/severe | II | TRUE | 3 | 0 | 0.9328 | 0.0774 | -2.549 | -2.542 | 0.1799 | 0.0754 | 0.1289 |  | 8.476 | -6.25 | 1524 |  |  | 2 | UNRELIABLE | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | 2 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | I | FALSE | 0 | 0 | 0.9984 | 0.8483 | 5.378 | 5.655 | 0.1152 | 0.8164 | 0.9475 | FALSE | 2.428 | 3 | 42 | 0.9639 | 0.9782 | -6 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | aVR | FALSE | 0 | 1 | 0.9976 | 0.9559 | 10.74 | 10.91 | 0.0489 | 0.9051 | 0.9507 | FALSE | 2.476 | 1 | 8 | 0.9849 | 0.9476 | 2 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | V1 | FALSE | 0 | 2 | 0.9992 | 0.9858 | 15.477 | 17.651 | 0.0628 | 0.9557 | 0.9791 | FALSE | 2.498 | -0.25 | 2 | 0.9942 | 0.978 | -4 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | V4 | FALSE | 0 | 3 | 0.9992 | 0.9302 | 8.709 | 9.526 | 0.0999 | 0.9214 | 0.9877 | FALSE | 2.5 | -1 | -16 | 0.9633 | 0.9868 | 6 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | II | FALSE | 1 | 0 | 0.9984 | 0.296 | -1.645 | 8.824 | 0.1499 | 0.3097 | 0.4705 | FALSE | 2.466 | 0 | -34 | 0.9324 | 0.9495 | 34 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | aVL | FALSE | 1 | 1 | 0.9992 | 0.9944 | 19.417 | 22.081 | 0.022 | 0.967 | 0.9526 | FALSE | 2.494 | 0.25 | 0 | 0.9973 | 0.949 | 2 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | V2 | FALSE | 1 | 2 | 1 | 0.9898 | 16.937 | 18 | 0.0651 | 0.9708 | 0.978 | FALSE | 2.488 | 0.5 | 2 | 0.9966 | 0.9778 | 2 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | V5 | FALSE | 1 | 3 | 0.988 | 0.8965 | 7.032 | 7.088 | 0.0506 | 0.8777 | 0.9944 | FALSE | 2.5 | -0.75 | -8 | 0.918 | 0.9923 | 2 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | III | FALSE | 2 | 0 | 0.9664 | 0.2883 | -1.111 | 3.122 | 0.2018 | 0.2747 | 0.8663 | FALSE | 2.492 | -3.5 | -98 | 0.9351 | 0.962 | 68 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | aVF | FALSE | 2 | 1 | 1 | 0.926 | 8.482 | 10.839 | 0.0428 | 0.9057 | 0.9821 | FALSE | 2.496 | 0 | 4 | 0.9573 | 0.9821 | -4 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | V3 | FALSE | 2 | 2 | 0.9976 | 0.9658 | 11.907 | 17.349 | 0.0882 | 0.9611 | 0.9801 | FALSE | 2.494 | -0.25 | 6 | 0.9925 | 0.9789 | -8 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | V6 | FALSE | 2 | 3 | 0.9784 | 0.8083 | 3.943 | 6.512 | 0.0589 | 0.8456 | 0.9629 | FALSE | 2.498 | -1 | 2 | 0.9326 | 0.956 | -10 | ok | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | 20261015_125638 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2919x2035 | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=207 grid=2661 locate=385 calibrate+segment=4 trace=19753 cut=10 quality=0 score*=9 classify=55 | baseline=102.8/43.8 corners=102.8/43.8 rectify=158.8/43.9 grid=179.3/97.2 locate=117.8/64.2 calibrate+segment=89.8/47.2 trace=511.8/71.4 cut=449.3/45.4 quality=449.4/45.4 score*=449.4/46.9 classify=459.8/48.4 | 557.3 | unreliable | SCALE_DISAGREE/severe | II | TRUE | 3 | 0 | 0.9902 | 0.4656 | -0.392 | -0.306 | 0.1404 | 0.4763 | 1.0264 |  | 3.22 | -3.75 | 6780 |  |  | 4 | UNRELIABLE | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | 8 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | I | FALSE | 0 | 0 | 0.9752 | 0.1889 | 0.024 | 0.037 | 0.1635 | 0.0569 | -0.0114 | TRUE | 2.21 | 8.25 | -100 | 0.3488 | 0.0625 | -166 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | aVR | FALSE | 0 | 1 | 0.9608 | 0.5256 | 1.58 | 1.58 | 0.1112 | 0.248 | 0.1047 | FALSE | 2.5 | 0 | 0 | 0.5256 | 0.1047 | 0 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | V1 | FALSE | 0 | 2 | 0.9616 | 0.494 | 1.078 | 1.123 | 0.2482 | 0.1837 | 0.0336 | FALSE | 2.5 | -4 | -20 | 0.5403 | 0.0288 | -14 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | V4 | FALSE | 0 | 3 | 0.7488 | 0.1115 | -2.958 | -2.922 | 0.2518 | 0.1288 | -0.0537 | TRUE | 2.172 | 10 | -100 | 0.2676 | 0.2366 |  | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | II | FALSE | 1 | 0 | 0.9608 | 0.6892 | 2.896 | 3.004 | 0.0768 | 0.5318 | 0.0985 | FALSE | 2.5 | -4.25 | -68 | 0.8098 | 0.11 | 8 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | aVL | FALSE | 1 | 1 | 0.9624 | 0.1392 | 0.013 | 0.018 | 0.1553 | 0.0254 | 0.0155 | FALSE | 2.5 | -5.75 | -96 | 0.2948 | 0.0435 | -196 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | V2 | FALSE | 1 | 2 | 0.9368 | 0.355 | 0.399 | 0.402 | 0.378 | 0.2035 | 0.1368 | FALSE | 2.272 | 10 | 90 | 0.4232 | 0.1159 | 8 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | V5 | FALSE | 1 | 3 | 0.972 | 0.4106 | 0.948 | 1.007 | 0.1452 | 0.1688 | 0.0194 | FALSE | 2.48 | 0.75 | -2 | 0.4277 | 0.0201 | 12 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | III | FALSE | 2 | 0 | 0.9784 | 0.2933 | 0.369 | 0.507 | 0.111 | 0.1099 | 0.0347 | FALSE | 2.464 | 1.5 | 0 | 0.3416 | 0.0406 | -200 | FAILED | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | aVF | FALSE | 2 | 1 | 0.964 | 0.765 | 3.843 | 4.055 | 0.0527 | 0.6669 | 0.3468 | FALSE | 2.47 | 1.25 | 10 | 0.8286 | 0.3408 | 6 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | V3 | FALSE | 2 | 2 | 0.9592 | 0.2221 | -1.617 | -1.315 | 0.2899 | 0.1912 | 0.0016 | FALSE | 2.272 | 10 | 12 | 0.3187 | -0.0367 | 28 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | V6 | FALSE | 2 | 3 | 0.976 | 0.5165 | 1.526 | 1.655 | 0.1138 | 0.2468 | 0.1007 | FALSE | 2.5 | -2.75 | -12 | 0.5727 | 0.1199 | -18 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | 20261015_124715 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=234 grid=1811 locate=276 calibrate+segment=4 trace=347 cut=8 quality=1 score*=7 classify=53 | baseline=67.8/41.1 corners=67.8/41.1 rectify=120.0/41.3 grid=134.6/80.1 locate=114.0/49.4 calibrate+segment=86.0/36.5 trace=90.5/62.2 cut=87.7/41.4 quality=87.7/41.4 score*=87.7/42.8 classify=96.1/44.2 | 214.6 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | II | TRUE | 3 | 0 | 0.9986 | 0.476 | -1.529 | -1.445 | 0.1253 | 0.628 | 0.6044 |  | 9.976 | 0.25 | 4 | 0.4964 | 0.603 | 10 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | I | FALSE | 0 | 0 | 0.98 | -0.0463 | -5.445 | -4.642 | 0.3068 | -0.0728 | 0.017 | TRUE | 2.406 | -0.25 | -100 | 0.0998 | 0.0601 | 130 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | aVR | FALSE | 0 | 1 | 0.9288 | 0.0709 | -6.601 | -6.43 | 0.2852 | 0.1336 | 0.1504 | FALSE | 2.494 | -10 | 6 | 0.161 | 0.1527 | -8 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | V1 | FALSE | 0 | 2 | 0.9632 | -0.0226 | -4.347 | -4.347 | 0.4634 | -0.026 | 0.0086 | FALSE | 2.178 | 10 | -94 | 0.2131 | 0.2131 | 200 | FAILED | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | V4 | FALSE | 0 | 3 | 0.868 | 0.1967 | -1.537 | -1.489 | 0.2139 | 0.1894 | 0.0437 | FALSE | 2.5 | -2.25 | -2 | 0.2183 | 0.0437 | -12 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | II | FALSE | 1 | 0 | 0.8584 | 0.4641 | 0.937 | 1.234 | 0.0963 | 0.3126 | 0.0642 | FALSE | 2.21 | 10 | -62 | 0.6162 | 0.2854 | 142 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | aVL | FALSE | 1 | 1 | 0.9632 | 0.1398 | -2.247 | -2.008 | 0.2015 | 0.1334 | 0.0125 | FALSE | 2.472 | -10 | 28 | 0.8029 | 0.4905 | -126 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | V2 | FALSE | 1 | 2 | 0.7408 | -0.0647 | -4.623 | -4.428 | 0.6739 | -0.0869 | 0.035 | FALSE | 2.432 | -10 | 68 | -0.0052 | -0.0195 |  | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | V5 | FALSE | 1 | 3 | 0.7976 | 0.1763 | -1.125 | -1.1 | 0.1844 | 0.1433 | 0.0042 | FALSE | 2.5 | -4 | -38 | 0.2171 | 0.0033 |  | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | III | FALSE | 2 | 0 | 0.9392 | 0.1747 | -0.145 | -0.121 | 0.1178 | 0.0744 | 0.0164 | TRUE | 2.344 | 2.25 | -100 | 0.279 | 0.0335 | 114 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | aVF | FALSE | 2 | 1 | 0.964 | -0.0758 | -4.845 | -4.138 | 0.1433 | -0.1052 | 0.0438 | FALSE | 2.498 | -10 | 2 | 0.6438 | 0.2205 | -200 | FAILED | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | V3 | FALSE | 2 | 2 | 0.9632 | -0.1755 | -2.37 | -2.091 | 0.3161 | -0.1405 | -0.0002 | FALSE | 2.42 | -10 | 80 | 0.062 | 0.4102 | 200 | FAILED | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | V6 | FALSE | 2 | 3 | 0.8448 | 0.0939 | -1.755 | -1.747 | 0.166 | 0.0817 | 0.0015 | FALSE | 2.5 | -8.5 | -72 | 0.1863 | 0.0027 |  | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | 20261015_124811 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=259 grid=1974 locate=245 calibrate+segment=4 trace=374 cut=8 quality=1 score*=6 classify=64 | baseline=73.2/48.2 corners=73.2/48.2 rectify=125.3/48.3 grid=139.9/89.9 locate=119.1/79.5 calibrate+segment=91.0/27.9 trace=96.2/45.7 cut=91.3/47.2 quality=91.3/47.2 score*=91.3/48.6 classify=99.4/49.8 | 207.5 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | II | TRUE | 3 | 0 | 0.8704 | 0.1401 | -3.028 | -2.261 | 0.149 | 0.165 | 0.0205 |  | 10 | -10 | -224 | 0.5943 | 0.3684 | 44 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | I | FALSE | 0 | 0 | 0.9808 | -0.1185 | -6.424 | -5.92 | 0.3435 | -0.2062 | 0.022 | TRUE | 2.498 | -3.75 | -100 | -0.0638 | 0.0201 | 200 | FAILED | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVR | FALSE | 0 | 1 | 0.952 | -0.0168 | -0.652 | -0.539 | 0.1438 | -0.0076 | -0.0093 | FALSE | 2.498 | -5.75 | 2 | 0.194 | -0.0078 | -116 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V1 | FALSE | 0 | 2 | 0.964 | 0.0667 | -0.318 | -0.317 | 0.2915 | 0.023 | 0.0288 | TRUE | 2.172 | 10 | -100 | 0.2787 | 0.0387 | 200 | FAILED | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V4 | FALSE | 0 | 3 | 0.6776 | -0.0297 | -0.43 | -0.228 | 0.1883 | -0.0109 | 0.0203 | FALSE | 2.242 | 10 | -30 | 0.2951 | 0.1291 |  | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | FALSE | 1 | 0 | 0.9808 | 0.1099 | -0.559 | -0.331 | 0.1144 | 0.0581 | -0.0049 | TRUE | 2.172 | 10 | -100 | 0.3564 | 0.0652 | 200 | FAILED | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVL | FALSE | 1 | 1 | 0.956 | 0.1299 | -1.645 | -1.408 | 0.188 | 0.0957 | 0.0033 | FALSE | 2.436 | 0.25 | -58 | 0.1949 | 0.0472 | 60 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V2 | FALSE | 1 | 2 | 0.9048 | 0.0202 | -0.086 | -0.072 | 0.3997 | 0.0044 | 0.0562 | TRUE | 2.258 | 6 | -100 | 0.1509 | 0.0391 | 178 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V5 | FALSE | 1 | 3 | 0.7144 | -0.0701 | -0.908 | -0.768 | 0.1799 | -0.0209 | -0.0321 | FALSE | 2.474 | -10 | 26 | 0.3397 | 0.0239 |  | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | III | FALSE | 2 | 0 | 0.956 | 0.1194 | -0.305 | 0.116 | 0.12 | 0.0506 | 0.0414 | FALSE | 2.476 | -0.5 | -36 | 0.2317 | 0.0414 | 34 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVF | FALSE | 2 | 1 | 0.956 | -0.198 | -4.737 | -4.217 | 0.1415 | -0.2451 | 0.0258 | TRUE | 2.4 | -10 | 100 | 0.1668 | -0.0345 | 162 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V3 | FALSE | 2 | 2 | 0.9648 | 0.047 | -2.393 | -1.525 | 0.317 | 0.0461 | 0.2907 | TRUE | 2.464 | -2.5 | -100 | 0.2817 | 0.3349 | 40 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V6 | FALSE | 2 | 3 | 0.6832 | 0.0326 | -5.428 | -5.301 | 0.2533 | 0.0448 | -0.1398 | TRUE | 2.4 | -1 | 100 | 0.2621 | 0.2309 |  | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | 20261015_124915 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=202 grid=2057 locate=309 calibrate+segment=4 trace=427 cut=10 quality=1 score*=7 classify=61 | baseline=74.4/48.5 corners=74.4/48.5 rectify=126.5/48.7 grid=145.6/91.8 locate=125.3/56.3 calibrate+segment=96.9/39.8 trace=104.1/66.1 cut=97.7/29.6 quality=97.7/29.6 score*=97.7/30.9 classify=106.6/32.5 | 218.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | TRUE | 3 | 0 | 0.8852 | -0.0303 | -4.696 | -3.946 | 0.1805 | -0.043 | -0.0002 |  | 9.206 | -3.25 | 794 | 0.266 | 0.0669 | 158 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | I | FALSE | 0 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVR | FALSE | 0 | 1 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V1 | FALSE | 0 | 2 | 0.0896 | 0.1637 | -8.36 | -8.249 | 0.7356 | 0.3222 | 0 | FALSE | 2.5 | -3.75 | -98 | 0.1851 | 0 |  | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V4 | FALSE | 0 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | FALSE | 1 | 0 | 0.8616 | -0.1183 | -3.113 | -2.679 | 0.1535 | -0.1123 | -0.0134 | TRUE | 2.37 | 1.25 | -100 | 0.116 | 0.3061 | 122 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVL | FALSE | 1 | 1 | 0.9552 | 0.196 | -4.377 | -3.381 | 0.2574 | 0.292 | 0.0376 | FALSE | 2.406 | -10 | 94 | 0.2968 | 0.0526 | -190 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V2 | FALSE | 1 | 2 | 0.3256 | -0.084 | -4.668 | -4.632 | 0.6773 | -0.0881 | 0 | TRUE | 2.172 | 10 | -100 | 0.0916 | 0.3594 |  | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V5 | FALSE | 1 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | III | FALSE | 2 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVF | FALSE | 2 | 1 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V3 | FALSE | 2 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V6 | FALSE | 2 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | 20261015_125147 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=241 grid=2350 locate=599 calibrate+segment=4 trace=444 cut=1 quality=0 score*=2 classify=50 | baseline=72.5/49.3 corners=72.5/49.3 rectify=128.7/49.4 grid=147.7/96.3 locate=123.0/63.2 calibrate+segment=94.6/50.0 trace=100.1/68.2 cut=95.3/34.8 quality=95.3/34.9 score*=95.3/35.4 classify=103.8/36.6 | 217.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | TRUE | 3 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | I | FALSE | 0 | 0 | 0.9864 | 0.5733 | 1.249 | 11.808 | 0.142 | 0.5216 | 0.8767 | FALSE | 2.476 | 0.25 | -18 | 0.9674 | 0.8979 | 22 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | aVR | FALSE | 0 | 1 | 0.9968 | 0.8872 | 6.965 | 6.967 | 0.0598 | 0.7435 | 0.8959 | FALSE | 2.488 | 0.5 | 6 | 0.8958 | 0.8949 | 2 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | V1 | FALSE | 0 | 2 | 0.9984 | 0.9409 | 9.36 | 14.379 | 0.0956 | 0.8998 | 0.9512 | FALSE | 2.498 | -0.25 | 2 | 0.9872 | 0.9456 | -6 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | V4 | FALSE | 0 | 3 | 0.9984 | 0.963 | 11.676 | 13.229 | 0.0467 | 0.8976 | 0.888 | FALSE | 2.494 | 0.25 | 6 | 0.9784 | 0.8831 | -4 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | II | FALSE | 1 | 0 | 0.976 | 0.7186 | 2.94 | 3.975 | 0.0765 | 0.6669 | 0.5796 | FALSE | 2.5 | -3.75 | -72 | 0.9804 | 0.792 | 10 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | aVL | FALSE | 1 | 1 | 0.9984 | 0.7694 | 3.803 | 5.53 | 0.1004 | 0.7097 | 0.8872 | FALSE | 2.5 | -0.5 | 0 | 0.8607 | 0.8851 | -8 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | V2 | FALSE | 1 | 2 | 0.9976 | 0.9659 | 11.958 | 11.958 | 0.0999 | 0.9566 | 0.9189 | FALSE | 2.5 | -1 | -12 | 0.9916 | 0.917 | 0 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | V5 | FALSE | 1 | 3 | 0.9992 | 0.6928 | 2.637 | 6.896 | 0.1196 | 0.6627 | 0.9059 | FALSE | 2.458 | 1.75 | 8 | 0.969 | 0.8958 | 14 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | III | FALSE | 2 | 0 | 0.9968 | 0.3585 | -1.242 | 5.291 | 0.1337 | 0.3854 | 0.7124 | FALSE | 2.45 | 1.5 | 50 | 0.9166 | 0.9272 | -30 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | aVF | FALSE | 2 | 1 | 0.9992 | 0.9133 | 7.797 | 14.269 | 0.0334 | 0.9008 | 0.9377 | FALSE | 2.494 | 0 | 6 | 0.9807 | 0.938 | -6 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | V3 | FALSE | 2 | 2 | 0.9976 | 0.8929 | 7.316 | 7.419 | 0.1037 | 0.8565 | 0.884 | FALSE | 2.5 | -1 | -14 | 0.9826 | 0.8755 | 2 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | V6 | FALSE | 2 | 3 | 0.9872 | 0.745 | 3.386 | 11.014 | 0.0918 | 0.7099 | 0.9254 | FALSE | 2.498 | -1 | 2 | 0.9901 | 0.9258 | -14 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | 20261015_124711 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2564x1817 | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | corners=0 rectify=194 grid=1825 locate=278 calibrate+segment=13 trace=20295 cut=47 quality=10 score*=65 classify=90 | baseline=57.6/10.9 corners=57.6/10.9 rectify=109.6/11.0 grid=129.6/52.8 locate=108.9/48.0 calibrate+segment=80.3/48.2 trace=502.4/66.0 cut=439.6/37.1 quality=439.6/37.1 score*=440.9/39.0 classify=448.9/40.1 | 531 | unreliable | SCALE_DISAGREE/severe | II | TRUE | 3 | 0 | 0.9978 | 0.7321 | 2.698 | 2.698 | 0.077 | 0.7649 | 0.8155 |  | 9.95 | 0.5 | 22 | 0.7974 | 0.7501 | 0 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | 9 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | I | FALSE | 0 | 0 | 0.9864 | 0.003 | -4.046 | -3.479 | 0.2612 | 0.0039 | 0.013 | TRUE | 2.276 | 5.25 | -100 | 0.3719 | 0.3984 | 146 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVR | FALSE | 0 | 1 | 0.9984 | 0.2218 | -0.846 | -0.594 | 0.147 | 0.1801 | 0.2213 | FALSE | 2.488 | -10 | 12 | 0.5965 | 0.7822 | -8 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V1 | FALSE | 0 | 2 | 0.9984 | 0.001 | -2.252 | -2.248 | 0.3641 | 0.0008 | 0.0164 | FALSE | 2.182 | 10 | -90 | 0.033 | 0.0382 | -72 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V4 | FALSE | 0 | 3 | 0.6904 | 0.1193 | -1.435 | -1.423 | 0.2114 | 0.0969 | 0.0067 | FALSE | 2.5 | -7.25 | -40 | 0.2152 | -0.0065 |  | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | FALSE | 1 | 0 | 0.9808 | 0.2034 | -2.063 | -0.269 | 0.136 | 0.2117 | 0.0298 | TRUE | 2.444 | -1.75 | -100 | 0.5615 | 0.3596 | 164 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVL | FALSE | 1 | 1 | 0.9984 | 0.2482 | -1.024 | -0.871 | 0.175 | 0.2115 | 0.0146 | FALSE | 2.478 | -10 | 22 | 0.7313 | 0.8031 | -124 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V2 | FALSE | 1 | 2 | 0.9984 | -0.1265 | -2.47 | -2.427 | 0.5259 | -0.1034 | 0.0212 | TRUE | 2.172 | 10 | -100 | 0.0535 | 0.078 | -76 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V5 | FALSE | 1 | 3 | 0.6888 | 0.1386 | -1.473 | -1.372 | 0.1919 | 0.115 | 0.0048 | FALSE | 2.5 | -4.75 | -46 | 0.2144 | 0.0056 |  | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | III | FALSE | 2 | 0 | 0.9984 | 0.0605 | -5.612 | -2.137 | 0.2211 | 0.1034 | 0.1513 | TRUE | 2.49 | -3.5 | -100 | 0.9013 | 0.7642 | 68 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVF | FALSE | 2 | 1 | 0.9984 | 0.241 | -1.559 | -0.87 | 0.0981 | 0.2337 | 0.0711 | FALSE | 2.5 | -10 | 0 | 0.9403 | 0.8408 | -102 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V3 | FALSE | 2 | 2 | 0.9984 | -0.175 | -2.501 | -2.203 | 0.3209 | -0.1454 | 0.0007 | FALSE | 2.42 | -10 | 80 | 0.0624 | 0.4163 | -92 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V6 | FALSE | 2 | 3 | 0.6936 | 0.1335 | -1.462 | -1.462 | 0.1605 | 0.1118 | 0.0179 | FALSE | 2.5 | -8.5 | -74 | 0.2153 | 0.0038 |  | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | 20261015_124807 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=155 grid=1867 locate=253 calibrate+segment=4 trace=20003 cut=11 quality=1 score*=9 classify=63 | baseline=66.7/43.1 corners=66.7/43.1 rectify=118.7/43.2 grid=134.9/80.1 locate=114.4/77.9 calibrate+segment=86.2/26.4 trace=508.0/66.3 cut=445.7/43.6 quality=445.4/43.7 score*=445.4/45.0 classify=453.6/46.6 | 546.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | TRUE | 3 | 0 | 0.9106 | 0.129 | -2.404 | -1.874 | 0.1386 | 0.1325 | 0.0313 |  | 10 | -10 | -222 | 0.5899 | 0.7498 | 30 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | I | FALSE | 0 | 0 | 0.8728 | -0.0878 | -7.063 | -6.623 | 0.3697 | -0.1687 | 0.0221 | FALSE | 2.372 | 1.25 | -98 | 0.1085 | 0.3705 | 120 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVR | FALSE | 0 | 1 | 0.9984 | 0.0132 | -3.039 | -2.971 | 0.1893 | 0.0144 | 0.1194 | FALSE | 2.284 | 9.5 | 26 | 0.2389 | 0.1705 | -162 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V1 | FALSE | 0 | 2 | 0.9984 | -0.1302 | -2.318 | -1.678 | 0.3669 | -0.0933 | 0.0274 | TRUE | 2.358 | 1.75 | -100 | 0.3249 | 0.0461 | 126 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V4 | FALSE | 0 | 3 | 0.5448 | -0.0056 | -1.58 | -1.459 | 0.2149 | -0.0033 | -0.0192 | FALSE | 2.452 | -10 | 48 | 0.6841 | 0.3706 |  | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | FALSE | 1 | 0 | 0.8744 | -0.2347 | -9.321 | -8.719 | 0.3136 | -0.5801 | -0.028 | TRUE | 2.394 | 0.25 | -100 | -0.018 | 0.3799 | 106 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVL | FALSE | 1 | 1 | 0.9984 | 0.1893 | -2.06 | -1.683 | 0.1972 | 0.1909 | 0.0367 | TRUE | 2.4 | -7.75 | 100 | 0.322 | -0.016 | -172 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V2 | FALSE | 1 | 2 | 0.9984 | -0.0894 | -1.892 | -1.338 | 0.4921 | -0.0642 | 0.0634 | TRUE | 2.364 | 1.5 | -100 | 0.2721 | 0.0688 | 126 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V5 | FALSE | 1 | 3 | 0.5464 | -0.05 | -1.309 | -1.234 | 0.1883 | -0.0283 | -0.022 | FALSE | 2.472 | -10 | 28 | 0.6845 | 0.3938 |  | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | III | FALSE | 2 | 0 | 0.9568 | -0.0469 | -7.17 | -6.952 | 0.2645 | -0.0932 | 0.0491 | TRUE | 2.21 | 8.25 | -100 | 0.2426 | 0.3705 | 176 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVF | FALSE | 2 | 1 | 0.9984 | 0.1885 | -3.658 | -1.386 | 0.125 | 0.2585 | -0.0325 | FALSE | 2.446 | -10 | 54 | 0.5595 | 0.3733 | -156 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V3 | FALSE | 2 | 2 | 0.9984 | -0.1108 | -1.818 | -0.65 | 0.2967 | -0.0804 | 0.0217 | FALSE | 2.26 | 10 | -12 | 0.2638 | 0.3188 | 148 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V6 | FALSE | 2 | 3 | 0.5432 | -0.0353 | -1.169 | -1.005 | 0.1551 | -0.0198 | -0.0109 | FALSE | 2.464 | -10 | 36 | 0.6943 | 0.4155 |  | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | 20261015_124910 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | corners=0 rectify=189 grid=2062 locate=302 calibrate+segment=4 trace=19553 cut=11 quality=1 score*=6 classify=59 | baseline=67.5/43.4 corners=67.5/43.4 rectify=119.6/43.4 grid=139.2/83.0 locate=118.9/53.3 calibrate+segment=90.5/47.2 trace=512.6/71.5 cut=450.4/44.2 quality=450.1/44.2 score*=450.1/45.5 classify=458.8/47.1 | 551.8 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | TRUE | 3 | 0 | 0.8584 | 0.0712 | -2.021 | -1.346 | 0.1326 | 0.0631 | 0 |  | 3.59 | 8.75 | 6410 |  |  | 44 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 3 | pulses
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | I | FALSE | 0 | 0 | 0.896 | -0.0934 | -5.004 | -4.587 | 0.2916 | -0.1293 | 0.0155 | TRUE | 2.34 | 2.5 | -100 | 0.205 | 0.3912 | 142 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVR | FALSE | 0 | 1 | 0.9992 | 0.0156 | -3.354 | -3.013 | 0.1963 | 0.0182 | 0.1662 | TRUE | 2.4 | -8.25 | 100 | 0.3059 | 0.366 | -176 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V1 | FALSE | 0 | 2 | 1 | 0.0287 | -1.73 | -1.6 | 0.3429 | 0.0207 | 0.0269 | TRUE | 2.394 | 0.25 | -100 | 0.3398 | 0.0293 | 106 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V4 | FALSE | 0 | 3 | 0.5048 | 0.0367 | -1.073 | -0.627 | 0.2027 | 0.0201 | 0.0714 | FALSE | 2.406 | -10 | 94 | 0.6854 | 0.3923 |  | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | FALSE | 1 | 0 | 0.8632 | -0.3068 | -12.746 | -12.472 | 0.4652 | -1.1603 | 0.0043 | TRUE | 2.37 | 1.25 | -100 | -0.222 | 0.3543 | 126 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVL | FALSE | 1 | 1 | 0.9984 | 0.1693 | -2.344 | -1.866 | 0.2037 | 0.1776 | 0.0345 | FALSE | 2.408 | -10 | 92 | 0.3235 | 0.0549 | -188 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V2 | FALSE | 1 | 2 | 0.9984 | 0.1279 | -0.914 | -0.871 | 0.4397 | 0.09 | 0.0571 | TRUE | 2.406 | -0.25 | -100 | 0.2342 | 0.0663 | 184 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V5 | FALSE | 1 | 3 | 0.5128 | -0.0161 | -1.093 | -0.462 | 0.1837 | -0.0086 | 0.0339 | FALSE | 2.41 | -10 | 90 | 0.6746 | 0.3857 |  | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | III | FALSE | 2 | 0 | 0.9264 | -0.0227 | -9.659 | -9.257 | 0.3523 | -0.066 | 0.0302 | FALSE | 2.398 | 0.25 | -96 | 0.0981 | 0.5746 | 100 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | aVF | FALSE | 2 | 1 | 0.9984 | 0.057 | -4.022 | -2.376 | 0.1303 | 0.0755 | -0.0022 | FALSE | 2.428 | -10 | 72 | 0.4866 | 0.2866 | -172 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V3 | FALSE | 2 | 2 | 0.9984 | 0.1104 | -0.846 | -0.678 | 0.2653 | 0.0782 | -0.0094 | FALSE | 2.284 | 9.5 | 22 | 0.2776 | 0.1196 | 106 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | V6 | FALSE | 2 | 3 | 0.5168 | -0.1061 | -15.143 | -15.12 | 0.7752 | -0.6011 | 0.2193 | TRUE | 2.4 | -10 | 100 | 0.0123 | 1.191 |  | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | 20261015_125142 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | corners=0 rectify=210 grid=2222 locate=592 calibrate+segment=4 trace=20110 cut=16 quality=1 score*=9 classify=69 | baseline=65.6/43.8 corners=65.6/43.8 rectify=121.6/43.9 grid=141.6/85.6 locate=117.0/69.5 calibrate+segment=88.8/58.2 trace=510.6/68.1 cut=448.0/45.4 quality=447.9/45.4 score*=447.9/46.7 classify=457.0/48.2 | 553.2 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | II | TRUE | 3 | 0 | 0.8376 | 0.0347 | -4.233 | -4.056 | 0.1711 | 0.0476 | 0.1405 |  | 9.304 | 0 | 696 | 0.3095 | 0.0769 |  | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 3 | rhythm_locator
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | I | FALSE | 0 | 0 | 0.86 | 0.5607 | 1.553 | 3.105 | 0.179 | 0.3955 | 0.4924 | TRUE | 2.352 | 6.25 | 100 | 0.9151 | 0.5158 | -32 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | aVR | FALSE | 0 | 1 | 0.872 | 0.4245 | 0.398 | 0.492 | 0.1609 | 0.2871 | -0.006 | FALSE | 2.376 | 5.25 | 82 | 0.569 | 0.4739 | 6 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | V1 | FALSE | 0 | 2 | 0.9624 | 0.8173 | 4.803 | 4.981 | 0.2146 | 0.6869 | 0.7636 | FALSE | 2.498 | -0.5 | 2 | 0.8267 | 0.7629 | -6 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | V4 | FALSE | 0 | 3 | 0.9792 | 0.618 | 2.116 | 2.419 | 0.2134 | 0.391 | 0.1211 | FALSE | 2.458 | 1.75 | 18 | 0.6633 | 0.1208 | 12 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | II | FALSE | 1 | 0 | 0.9792 | 0.6969 | 2.726 | 3.217 | 0.0906 | 0.5864 | 0.4872 | FALSE | 2.458 | 1.75 | 20 | 0.8051 | 0.4694 | -10 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | aVL | FALSE | 1 | 1 | 0.9616 | 0.9701 | 12.43 | 12.634 | 0.0492 | 0.9336 | 0.8996 | FALSE | 2.482 | 0.75 | 8 | 0.9765 | 0.8866 | 2 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | V2 | FALSE | 1 | 2 | 0.7288 | 0.6771 | 1.741 | 1.741 | 0.3745 | 0.5997 | 0.9632 | FALSE | 2.476 | 1 | 10 | 0.6859 | 0.9627 |  | UNRELIABLE | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | V5 | FALSE | 1 | 3 | 0.9792 | 0.6604 | 2.388 | 3.046 | 0.0864 | 0.5344 | 0.86 | FALSE | 2.428 | 3 | 4 | 0.7904 | 0.9618 | 12 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | III | FALSE | 2 | 0 | 0.9792 | 0.3621 | -0.882 | 4.951 | 0.1965 | 0.362 | 0.8187 | FALSE | 2.498 | -1.75 | -46 | 0.9077 | 0.9382 | 32 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | aVF | FALSE | 2 | 1 | 0.9632 | 0.9233 | 8.392 | 10.081 | 0.0433 | 0.8699 | 0.9134 | FALSE | 2.496 | 0 | 4 | 0.9488 | 0.9134 | -4 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | V3 | FALSE | 2 | 2 | 0.9624 | 0.9622 | 11.49 | 17.715 | 0.0925 | 0.9607 | 0.9835 | FALSE | 2.494 | -0.25 | 6 | 0.9927 | 0.9825 | -10 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | V6 | FALSE | 2 | 3 | 0.9776 | 0.92 | 7.696 | 8.015 | 0.0382 | 0.9609 | 0.8725 | FALSE | 2.494 | 0.25 | 4 | 0.9305 | 0.8749 | -2 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | 20261015_155455 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=214 grid=2493 locate=391 calibrate+segment=11 trace=409 cut=11 quality=0 score*=6 classify=51 | baseline=73.2/52.6 corners=73.3/52.6 rectify=129.2/52.8 grid=147.6/97.3 locate=123.3/65.4 calibrate+segment=95.2/49.1 trace=101.2/76.4 cut=95.8/27.9 quality=95.8/27.5 score*=95.8/28.9 classify=104.6/30.4 | 235.4 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/V2 | II | TRUE | 3 | 0 | 0.997 | 0.4773 | -0.187 | 0.076 | 0.1371 | 0.4767 | 0.9906 |  | 3.182 | -2.5 | 6818 |  |  | -12 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | I | FALSE | 0 | 0 | 0.9992 | 0.5781 | 0.918 | 2.73 | 0.1926 | 0.5605 | 0.5214 | TRUE | 2.364 | 5.75 | 100 | 0.9767 | 0.9969 | -34 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | aVR | FALSE | 0 | 1 | 0.9984 | 0.9443 | 9.636 | 9.773 | 0.0556 | 0.9416 | 0.9552 | FALSE | 2.476 | 1 | 8 | 0.9781 | 0.9543 | 2 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | V1 | FALSE | 0 | 2 | 0.9992 | 0.9869 | 15.867 | 17.662 | 0.06 | 0.9653 | 0.9874 | FALSE | 2.5 | -0.5 | -2 | 0.9951 | 0.9848 | -4 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | V4 | FALSE | 0 | 3 | 0.9984 | 0.8573 | 5.524 | 11.63 | 0.1441 | 0.8646 | 0.9971 | FALSE | 2.478 | 0.25 | -16 | 0.9654 | 0.9986 | 18 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | II | FALSE | 1 | 0 | 0.9984 | 0.7486 | 3.041 | 4.57 | 0.0874 | 0.7464 | 0.9699 | FALSE | 2.45 | 2 | 20 | 0.9542 | 0.9624 | 12 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | aVL | FALSE | 1 | 1 | 1 | 0.9949 | 20.019 | 22.181 | 0.0205 | 0.9816 | 0.9706 | FALSE | 2.494 | 0.25 | 0 | 0.9974 | 0.9628 | 2 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | V2 | FALSE | 1 | 2 | 0.9984 | 0.9918 | 17.891 | 18.804 | 0.0583 | 0.9809 | 0.9914 | FALSE | 2.488 | 0.5 | 2 | 0.997 | 0.9899 | 2 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | V5 | FALSE | 1 | 3 | 0.9984 | 0.8341 | 4.914 | 7.933 | 0.0646 | 0.8315 | 1.0018 | FALSE | 2.482 | 0.5 | -6 | 0.9239 | 0.9998 | 12 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | III | FALSE | 2 | 0 | 0.9888 | 0.6786 | 2.272 | 8.112 | 0.1367 | 0.6567 | 0.9617 | FALSE | 2.494 | -1.5 | -44 | 0.944 | 0.956 | 24 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | aVF | FALSE | 2 | 1 | 1 | 0.9471 | 9.935 | 10.854 | 0.0362 | 0.9226 | 0.9732 | FALSE | 2.498 | 0 | 2 | 0.9573 | 0.9731 | -2 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | V3 | FALSE | 2 | 2 | 0.9992 | 0.9709 | 12.551 | 17.613 | 0.0819 | 0.9766 | 0.9906 | FALSE | 2.494 | -0.25 | 6 | 0.9928 | 0.9898 | -8 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | V6 | FALSE | 2 | 3 | 0.9904 | 0.9286 | 8.154 | 8.154 | 0.0362 | 0.976 | 0.9702 | FALSE | 2.494 | 0.25 | 2 | 0.9323 | 0.9691 | 0 | ok | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | 20261015_155451 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2876x2020 | ticks 10.2627 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=210 grid=2508 locate=394 calibrate+segment=11 trace=21050 cut=11 quality=0 score*=12 classify=88 | baseline=70.4/27.5 corners=70.4/27.5 rectify=126.5/27.6 grid=146.5/96.3 locate=121.5/62.4 calibrate+segment=93.4/47.4 trace=515.4/71.7 cut=451.9/45.6 quality=451.9/45.6 score*=452.3/47.3 classify=461.0/48.7 | 560.3 | unreliable | T0_GUARD/severe | II | TRUE | 3 | 0 | 0.9966 | 0.4693 | -0.413 | -0.214 | 0.1407 | 0.4856 | 1.0423 |  | 3.18 | -2.25 | 6820 |  |  | -10 | UNRELIABLE | grid (fit rms 0.214 mm, 55x37 rules, 10.2627 px/mm, inliers 0.694) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | I | FALSE | 0 | 0 | 0.9792 | 0.344 | 0.458 | 0.647 | 0.1555 | 0.1644 | 0.0472 | FALSE | 2.348 | 6.5 | 98 | 0.4473 | 0.0712 | -26 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | aVR | FALSE | 0 | 1 | 0.7856 | 0.133 | -6.702 | -6.623 | 0.2886 | 0.2698 | 0.1831 | FALSE | 2.272 | 10 | 6 | 0.166 | 0.1805 |  | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | V1 | FALSE | 0 | 2 | 0.9624 | 0.7456 | 3.46 | 6.083 | 0.1887 | 0.6143 | 0.878 | FALSE | 2.492 | -0.5 | 8 | 0.8746 | 0.8779 | -12 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | V4 | FALSE | 0 | 3 | 0.9792 | 0.0561 | -4.748 | -4.187 | 0.3095 | 0.0831 | 0.3896 | FALSE | 2.386 | 4.75 | 24 | 0.434 | 0.7762 | 50 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | II | FALSE | 1 | 0 | 0.9784 | 0.637 | 2.276 | 4.107 | 0.0825 | 0.4771 | 0.0081 | FALSE | 2.472 | -1 | 28 | 0.7818 | 0.1988 | -36 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | aVL | FALSE | 1 | 1 | 0.9624 | 0.8817 | 5.743 | 6.368 | 0.0803 | 0.5824 | 0.5823 | FALSE | 2.488 | 0.5 | 12 | 0.9231 | 0.5775 | -6 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | V2 | FALSE | 1 | 2 | 0.9616 | 0.8022 | 2.722 | 3.645 | 0.2893 | 0.9597 | 0.825 | FALSE | 2.486 | 0.25 | 14 | 0.8475 | 0.8204 | -12 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | V5 | FALSE | 1 | 3 | 0.9792 | 0.151 | -1.86 | -0.016 | 0.2007 | 0.1477 | 0.39 | FALSE | 2.364 | 5.75 | 16 | 0.9914 | 0.8749 | 78 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | III | FALSE | 2 | 0 | 0.9792 | 0.1787 | 0.08 | 0.617 | 0.1148 | 0.0642 | 0.0056 | FALSE | 2.41 | 3.5 | 90 | 0.3999 | 0.0382 | -64 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | aVF | FALSE | 2 | 1 | 0.8752 | 0.5089 | -0.389 | 0.88 | 0.0858 | 0.5741 | 0.2606 | FALSE | 2.492 | -1.5 | -46 | 0.655 | 0.3671 | 16 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | V3 | FALSE | 2 | 2 | 0.9632 | 0.8589 | 6.172 | 6.219 | 0.1182 | 0.8089 | 0.7943 | FALSE | 2.5 | -1 | -12 | 0.9393 | 0.792 | -2 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | V6 | FALSE | 2 | 3 | 0.9792 | 0.478 | 0.12 | 3.812 | 0.1337 | 0.4729 | 0.6254 | FALSE | 2.428 | 3 | 10 | 0.9828 | 0.9276 | 26 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | 20261015_153340 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=178 grid=2024 locate=251 calibrate+segment=11 trace=353 cut=9 quality=1 score*=8 classify=85 | baseline=72.4/49.1 corners=72.4/49.1 rectify=124.6/49.3 grid=138.5/86.2 locate=118.0/79.6 calibrate+segment=89.7/28.3 trace=93.7/44.3 cut=90.3/45.9 quality=90.3/45.9 score*=90.3/47.3 classify=98.6/48.8 | 206.3 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/aVR | II | TRUE | 3 | 0 | 0.9512 | 0.391 | -0.883 | 0.629 | 0.1164 | 0.3987 | 0.286 |  | 9.852 | 1.5 | 80 | 0.6726 | 0.6948 | 30 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | I | FALSE | 0 | 0 | 0.6184 | 0.1266 | 0.017 | 0.026 | 0.1636 | 0.0109 | -0.0043 | TRUE | 2.4 | -10 | 100 | 0.1836 | -0.0058 |  | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | aVR | FALSE | 0 | 1 | 0.9632 | 0.3194 | 0.448 | 0.451 | 0.1267 | 0.1351 | -0.0081 | FALSE | 2.5 | -0.5 | -10 | 0.3219 | -0.0068 | 2 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | V1 | FALSE | 0 | 2 | 0.964 | 0.5722 | 1.51 | 1.626 | 0.2361 | 0.265 | 0.0374 | FALSE | 2.482 | 0.25 | 18 | 0.593 | 0.0409 | -18 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | V4 | FALSE | 0 | 3 | 0.9808 | 0.3512 | 0.554 | 0.742 | 0.1681 | 0.2228 | 0.0369 | FALSE | 2.386 | 4.75 | 26 | 0.4583 | 0.0219 | 28 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | II | FALSE | 1 | 0 | 0.98 | 0.6844 | 2.756 | 3.164 | 0.0781 | 0.4666 | -0.0024 | FALSE | 2.486 | -0.25 | 14 | 0.72 | 0.0004 | -18 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | aVL | FALSE | 1 | 1 | 0.9648 | 0.3915 | 0.561 | 0.578 | 0.1458 | 0.0855 | 0.0179 | FALSE | 2.5 | -0.75 | -18 | 0.402 | 0.0172 | 6 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | V2 | FALSE | 1 | 2 | 0.868 | 0.3843 | 0.65 | 0.762 | 0.3672 | 0.247 | -0.0111 | TRUE | 2.272 | 10 | 100 | 0.5234 | 0.0206 | -22 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | V5 | FALSE | 1 | 3 | 0.9792 | 0.3467 | 0.623 | 0.675 | 0.1508 | 0.1937 | 0.133 | FALSE | 2.32 | 7.75 | 48 | 0.6211 | 0.1886 | 10 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | III | FALSE | 2 | 0 | 0.9808 | 0.1993 | 0.165 | 0.351 | 0.1137 | 0.0726 | -0.0009 | TRUE | 2.4 | 1.5 | 100 | 0.2991 | 0.01 | -98 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | aVF | FALSE | 2 | 1 | 0.928 | 0.3228 | -0.271 | 0.155 | 0.0846 | 0.2475 | -0.0681 | FALSE | 2.41 | 3.75 | 12 | 0.4168 | 0 | 34 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | V3 | FALSE | 2 | 2 | 0.964 | 0.2086 | -3 | -2.892 | 0.3399 | 0.2614 | -0.0868 | FALSE | 2.478 | 0.75 | 22 | 0.2292 | -0.0505 | -14 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | V6 | FALSE | 2 | 3 | 0.9808 | 0.472 | -0.06 | 2.983 | 0.1365 | 0.4806 | 0.5738 | FALSE | 2.422 | 3.25 | 12 | 0.9541 | 0.8636 | 28 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | 20261015_154123 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=218 grid=2182 locate=324 calibrate+segment=16 trace=430 cut=8 quality=1 score*=8 classify=85 | baseline=73.6/50.7 corners=73.6/50.7 rectify=125.6/50.9 grid=144.4/90.5 locate=124.3/66.4 calibrate+segment=95.9/50.3 trace=101.3/66.2 cut=97.3/36.0 quality=97.3/36.0 score*=97.3/37.4 classify=106.0/38.9 | 221 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I | II | TRUE | 3 | 0 | 0.9556 | 0.4204 | -0.959 | 0.321 | 0.1174 | 0.4703 | 0.5343 |  | 9.9 | 1 | 38 | 0.6728 | 0.6982 | 30 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | I | FALSE | 0 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVR | FALSE | 0 | 1 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V1 | FALSE | 0 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V4 | FALSE | 0 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | FALSE | 1 | 0 | 0.9528 | 0.6256 | 1.88 | 2.097 | 0.0864 | 0.5504 | 0.6953 | FALSE | 2.5 | -3 | -36 | 0.775 | 0.694 | -8 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVL | FALSE | 1 | 1 | 0.888 | 0.7555 | 3.473 | 6.071 | 0.1043 | 0.7152 | 0.8448 | FALSE | 2.496 | -0.5 | 4 | 0.8795 | 0.8411 | -10 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V2 | FALSE | 1 | 2 | 0.9016 | 0.9335 | 9.131 | 9.45 | 0.1383 | 0.8775 | 0.8815 | FALSE | 2.5 | -1 | -10 | 0.9795 | 0.8796 | -4 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V5 | FALSE | 1 | 3 | 0.0264 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | III | FALSE | 2 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVF | FALSE | 2 | 1 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V3 | FALSE | 2 | 2 | 0.1744 | 0.32 | 0.55 | 1.056 | 0.2259 | 0.1447 | 0 | FALSE | 2.416 | 3.25 | 84 | 0.4448 | 0 |  | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V6 | FALSE | 2 | 3 | 0.044 | -0.2227 | -1.051 | -0.483 | 0.153 | -0.0724 | 0 | TRUE | 2.398 | 4.25 | 100 | -0.048 | 0 |  | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | 20261015_154216 | ptbxl40_kit_500hz_300dpi | 40 | classical | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=213 grid=2431 locate=649 calibrate+segment=12 trace=481 cut=2 quality=1 score*=2 classify=65 | baseline=73.0/51.0 corners=73.0/51.0 rectify=129.1/51.2 grid=151.6/92.8 locate=123.8/60.9 calibrate+segment=95.5/49.7 trace=101.7/62.4 cut=96.5/35.5 quality=96.5/35.5 score*=96.5/36.1 classify=104.6/37.2 | 220 | unreliable | T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | TRUE | 3 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | I | FALSE | 0 | 0 | 0.888 | 0.5649 | 1.156 | 1.52 | 0.1873 | 0.493 | 0.1258 | FALSE | 2.336 | 7 | 72 | 0.7743 | 0.6396 | -16 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | aVR | FALSE | 0 | 1 | 0.5088 | 0.4453 | 0.579 | 0.646 | 0.1576 | 0.2822 | 0 | FALSE | 2.376 | 5.25 | 82 | 0.5797 | 0.4703 |  | UNRELIABLE | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | V1 | FALSE | 0 | 2 | 0.9624 | 0.8645 | 5.878 | 6.09 | 0.1896 | 0.7035 | 0.9627 | FALSE | 2.5 | -0.75 | 0 | 0.8769 | 0.9622 | -4 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | V4 | FALSE | 0 | 3 | 0.848 | 0.2313 | -2.631 | -2.606 | 0.3686 | 0.2629 | 0.1168 | FALSE | 2.488 | 0.5 | 12 | 0.2323 | 0.1168 |  | UNRELIABLE | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | II | FALSE | 1 | 0 | 0.98 | 0.3487 | -1.472 | 3.843 | 0.1469 | 0.3778 | 0.3055 | FALSE | 2.48 | -1 | -46 | 0.803 | 0.5408 | 38 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | aVL | FALSE | 1 | 1 | 0.9592 | 0.7034 | 0.667 | 0.7 | 0.1907 | 0.8288 | 0.8558 | FALSE | 2.476 | 1 | 16 | 0.7179 | 0.8546 | -2 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | V2 | FALSE | 1 | 2 | 0.9624 | 0.6859 | 2.484 | 2.5 | 0.3438 | 0.4192 | 0.0192 | FALSE | 2.498 | -0.25 | 2 | 0.6878 | 0.019 | -4 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | V5 | FALSE | 1 | 3 | 0.7616 | 0.4223 | -5.884 | -5.83 | 0.2239 | 0.8448 | 0.9177 | FALSE | 2.45 | 2 | 4 | 0.4577 | 0.9163 |  | UNRELIABLE | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | III | FALSE | 2 | 0 | 0.9448 | 0.4371 | -0.176 | 4.085 | 0.1812 | 0.4247 | 0.7836 | TRUE | 2.49 | -3.5 | -100 | 0.9391 | 0.9278 | 72 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | aVF | FALSE | 2 | 1 | 0.9632 | 0.914 | 7.911 | 9.974 | 0.0457 | 0.861 | 0.9023 | FALSE | 2.496 | 0 | 4 | 0.9474 | 0.9015 | -4 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | V3 | FALSE | 2 | 2 | 0.9624 | 0.9642 | 11.736 | 17.269 | 0.0899 | 0.9555 | 0.973 | FALSE | 2.496 | -0.5 | 4 | 0.9924 | 0.9718 | -8 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | V6 | FALSE | 2 | 3 | 0.9656 | 0.7985 | 3.865 | 7.292 | 0.0594 | 0.8129 | 0.8647 | FALSE | 2.496 | -0.75 | 4 | 0.9315 | 0.8659 | -12 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | 20261015_154538 | ptbxl4016_kit_500hz_300dpi | 4016 | classical | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=227 grid=2817 locate=416 calibrate+segment=4 trace=459 cut=9 quality=0 score*=6 classify=84 | baseline=74.7/52.2 corners=74.7/52.2 rectify=130.8/52.4 grid=149.2/96.0 locate=125.0/63.4 calibrate+segment=96.8/63.5 trace=102.8/63.5 cut=97.8/28.0 quality=97.8/28.0 score*=97.8/29.4 classify=105.7/30.9 | 224.1 | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V5 | II | TRUE | 3 | 0 | 0.991 | 0.3653 | -1.394 | -1.276 | 0.1576 | 0.3936 | 0.95 |  | 3.236 | -4 | 6764 |  |  | 8 | UNRELIABLE | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | I | FALSE | 0 | 0 | 0.9976 | 0.5738 | 0.841 | 1.688 | 0.1488 | 0.5767 | 0.6837 | FALSE | 2.416 | 3.5 | 58 | 0.9553 | 0.8919 | -28 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | aVR | FALSE | 0 | 1 | 0.9984 | 0.8595 | 5.999 | 11.748 | 0.0669 | 0.7993 | 0.8911 | FALSE | 2.496 | -0.25 | -10 | 0.9728 | 0.8899 | 8 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | V1 | FALSE | 0 | 2 | 0.9984 | 0.8429 | 5.139 | 13.041 | 0.1555 | 0.8191 | 0.9454 | FALSE | 2.49 | 0 | 10 | 0.9752 | 0.945 | -10 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | V4 | FALSE | 0 | 3 | 0.9984 | 0.293 | -1.295 | 0.985 | 0.208 | 0.3019 | 0.4686 | FALSE | 2.398 | 4.25 | 14 | 0.9694 | 0.9268 | 52 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | II | FALSE | 1 | 0 | 0.9984 | 0.7916 | 4.041 | 15.934 | 0.0673 | 0.773 | 0.8977 | FALSE | 2.486 | 0 | 14 | 0.9881 | 0.897 | -14 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | aVL | FALSE | 1 | 1 | 0.9976 | 0.9635 | 11.587 | 12.42 | 0.041 | 0.9165 | 0.9036 | FALSE | 2.5 | -0.25 | -6 | 0.9779 | 0.9011 | 2 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | V2 | FALSE | 1 | 2 | 0.9968 | 0.9521 | 10.512 | 16.279 | 0.118 | 0.9377 | 0.921 | FALSE | 2.496 | -0.25 | 4 | 0.9917 | 0.9167 | -8 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | V5 | FALSE | 1 | 3 | 0.9984 | 0.132 | -2.235 | -0.205 | 0.2095 | 0.1376 | 0.4095 | FALSE | 2.364 | 5.75 | 14 | 0.9836 | 0.9564 | 80 | UNRELIABLE | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | III | FALSE | 2 | 0 | 0.9976 | 0.071 | -2.796 | -0.959 | 0.1598 | 0.0755 | 0.1631 | FALSE | 2.418 | -5.75 | 82 | 0.5827 | 0.4501 | -120 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | aVF | FALSE | 2 | 1 | 0.9984 | 0.8153 | 3.618 | 6.686 | 0.0541 | 0.9287 | 0.957 | FALSE | 2.484 | 0.5 | -4 | 0.9464 | 0.9361 | 12 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | V3 | FALSE | 2 | 2 | 0.9952 | 0.8589 | 6.05 | 6.05 | 0.1199 | 0.8415 | 0.9396 | FALSE | 2.5 | -1 | -14 | 0.9413 | 0.9224 | 0 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | V6 | FALSE | 2 | 3 | 0.9984 | 0.4378 | -0.359 | 3.317 | 0.1413 | 0.4492 | 0.5989 | FALSE | 2.428 | 3 | 8 | 0.9862 | 0.974 | 28 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | 20261015_153336 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2657x1867 | ticks 9.4093 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=163 grid=1875 locate=240 calibrate+segment=14 trace=19817 cut=10 quality=1 score*=8 classify=57 | baseline=68.4/54.7 corners=68.4/54.7 rectify=120.4/54.8 grid=140.3/91.7 locate=114.1/60.9 calibrate+segment=85.9/61.3 trace=512.7/66.4 cut=446.2/43.8 quality=446.1/43.8 score*=446.1/45.3 classify=455.7/46.8 | 551.3 | unreliable | T0_GUARD/severe | II | TRUE | 3 | 0 | 0.9992 | 0.3705 | -1.171 | 0.137 | 0.1203 | 0.4021 | 0.2902 |  | 9.852 | 1.5 | 80 | 0.6857 | 0.8864 | 30 | ok | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | I | FALSE | 0 | 0 | 0.9976 | 0.589 | 1.142 | 2.179 | 0.1437 | 0.5722 | 0.7556 | FALSE | 2.422 | 3.25 | 54 | 0.9545 | 0.8628 | -24 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | aVR | FALSE | 0 | 1 | 0.9976 | 0.7994 | 4.569 | 7.285 | 0.0788 | 0.6902 | 0.8854 | FALSE | 2.494 | -0.25 | -12 | 0.9042 | 0.8793 | 8 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | V1 | FALSE | 0 | 2 | 0.9976 | 0.8458 | 5.234 | 12.979 | 0.1538 | 0.8179 | 0.9288 | FALSE | 2.49 | 0 | 10 | 0.975 | 0.9282 | -10 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | V4 | FALSE | 0 | 3 | 0.9984 | 0.3238 | -1.098 | 0.995 | 0.2033 | 0.3331 | 0.4703 | FALSE | 2.392 | 4.5 | 18 | 0.9704 | 0.9446 | 52 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | II | FALSE | 1 | 0 | 0.9976 | 0.8077 | 4.454 | 12.485 | 0.0642 | 0.7749 | 0.8705 | FALSE | 2.496 | -0.75 | 4 | 0.9849 | 0.8399 | -14 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | aVL | FALSE | 1 | 1 | 0.9992 | 0.9719 | 12.626 | 13.597 | 0.0364 | 0.9151 | 0.9112 | FALSE | 2.498 | 0 | -2 | 0.9779 | 0.9111 | 2 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | V2 | FALSE | 1 | 2 | 0.9976 | 0.9443 | 9.875 | 15.37 | 0.127 | 0.926 | 0.9043 | FALSE | 2.498 | -0.5 | 2 | 0.99 | 0.902 | -8 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | V5 | FALSE | 1 | 3 | 0.9968 | 0.164 | -2.026 | -0.152 | 0.2045 | 0.1692 | 0.4404 | FALSE | 2.358 | 6 | 18 | 0.9859 | 0.94 | 80 | UNRELIABLE | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | III | FALSE | 2 | 0 | 0.9984 | 0.0795 | -2.543 | -0.484 | 0.1553 | 0.0807 | 0.1792 | TRUE | 2.4 | 3 | 100 | 0.6509 | 0.7515 | -112 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | aVF | FALSE | 2 | 1 | 0.9968 | 0.8002 | 3.222 | 5.954 | 0.0566 | 0.9226 | 0.96 | FALSE | 2.48 | 0.75 | -2 | 0.9352 | 0.959 | 14 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | V3 | FALSE | 2 | 2 | 0.9984 | 0.8423 | 5.59 | 5.709 | 0.1264 | 0.8213 | 0.9144 | FALSE | 2.5 | -1.25 | -18 | 0.9348 | 0.9081 | 2 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | V6 | FALSE | 2 | 3 | 0.9992 | 0.4329 | -0.493 | 2.983 | 0.1435 | 0.454 | 0.5906 | FALSE | 2.422 | 3.25 | 10 | 0.9772 | 0.9721 | 28 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | 20261015_154119 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2804x1985 | ticks 9.8823 px/mm, 8/9 ticks, t0 from ticks | corners=0 rectify=180 grid=2065 locate=312 calibrate+segment=13 trace=19885 cut=10 quality=1 score*=8 classify=56 | baseline=68.0/44.1 corners=68.0/44.1 rectify=120.0/44.2 grid=140.0/91.0 locate=119.3/58.4 calibrate+segment=91.3/58.8 trace=513.4/67.7 cut=450.4/45.0 quality=450.4/45.0 score*=449.8/46.5 classify=460.3/48.1 | 561.8 | unreliable | T0_GUARD/severe | II | TRUE | 3 | 0 | 0.9994 | 0.368 | -1.291 | 0.069 | 0.1219 | 0.4068 | 0.3857 |  | 9.852 | 1.5 | 76 | 0.69 | 0.866 | 32 | ok | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | 8 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | I | FALSE | 0 | 0 | 0.9984 | 0.7801 | 3.912 | 4.131 | 0.1045 | 0.751 | 0.9133 | FALSE | 2.444 | 2.25 | 28 | 0.9518 | 0.9124 | -4 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | aVR | FALSE | 0 | 1 | 0.9984 | 0.9521 | 10.307 | 10.72 | 0.0407 | 0.9488 | 0.9963 | FALSE | 2.498 | 0 | -2 | 0.9566 | 0.9964 | 2 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | V1 | FALSE | 0 | 2 | 0.9992 | 0.8799 | 6.217 | 10.724 | 0.1374 | 0.8755 | 0.9573 | FALSE | 2.492 | 0 | 8 | 0.9575 | 0.9568 | -8 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | V4 | FALSE | 0 | 3 | 0.9984 | 0.1526 | -1.928 | 2.235 | 0.2237 | 0.1524 | 0.427 | FALSE | 2.414 | 3.5 | -2 | 0.9687 | 0.8967 | 52 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | II | FALSE | 1 | 0 | 0.9904 | 0.7397 | 2.562 | 2.764 | 0.0798 | 0.8063 | 0.9087 | FALSE | 2.5 | -2.75 | -36 | 0.8986 | 0.8946 | -8 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | aVL | FALSE | 1 | 1 | 0.9992 | 0.7811 | 3.826 | 5.891 | 0.1001 | 0.765 | 0.956 | FALSE | 2.498 | -0.5 | 2 | 0.8763 | 0.9557 | -8 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | V2 | FALSE | 1 | 2 | 0.9968 | 0.9395 | 9.49 | 9.615 | 0.1327 | 0.9307 | 0.9574 | FALSE | 2.5 | -1.25 | -14 | 0.9861 | 0.9478 | -2 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | V5 | FALSE | 1 | 3 | 0.9992 | 0.1392 | -2.057 | 0.263 | 0.2053 | 0.1406 | 0.3568 | FALSE | 2.392 | 4.5 | 4 | 0.9898 | 0.9559 | 68 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | III | FALSE | 2 | 0 | 1 | 0.0833 | -4.243 | -1.977 | 0.1888 | 0.1171 | 0.3447 | FALSE | 2.41 | 3.25 | 90 | 0.6941 | 0.8079 | -56 | UNRELIABLE | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | aVF | FALSE | 2 | 1 | 0.9992 | 0.8414 | 4.809 | 8.666 | 0.0471 | 0.8945 | 0.9496 | FALSE | 2.494 | 0 | 6 | 0.936 | 0.9506 | -6 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | V3 | FALSE | 2 | 2 | 0.9968 | 0.7931 | 4.325 | 4.386 | 0.1463 | 0.7907 | 0.9466 | FALSE | 2.5 | -1.5 | -20 | 0.9692 | 0.9419 | 2 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | V6 | FALSE | 2 | 3 | 0.9984 | 0.8051 | 4.408 | 10.303 | 0.0816 | 0.7938 | 0.9661 | FALSE | 2.476 | 1 | 0 | 0.979 | 0.9642 | 12 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | 20261015_154211 | ptbxl40_kit_500hz_300dpi | 40 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2863x2010 | ticks 10.2094 px/mm, 9/9 ticks, t0 from ticks | corners=0 rectify=184 grid=2211 locate=604 calibrate+segment=11 trace=20388 cut=17 quality=1 score*=18 classify=81 | baseline=68.9/43.8 corners=68.9/43.8 rectify=124.9/43.9 grid=145.1/86.3 locate=120.3/66.9 calibrate+segment=92.1/58.6 trace=514.1/68.2 cut=450.6/45.4 quality=450.5/45.5 score*=453.3/47.2 classify=459.2/48.3 | 555.7 | unreliable | T0_GUARD/severe | II | TRUE | 3 | 0 | 0.9998 | 0.5931 | 0.442 | 0.583 | 0.0999 | 0.6781 | 0.8181 |  | 9.9 | 1 | 42 | 0.7589 | 0.8723 | 6 | ok | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | 9 | ticks
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | I | FALSE | 0 | 0 | 0.9984 | 0.9089 | 7.527 | 7.527 | 0.09 | 0.8893 | 0.947 | FALSE | 2.444 | 2.25 | 26 | 0.9691 | 0.9801 | 0 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | aVR | FALSE | 0 | 1 | 0.9976 | 0.9533 | 10.454 | 10.555 | 0.0506 | 0.9378 | 0.9432 | FALSE | 2.482 | 0.75 | 6 | 0.9785 | 0.943 | 2 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | V1 | FALSE | 0 | 2 | 0.9984 | 0.9867 | 15.749 | 18.074 | 0.0609 | 0.954 | 0.9724 | FALSE | 2.498 | -0.25 | 2 | 0.9946 | 0.972 | -4 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | V4 | FALSE | 0 | 3 | 1 | 0.9299 | 8.694 | 9.878 | 0.1001 | 0.9192 | 0.986 | FALSE | 2.5 | -1 | -18 | 0.9653 | 0.9852 | 8 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | II | FALSE | 1 | 0 | 0.9952 | 0.2107 | -2.828 | 7.81 | 0.1717 | 0.253 | 0.3716 | FALSE | 2.47 | -0.75 | -48 | 0.9377 | 0.943 | 40 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | aVL | FALSE | 1 | 1 | 0.9992 | 0.995 | 19.881 | 22.624 | 0.0209 | 0.967 | 0.9533 | FALSE | 2.498 | 0 | -2 | 0.9975 | 0.9534 | 2 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | V2 | FALSE | 1 | 2 | 0.9992 | 0.9909 | 17.411 | 18.245 | 0.0617 | 0.9686 | 0.9777 | FALSE | 2.488 | 0.5 | 2 | 0.997 | 0.9768 | 2 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | V5 | FALSE | 1 | 3 | 0.9912 | 0.9078 | 7.526 | 7.575 | 0.0478 | 0.8904 | 0.9905 | FALSE | 2.5 | -0.75 | -8 | 0.9256 | 0.9884 | 2 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | III | FALSE | 2 | 0 | 0.9616 | 0.0338 | -2.973 | -0.536 | 0.25 | 0.0363 | 0.8472 | TRUE | 2.49 | -3.5 | -100 | 0.9409 | 0.9485 | 78 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | aVF | FALSE | 2 | 1 | 0.9992 | 0.9395 | 9.386 | 10.674 | 0.0386 | 0.8995 | 0.9461 | FALSE | 2.496 | 0 | 4 | 0.9554 | 0.9459 | -4 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | V3 | FALSE | 2 | 2 | 0.9992 | 0.9721 | 12.777 | 17.168 | 0.0798 | 0.966 | 0.9781 | FALSE | 2.498 | -0.5 | 2 | 0.9923 | 0.9771 | -8 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | V6 | FALSE | 2 | 3 | 0.9768 | 0.8203 | 4.247 | 6.783 | 0.0568 | 0.8538 | 0.9609 | FALSE | 2.496 | -0.75 | 4 | 0.9321 | 0.9612 | -10 | ok | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
crumple rescue | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | 20261015_154533 | ptbxl4016_kit_500hz_300dpi | 4016 | model ml/b3_512x1024_fp16.tflite | user | 2867x2012 | ticks 10.3522 px/mm, 8/9 ticks, t0 from pulses | corners=0 rectify=210 grid=2639 locate=367 calibrate+segment=4 trace=20070 cut=11 quality=0 score*=7 classify=55 | baseline=68.6/26.8 corners=68.6/26.9 rectify=124.6/26.9 grid=144.8/96.2 locate=119.7/65.1 calibrate+segment=91.8/55.4 trace=513.9/79.7 cut=451.8/45.4 quality=451.8/45.4 score*=451.9/46.8 classify=462.8/48.4 | 561.5 | unreliable | SCALE_DISAGREE/severe | II | TRUE | 3 | 0 | 0.9894 | 0.44 | -0.655 | -0.55 | 0.1447 | 0.456 | 1.0317 |  | 3.238 | -4.25 | 6762 |  |  | 6 | UNRELIABLE | grid (fit rms 0.217 mm, 55x38 rules, 10.2244 px/mm, inliers 0.675) | 8 | pulses
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | I | FALSE | 0 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVR | FALSE | 0 | 1 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V1 | FALSE | 0 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V4 | FALSE | 0 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | FALSE | 1 | 0 | 0.9664 | 0.3114 | -14.207 | -13.921 | 0.5424 | 1.7077 | 0.6833 | TRUE | 2.4 | -2.75 | 100 | 0.4051 | 0.8438 | -126 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVL | FALSE | 1 | 1 | 0.932 | 0.1571 | -20.66 | -20.365 | 0.6132 | 1.6895 | 0.083 | FALSE | 2.266 | 8.75 | -32 | 0.2267 | 0.7366 | -198 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V2 | FALSE | 1 | 2 | 0.5264 | -0.0713 | -1.673 | -1.545 | 0.3901 | -0.0366 | 0.0367 | FALSE | 2.5 | -10 | -80 | -0.0699 | 0.0366 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V5 | FALSE | 1 | 3 | 0.8024 | -0.1566 | -5.305 | -5.113 | 0.3147 | -0.2247 | 0.0093 | TRUE | 2.172 | 10 | -100 | -0.0381 | 0.1575 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | III | FALSE | 2 | 0 | 0.088 | 0.2189 | -6.316 | -6.164 | 0.1354 | 0.3264 | -0.2254 | TRUE | 2.336 | 7 | 100 | 0.227 | -0.2274 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVF | FALSE | 2 | 1 | 0.9144 | 0.0057 | -16.554 | -16.377 | 0.5857 | 0.0391 | 0.089 | TRUE | 2.32 | 7.75 | 100 | 0.0817 | 0.1149 | -78 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V3 | FALSE | 2 | 2 | 0.3288 | 0.1041 | -2.271 | -2.185 | 0.3339 | 0.0847 | -0.1447 | TRUE | 2.178 | 9.75 | -100 | 0.152 | 0.0026 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V6 | FALSE | 2 | 3 | 0.2256 | -0.2253 | -1.272 | -0.91 | 0.142 | -0.0969 | 0 | FALSE | 2.224 | 9.5 | -60 | 0.1853 | 0.1828 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261015_153112.csv | 20261015_153112 | ptbxl9_kit_500hz_300dpi | 9 | classical | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=108 grid=931 locate=423 calibrate+segment=28 trace=455 cut=5 quality=1 score*=5 classify=68 | baseline=72.5/57.6 corners=72.5/57.6 rectify=111.5/57.7 grid=121.2/73.3 locate=113.6/39.8 calibrate+segment=86.9/33.6 trace=91.3/48.0 cut=88.0/34.5 quality=88.0/34.5 score*=88.0/35.4 classify=96.8/36.6 | 205.7 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | TRUE | 3 | 0 | 0.7822 | -0.0596 | -22.608 | -22.531 | 1.6065 | -0.809 | 0.0783 |  | 1.268 | 5.25 | 8732 |  |  |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | I | FALSE | 0 | 0 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVR | FALSE | 0 | 1 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V1 | FALSE | 0 | 2 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V4 | FALSE | 0 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | FALSE | 1 | 0 | 0.9976 | 0.1393 | -15.773 | -15.033 | 0.6495 | 0.8854 | 1.219 | TRUE | 2.418 | -0.75 | -100 | 0.3513 | 0.9 | 158 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVL | FALSE | 1 | 1 | 0.9952 | -0.0149 | -20.819 | -20.588 | 0.6246 | -0.1629 | -0.3364 | FALSE | 2.464 | -8.75 | 36 | 0.2781 | 0.0889 | -96 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V2 | FALSE | 1 | 2 | 0.4784 | 0.0819 | -4.814 | -4.765 | 0.56 | 0.1074 | -0.0145 | TRUE | 2.4 | -10 | 100 | 0.1641 | 0.165 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V5 | FALSE | 1 | 3 | 0 | [13 empty] | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | III | FALSE | 2 | 0 | 0.9976 | -0.0863 | -20.071 | -19.811 | 0.6596 | -0.8503 | 0.7239 | TRUE | 2.198 | 8.75 | -100 | 0.0847 | 1.0983 | 172 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | aVF | FALSE | 2 | 1 | 0.9976 | 0.0101 | -18.51 | -18.046 | 0.7336 | 0.0871 | 0.7982 | TRUE | 2.4 | -10 | 100 | 0.1384 | 1.2337 | 2 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V3 | FALSE | 2 | 2 | 0.9976 | -0.0638 | -8.321 | -8.027 | 0.6701 | -0.1522 | 0.0423 | TRUE | 2.258 | 6 | -100 | 0.1216 | 0.3115 | 148 | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | V6 | FALSE | 2 | 3 | 0.6184 | -0.1837 | -11.063 | -10.345 | 0.4384 | -0.5741 | 0.0262 | FALSE | 2.5 | -4 | -92 | 0.1434 | 2.1937 |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks
tilt rescue | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261015_153108.csv | 20261015_153108 | ptbxl9_kit_500hz_300dpi | 9 | model ml/b3_512x1024_fp16.tflite | automatic (refined=false) | 2000x1900 | ticks 7.2980 px/mm, 3/9 ticks, t0 from ticks | corners=0 rectify=77 grid=904 locate=408 calibrate+segment=27 trace=20038 cut=24 quality=8 score*=13 classify=89 | baseline=60.3/31.4 corners=60.3/31.4 rectify=99.2/31.4 grid=114.3/47.1 locate=106.9/48.0 calibrate+segment=80.5/21.7 trace=502.2/64.9 cut=439.4/41.6 quality=439.4/41.6 score*=439.4/42.5 classify=448.2/43.7 | 534.4 | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; T0_GUARD/severe; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | II | TRUE | 3 | 0 | 0.364 | -0.2718 | -18.708 | -18.604 | 1.0254 | -1.9232 | 0 |  | 6.7 | 7.75 | -2580 |  |  |  | UNRELIABLE | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | 3 | ticks

## rescue_before_after
photo | tracer | before_pair_id | after_pair_id | note | before_export | before_target | before_scored_against | before_ticks_found | before_t0_source | before_tick_vs_grid_pct | before_quality | before_t0_guard | before_median_pcc | before_median_pcc_aligned | before_median_snr_db | before_median_shifted_snr_db | before_median_r_ratio | before_leads_pcc_ge_0.9 | before_rhythm_pcc | after_export | after_target | after_scored_against | after_ticks_found | after_t0_source | after_tick_vs_grid_pct | after_quality | after_t0_guard | after_median_pcc | after_median_pcc_aligned | after_median_snr_db | after_median_shifted_snr_db | after_median_r_ratio | after_leads_pcc_ge_0.9 | after_rhythm_pcc | delta_median_pcc
eacg_20261015_124720.jpg | learned | 20261015_124807 | 20261015_153336 | guard used | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 3 | pulses | 11.7 | unreliable |  | 0.126 | 0.294 | -1.81 | -1.44 | 0.017 | 0 | 0.129 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_153336.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 8 | ticks | -0.5 | unreliable | fired, used: t0 x 90.8 -> 156.9, ticks 3 -> 8, tick vs grid +11.7% -> -0.5% (t0 from the ticks) | 0.803 | 0.974 | 3.83 | 6.37 | 0.894 | 2 | 0.3705 | 0.677
eacg_20261015_124720.jpg | classical | 20261015_124807 | 20261015_153336 | guard used | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 3 | pulses | 11.7 | unreliable |  | 0.082 | 0.215 | -2.31 | -2.05 | 0.017 | 0 | 0.1401 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_153340.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 8 | ticks | -0.5 | unreliable | fired, used: t0 x 90.8 -> 156.9, ticks 3 -> 8, tick vs grid +11.7% -> -0.5% (t0 from the ticks) | 0.493 | 0.815 | 0.29 | 2.26 | 0.39 | 0 | 0.391 | 0.411
eacg_20261015_124821.jpg | learned | 20261015_124910 | 20261015_154119 | guard used | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 3 | pulses | 18.3 | unreliable |  | -0.048 | 0.297 | -2.19 | -1.57 | 0.022 | 0 | 0.0712 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154119.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 8 | ticks | -0.2 | unreliable | fired, used: t0 x 68.8 -> 178.7, ticks 3 -> 8, tick vs grid +18.3% -> -0.2% (t0 from the ticks) | 0.8 | 0.973 | 3.84 | 5.83 | 0.878 | 2 | 0.368 | 0.848
eacg_20261015_124821.jpg | classical | 20261015_124910 | 20261015_154119 | guard used | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 3 | pulses | 18.3 | unreliable |  | 0.026 | 0.247 | -0.78 | -0.65 | 0.021 | 0 | -0.0303 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154123.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 8 | ticks | -0.2 | unreliable | fired, used: t0 x 68.8 -> 178.7, ticks 3 -> 8, tick vs grid +18.3% -> -0.2% (t0 from the ticks) | 0.349 | 0.438 | 0.5 | 0.63 | -0.002 | 0 | 0.4204 | 0.323
eacg_20261015_125056.jpg | learned | 20261015_125142 | 20261015_154211 | guard used | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 3 | rhythm_locator | 20.1 | unreliable |  | 0.022 | 0.292 | -2.85 | -2.12 | 0.032 | 0 | 0.0347 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154211.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 9 | ticks | -0.2 | unreliable | fired, used: t0 x 30.0 -> 158.3, ticks 3 -> 9, tick vs grid +20.1% -> -0.2% (t0 from the ticks) | 0.787 | 0.957 | 4.12 | 5.14 | 0.948 | 2 | 0.5931 | 0.765
eacg_20261015_125056.jpg | classical | 20261015_125142 | 20261015_154211 | guard used | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 3 | rhythm_locator | 20.1 | unreliable |  | 0.04 | 0.151 | -4.52 | -4.01 | 0 | 0 |  | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154216.csv | ptbxl40_kit_500hz_300dpi | ptbxl40_kit_500hz_300dpi | 9 | ticks | -0.2 | unreliable | fired, used: t0 x 30.0 -> 158.3, ticks 3 -> 9, tick vs grid +20.1% -> -0.2% (t0 from the ticks) | 0.626 | 0.775 | 1.88 | 2.1 | 0.695 | 1 |  | 0.586
eacg_20261015_125342.jpg | learned | 20261015_125428 | 20261015_155451 | guard used | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | ptbxl4016_kit_500hz_300dpi | ptbxl4016_kit_500hz_300dpi | 2 | pulses |  | unreliable |  | 0.014 | 0.601 | -2.87 | -1.94 | 0.044 | 0 | 0.0774 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_155451.csv | ptbxl4016_kit_500hz_300dpi | ptbxl4016_kit_500hz_300dpi | 8 | ticks | 0 | unreliable | fired, used: t0 x 91.2 -> 164.7, ticks 2 -> 8, tick vs grid nan% -> +0.0% (t0 from the ticks) | 0.936 | 0.971 | 8.89 | 10.31 | 0.972 | 7 | 0.4693 | 0.922
eacg_20261015_125342.jpg | classical | 20261015_125428 | 20261015_155451 | guard used | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | ptbxl4016_kit_500hz_300dpi | ptbxl4016_kit_500hz_300dpi | 2 | pulses |  | unreliable |  | 0.025 | 0.427 | -2.56 | -2.15 | 0.027 | 0 | 0.0844 | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_155455.csv | ptbxl4016_kit_500hz_300dpi | ptbxl4016_kit_500hz_300dpi | 8 | ticks | 0 | unreliable | fired, used: t0 x 91.2 -> 164.7, ticks 2 -> 8, tick vs grid nan% -> +0.0% (t0 from the ticks) | 0.687 | 0.867 | 2.56 | 4.08 | 0.839 | 4 | 0.4773 | 0.662
eacg_20261015_125342.jpg | learned | 20261015_125428 | 20261015_154533 | second run: TR corner by hand; guard did not fire; scored against PTB-XL 4016 (labelled 40, python/target_corrections.csv) | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | ptbxl4016_kit_500hz_300dpi | ptbxl4016_kit_500hz_300dpi | 2 | pulses |  | unreliable |  | 0.014 | 0.601 | -2.87 | -1.94 | 0.044 | 0 | 0.0774 | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_154533.csv | ptbxl40_kit_500hz_300dpi | ptbxl4016_kit_500hz_300dpi | 8 | pulses | 1.2 | unreliable |  | 0.935 | 0.967 | 9.04 | 10.22 | 0.957 | 9 | 0.44 | 0.921
eacg_20261015_125342.jpg | classical | 20261015_125428 | 20261015_154533 | second run: TR corner by hand; guard did not fire; scored against PTB-XL 4016 (labelled 40, python/target_corrections.csv) | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | ptbxl4016_kit_500hz_300dpi | ptbxl4016_kit_500hz_300dpi | 2 | pulses |  | unreliable |  | 0.025 | 0.427 | -2.56 | -2.15 | 0.027 | 0 | 0.0844 | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_154538.csv | ptbxl40_kit_500hz_300dpi | ptbxl4016_kit_500hz_300dpi | 8 | pulses | 1.2 | unreliable |  | 0.625 | 0.789 | 0.91 | 3.17 | 0.82 | 2 | 0.3653 | 0.6