## README
item | what it is / how it is computed
WHAT THIS WORKBOOK IS
Set A | The 12 printed test sheets as clean digital images (3200 x 2304 px), picked from the phone's gallery: the app with no camera in the way. 12 images x 2 tracers.
Set B | The same 12 sheets printed and photographed with the app in good indoor light ('ideal'): 24 photographs (2 per sheet) x 2 tracers, each pair from ONE Digitise tap in both-tracers mode (pair_id).
Set C | Robustness: the printed sheets photographed under tilt (phone ~25-35 deg from parallel), hard shadow (one warm lamp) and crumpling, digitised on the A30s, both tracers where run (three failed runs are learned-only). Runs -> photographs: the README in each data/device_exports/Set C folder. Scored by the same scorer (score_device_set.py); the same statistics as A and B. The rescue re-runs with the t0 guard are NOT in this set (in-sample; results/device/setC_crumple_rescue_before_after.csv).
Tracers | learned = b3, the CNN (LiteRT FP16) that reads each row; classical = the threshold-and-centroid control. Everything else in the pipeline is identical for both.
WHAT THE PHONE WRITES (each CSV)
Header lines | '#' lines: device, photo, target sheet, tracer, corners, geometry (grid fit), calibration (px/mm, ticks found of 9, time-zero source), per-stage times (ms), per-stage peak memory, quality verdict and warnings, the classifier's probabilities, pair_id, photo_rotation.
Time series | time_s + 13 columns: I, II, III, aVR, aVL, aVF, V1-V6 and II_rhythm, at 500 Hz, 5,000 rows (10 s), in mV. A 3x4 lead has values only in its own 2.5 s cell (1,250 samples: I/II/III 0-2.5 s, aVR/aVL/aVF 2.5-5 s, V1-V3 5-7.5 s, V4-V6 7.5-10 s); empty elsewhere. II_rhythm has all 10 s. Empty = NaN = the tracer saw no trace there.
How the phone made them | Each traced row (page pixels) is cut on the row's time knots (the printed separator ticks), each pixel column mapped to seconds, then resampled to the 500 Hz grid by linear interpolation (np.interp / SheetLayout.cellSignal). Amplitude: mV = -(row - median row of the cell) / (px per mm x 10 mm/mV); so every cell is about its own median (a printout has no absolute baseline).
WHAT THE SCORER DOES (score_device_sheets.py -> score_12lead.score_lead)
Truth | The PTB-XL record's own 500 Hz signal (the one the kit printed), cut to exactly the same window: for a 3x4 lead the 1,250 samples the kit drew in that cell; for the rhythm strip all 5,000. No resampling of the truth: both are 500 Hz, sample k of the cell against sample k of the truth.
Gaps | Samples the tracer left empty are filled by linear interpolation between their neighbours (run_photo.fill_nan) before every metric except coverage. A lead with < 50 finite samples is not scored at all: it appears with an empty PCC and counts as a FAILURE.
coverage | Fraction of the cell's samples that are finite (before filling).
pcc_asis | Pearson correlation of truth and digitised over the cell, as cut (no shift, no rescale): np.corrcoef. Offset- and scale-invariant: it measures SHAPE AND TIMING, not amplitude.
snr_db_asis | 10 log10( sum truth^2 / sum (truth - digitised)^2 ), both with their own median removed. Sensitive to shape, timing AND amplitude.
shifted_snr_db_asis | The best SNR over integer shifts of -25..+25 samples (+-50 ms at 500 Hz), each scored on the overlapping part only: the PhysioNet Challenge 2024 convention, which forgives a small time offset.
rmse_mv_asis | sqrt(mean (truth - digitised)^2) in mV, both with their own median removed (as SNR; since 2026-10-13 -- before, the truth's raw baseline offset counted as error; progress_log §37-38).
gain_ratio | Least-squares slope of digitised on truth: cov(truth, dig) / var(truth). 1.0 = amplitude right; 0.9 = 10% low. Intercept fitted, so a baseline error does not affect it.
r_ratio | R-peak height ratio: truth peaks = local maxima above half the largest deflection (polarity taken from the truth, so a downward QRS is scored on its S wave), >= 0.25 s apart; heights from each signal's median; for each peak the digitised value is its largest same-polarity deviation within +-25 ms; the median of digitised / truth over the peaks.
pcc_aligned / align_* | The best PCC after letting the digitised signal be rescaled in time (x0.90 to x1.10 in 0.25% steps, np.interp) and shifted (+-100 ms for a 2.5 s cell; any lag by FFT cross-correlation for the 10 s strip). It shows what the trace would score with a perfect time base: high aligned PCC with lower plain PCC = a TIMING error, not a reading error. Withheld (empty, 'n.m.') when the aligned overlap is < 85% of the window.
shift_ms / shift_status | The integer lag (+-200 ms) maximising PCC of truth against digitised, medians removed, NaN pairs skipped, >= 85% overlap; status UNRELIABLE if that PCC < 0.5, FAILED at the search bound. A diagnostic of where the time error is.
HOW THE SUMMARIES ARE MADE
Unit | The LEAD: 12 per photo (the rhythm strip is reported separately). Medians and quartiles over leads; a lead with no score is a failure: counted in n, n_unavailable and as 'not >= 0.9', excluded from the median.
95% CI | Of a median: CLUSTER bootstrap over photos (leads of one photo are not independent), 2,000 resamples, seed 20261012 (make_report_figures.cluster_ci).
Learned vs classical | Paired by (photo, lead); Wilcoxon signed-rank on the PCC difference.
Times | Per export, from the header's timings_ms; 'pipeline' = every stage except score* (scoring against the truth is test harness, not the app). Memory: the peak native heap per stage (MB), from memory_peak_mb.
Photo quality (Set B) | Measured on each photograph here: sheet fill (quad area / frame), corner angles' largest departure from 90 degrees (the phone not held parallel), keystone (longer / shorter of top and bottom edges), light ratio (5th / 95th percentile of the paper's smoothed luma; 1 = perfectly even), sharpness (variance of the Laplacian), clipped fraction (pixels >= 250).

## Summary 12 leads
set_name | tracer_kind | photos | leads | leads_unavailable | pct_pcc_ge_0_9 | pcc_median | pcc_q1 | pcc_q3 | pcc_ci95_lo | pcc_ci95_hi | pcc_aligned_median | snr_db_asis_median | shifted_snr_db_asis_median | rmse_mv_asis_median | gain_ratio_median | r_ratio_median | coverage_median | pcc_aligned_not_measurable
C (robustness) | classical | 14 | 168 | 14 | 4.2 | 0.4477 | 0.13935 | 0.6934 | 0.19299 | 0.5576 | 0.5369 | 0.373 | 0.5655 | 0.1637 | 0.30955 | 0.1027 | 0.958 | 0
C (robustness) | learned | 17 | 204 | 13 | 40.2 | 0.8483 | 0.101 | 0.96355 | 0.1511 | 0.9391 | 0.9665 | 5.378 | 7.936 | 0.0961 | 0.8318 | 0.899 | 0.9976 | 0

## Summary rhythm strip
set_name | tracer_kind | photos | leads | leads_unavailable | pct_pcc_ge_0_9 | pcc_median | pcc_q1 | pcc_q3 | pcc_ci95_lo | pcc_ci95_hi | pcc_aligned_median | snr_db_asis_median | shifted_snr_db_asis_median | rmse_mv_asis_median | gain_ratio_median | r_ratio_median | coverage_median | pcc_aligned_not_measurable
C (robustness) | classical | 14 | 14 | 1 | 0 | 0.476 | 0.3139 | 0.6483 | 0.3139 | 0.6483 | 0.6937 | 0.033 | 0.054 | 0.1335 | 0.4884 | 0.5099 | 0.915 | 4
C (robustness) | learned | 17 | 17 | 0 | 5.9 | 0.4935 | 0.0712 | 0.8204 | 0.0712 | 0.8204 | 0.8634 | -0.369 | -0.029 | 0.1362 | 0.502 | 0.8155 | 0.9902 | 6

## Summary by condition (C)
set_name | condition | tracer_kind | photos | leads | leads_unavailable | pct_pcc_ge_0_9 | pcc_median | pcc_q1 | pcc_q3 | pcc_ci95_lo | pcc_ci95_hi | pcc_aligned_median | snr_db_asis_median | shifted_snr_db_asis_median | rmse_mv_asis_median | gain_ratio_median | r_ratio_median | coverage_median | pcc_aligned_not_measurable
C (robustness) | crumple | classical | 7 | 84 | 8 | 4.8 | 0.1826 | 0.025575 | 0.49493 | 0.0449 | 0.50525 | 0.3526 | -0.887 | -0.6535 | 0.18815 | 0.1352 | 0.03485 | 0.956 | 0
C (robustness) | crumple | learned | 7 | 84 | 0 | 21.4 | 0.21055 | 0 | 0.88825 | 0.013189 | 0.88205 | 0.806 | -1.2755 | -0.664 | 0.1654 | 0.2116 | 0.17075 | 0.9976 | 0
C (robustness) | shadow | classical | 3 | 36 | 1 | 8.3 | 0.5556 | 0.38385 | 0.74395 | 0.54875 | 0.74595 | 0.5899 | 0.98 | 1.125 | 0.1196 | 0.4771 | 0.3936 | 0.9596 | 0
C (robustness) | shadow | learned | 4 | 48 | 3 | 50 | 0.9321 | 0.1927 | 0.9805 | 0.1036 | 0.9738 | 0.9918 | 8.827 | 13.094 | 0.0775 | 0.9101 | 0.9273 | 0.9976 | 0
C (robustness) | tilt | classical | 4 | 48 | 5 | 0 | 0.5627 | 0.4366 | 0.72825 | 0.4419 | 0.5919 | 0.5788 | 1.359 | 1.485 | 0.1411 | 0.391 | 0.1759 | 0.9592 | 0
C (robustness) | tilt | learned | 6 | 72 | 10 | 55.6 | 0.9461 | 0.76992 | 0.97425 | 0.1883 | 0.96401 | 0.9852 | 9.9125 | 13.2865 | 0.05395 | 0.91865 | 0.91745 | 0.9968 | 0

## Tracer comparison
set | n_pairs | statistic | p | median_diff
C (robustness) | 154 | 1536 | 1.3089e-15 | 0.18685

## Per photo
set | ecg_id | ptbxl_label | photo_id | tracer | export | photo | photo_path | photo_rotation | corners | geometry | calibration | quality | warnings | pcc_median_12 | pcc_min_12 | leads_pcc_ge_0_9 | leads_unavailable | pcc_aligned_median_12 | snr_db_median_12 | shifted_snr_db_median_12 | rmse_mv_median_12 | gain_median_12 | pcc_rhythm | pipeline_ms | trace_ms | peak_native_mb | pss_end_mb | classifier_flags | p_NORM | p_MI | p_STTC | p_CD | p_HYP
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.262 mm, 54x48 rules, 7.0180 px/mm, inliers 0.538) | ticks 9.0720 px/mm, 3/9 ticks, t0 from rhythm_locator | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | -0.05795 | -0.213 | 0 | 4 | 0.09915 | -10.593 | -10.224 | 0.5708 | -0.10775 | -0.3106 | 21865 | 20345 | 501 | 534.7 | MI CD | 0.059 | 0.7678 | 0.3604 | 0.0901 | 0.1141
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm | 0.5919 | 0.1893 | 0 | 1 | 0.642 | 1.663 | 1.692 | 0.1154 | 0.3949 | 0.6937 | 1600 | 182 | 119.9 | 193.8 | MI CD | 0.1436 | 0.8417 | 0.0589 | 0.1872 | 0.0298
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.226 mm, 56x38 rules, 6.4000 px/mm, inliers 0.955) | ticks 6.4344 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9408 | 0.8099 | 9 | 0 | 0.985 | 9.2865 | 14.6105 | 0.0454 | 0.8873 | 0.9057 | 20797 | 19433 | 500 | 531.7 | NORM | 0.9253 | 0.0598 | 0.0143 | 0.0191 | 0.0081
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | camera 3200x2396 |  |  | automatic (refined=false) | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4 | 0.3828 | 0.1115 | 0 | 0 | 0.42545 | 0.6735 | 0.757 | 0.15025 | 0.18745 | 0.476 | 2734 | 347 | 134.6 | 214.6 | MI CD | 0.0533 | 0.7584 | 0.159 | 0.5912 | 0.0596
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | camera 3200x2396 |  |  | automatic (refined=false) | grid (fit rms 0.241 mm, 54x38 rules, 9.2106 px/mm, inliers 0.797) | ticks 9.3132 px/mm, 9/9 ticks, t0 from pulses | unreliable | SCALE_DISAGREE/severe | 0.8283 | 0.3585 | 4 | 0 | 0.9794 | 5.384 | 9.2165 | 0.09775 | 0.7267 | 0.7321 | 22752 | 20295 | 502.4 | 531.0 | MI CD | 0.3589 | 0.4961 | 0.0721 | 0.0673 | 0.0304
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | camera 3200x2396 |  | 270 clockwise (auto 270) | automatic (refined=false) | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5 | 0.0824 | -0.1755 | 0 | 0 | 0.2151 | -2.3085 | -2.0495 | 0.2077 | 0.07805 | 0.1401 | 2929 | 374 | 139.9 | 207.5 | MI CD | 0.1694 | 0.4623 | 0.3664 | 0.1205 | 0.029
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | camera 3200x2396 |  | 270 clockwise (auto 270) | automatic (refined=false) | grid (fit rms 0.248 mm, 55x37 rules, 9.4578 px/mm, inliers 0.565) | ticks 10.5629 px/mm, 3/9 ticks, t0 from pulses | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 0.1264 | -0.175 | 0 | 0 | 0.2936 | -1.811 | -1.4425 | 0.20165 | 0.1076 | 0.129 | 22357 | 20003 | 508 | 546.7 | MI CD | 0.222 | 0.6712 | 0.1053 | 0.0574 | 0.0172
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | camera 3200x2396 |  | 270 clockwise (auto 270) | automatic (refined=false) | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 0.0264 | -0.198 | 0 | 0 | 0.2469 | -0.78 | -0.6535 | 0.18815 | 0.0137 | -0.0303 | 3071 | 427 | 145.6 | 218.4 | MI CD | 0.0461 | 0.7833 | 0.2601 | 0.0676 | 0.0157
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | camera 3200x2396 |  | 270 clockwise (auto 270) | automatic (refined=false) | grid (fit rms 0.241 mm, 54x38 rules, 9.9007 px/mm, inliers 0.536) | ticks 11.7158 px/mm, 3/9 ticks, t0 from pulses | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | -0.04845 | -0.2347 | 0 | 0 | 0.29705 | -2.189 | -1.5685 | 0.2397 | -0.04625 | 0.0712 | 22181 | 19553 | 512.6 | 551.8 | MI CD | 0.037 | 0.9024 | 0.1745 | 0.0632 | 0.0207
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | camera 3200x2396 |  |  | automatic (refined=false) | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | 0.03985 | -0.1183 | 0 | 8 | 0.15055 | -4.5225 | -4.0065 | 0.46735 | 0.10195 |  | 3689 | 444 | 147.7 | 217.8 | MI CD | 0.004 | 0.9767 | 0.311 | 0.0579 | 0.0212
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | camera 3200x2396 |  |  | automatic (refined=false) | grid (fit rms 0.224 mm, 53x38 rules, 10.2268 px/mm, inliers 0.665) | ticks 12.2800 px/mm, 3/9 ticks, t0 from rhythm_locator | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 0.02215 | -0.3068 | 0 | 0 | 0.29175 | -2.849 | -2.121 | 0.27845 | 0.01915 | 0.0347 | 23224 | 20110 | 510.6 | 553.2 | MI | 0.0138 | 0.8618 | 0.2413 | 0.0494 | 0.0301
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | camera 3200x2396 |  | 180 clockwise (auto 270, user 270) | user | grid (fit rms 0.152 mm, 54x37 rules, 9.5593 px/mm, inliers 0.976) | duration 9.3960 px/mm, 1/9 ticks, t0 from rhythm_locator | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4 | 0.0997 | -0.0746 | 0 | 3 | 0.29 | -4.213 | -4.132 | 0.2524 | 0.1454 | 0.0477 | 21923 | 19734 | 509.3 | 547.8 | MI CD | 0.016 | 0.936 | 0.1212 | 0.1479 | 0.0146
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | camera 2396x3200 |  |  | user | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II | 0.74595 | 0.0616 | 1 | 0 | 0.77125 | 3.5575 | 3.8715 | 0.10715 | 0.59125 | 0.7455 | 2341 | 337 | 124 | 203.4 | CD | 0.2422 | 0.3453 | 0.1899 | 0.1264 | 0.0632
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | camera 2396x3200 |  |  | user | grid (fit rms 0.160 mm, 52x37 rules, 8.2495 px/mm, inliers 0.959) | ticks 8.2562 px/mm, 7/9 ticks, t0 from pulses | ok |  | 0.93575 | 0.793 | 7 | 0 | 0.9938 | 9.231 | 14.5635 | 0.07085 | 0.914 | 0.8204 | 21925 | 19940 | 498 | 534.7 | NORM CD | 0.6263 | 0.0737 | 0.1156 | 0.099 | 0.1372
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.254 mm, 53x39 rules, 7.0082 px/mm, inliers 0.584) | ticks 8.5440 px/mm, 3/9 ticks, t0 from rhythm_locator | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm | 0.084 | -0.1518 | 0 | 6 | 0.19835 | -4.9625 | -4.4645 | 0.2451 | 0.10865 | -0.1681 | 20551 | 19428 | 509.6 | 543.3 | MI STTC | 0.0089 | 0.8608 | 0.505 | 0.0376 | 0.0275
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | ok |  | 0.58555 | 0.4424 | 0 | 0 | 0.6202 | 1.748 | 1.8765 | 0.13305 | 0.48985 | 0.7643 | 1427 | 188 | 118.6 | 197.3 | MI CD | 0.0607 | 0.814 | 0.1788 | 0.1789 | 0.0425
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.237 mm, 54x37 rules, 6.9000 px/mm, inliers 0.965) | ticks 6.9269 px/mm, 8/9 ticks, t0 from pulses | ok |  | 0.96065 | 0.932 | 12 | 0 | 0.99025 | 11.2055 | 16.148 | 0.0372 | 0.9462 | 0.8603 | 20653 | 19422 | 500.2 | 586.4 | NORM CD | 0.538 | 0.1347 | 0.1295 | 0.0931 | 0.1111
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | 0.60085 | -0.2094 | 0 | 4 | 0.6408 | 0.966 | 1.0235 | 0.143 | 0.41385 | 0.6483 | 1426 | 184 | 115.2 | 185.0 | MI CD | 0.0007 | 0.9966 | 0.1591 | 0.0511 | 0.01
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.240 mm, 56x38 rules, 6.4105 px/mm, inliers 0.978) | ticks 6.4224 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.97275 | 0.7566 | 9 | 0 | 0.98815 | 12.8285 | 16.159 | 0.03635 | 0.9248 | 0.8937 | 20685 | 19434 | 495 | 529.8 | MI | 0.0156 | 0.9818 | 0.0579 | 0.0151 | 0.0109
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | camera 3200x2396 |  |  | automatic (refined=false) | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | check | NO_PULSE/warn | 0.4392 | 0.069 | 0 | 0 | 0.51875 | 0.258 | 0.457 | 0.1983 | 0.32925 | 0.4088 | 3516 | 369 | 151.8 | 222.2 | MI CD | 0.0065 | 0.9838 | 0.1647 | 0.1119 | 0.0554
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | camera 3200x2396 |  |  | automatic (refined=false) | grid (fit rms 0.219 mm, 55x38 rules, 10.0859 px/mm, inliers 0.696) | ticks 10.1260 px/mm, 8/9 ticks, t0 from rhythm_locator | check | NO_PULSE/warn | 0.9318 | 0.5255 | 7 | 0 | 0.96705 | 8.609 | 10.517 | 0.0639 | 0.9298 | 0.4935 | 24436 | 21244 | 513.3 | 555.9 | MI CD | 0.0001 | 0.993 | 0.0795 | 0.4132 | 0.405
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | camera 3200x2396 |  | 90 clockwise (auto 90) | user | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR | 0.02545 | -0.2346 | 0 | 0 | 0.42745 | -2.563 | -2.1545 | 0.2981 | 0.0192 | 0.0844 | 3700 | 455 | 148.6 | 226.3 | MI CD | 0.0014 | 0.9908 | 0.2286 | 0.0693 | 0.0969
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | camera 3200x2396 |  | 90 clockwise (auto 90) | user | grid (fit rms 0.216 mm, 55x38 rules, 10.2081 px/mm, inliers 0.680) | duration 10.9351 px/mm, 2/9 ticks, t0 from pulses | unreliable | NO_TICK_SCALE/severe | 0.01365 | -0.2393 | 0 | 0 | 0.60085 | -2.8735 | -1.945 | 0.2944 | 0.01415 | 0.0774 | 23005 | 19790 | 512 | 560.8 | MI CD | 0.0002 | 0.9904 | 0.2819 | 0.205 | 0.3646
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | camera 3200x2396 |  | 270 clockwise (auto 90, user 180) | automatic (refined=false) | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR | 0.88605 | 0.3228 | 4 | 0 | 0.9417 | 6.741 | 8.0165 | 0.08415 | 0.81445 | 0.4966 | 3795 | 393 | 146.7 | 232.1 | MI CD | 0.0001 | 0.9941 | 0.0534 | 0.4624 | 0.2932
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | camera 3200x2396 |  | 270 clockwise (auto 90, user 180) | automatic (refined=false) | grid (fit rms 0.212 mm, 55x37 rules, 10.4234 px/mm, inliers 0.663) | ticks 10.5444 px/mm, 8/9 ticks, t0 from pulses | unreliable | SCALE_DISAGREE/severe | 0.9281 | 0.2883 | 7 | 0 | 0.9636 | 8.5955 | 10.1825 | 0.06395 | 0.9054 | 0.4656 | 23075 | 19753 | 511.8 | 557.3 | MI CD | 0.0001 | 0.9948 | 0.0638 | 0.4481 | 0.3796
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | camera 2396x3200 |  |  | user | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6 | 0.54875 | 0.282 | 2 | 0 | 0.5702 | 0.182 | 0.3065 | 0.17975 | 0.52465 | 0.5824 | 2000 | 225 | 133.9 | 203.8 | MI CD | 0.003 | 0.9911 | 0.2265 | 0.124 | 0.1009
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | camera 2396x3200 |  |  | user | grid (fit rms 0.232 mm, 55x38 rules, 7.7868 px/mm, inliers 0.868) | ticks 7.8678 px/mm, 8/9 ticks, t0 from pulses | unreliable | SCALE_DISAGREE/severe | 0.9742 | 0.6152 | 10 | 0 | 0.993 | 12.89 | 18.119 | 0.0448 | 0.94395 | 0.6662 | 21312 | 19474 | 505.5 | 549.6 | MI CD | 0.0001 | 0.9805 | 0.097 | 0.509 | 0.3815
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | camera 2396x3200 |  | 90 clockwise (user 90) | user | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF | 0.5498 | 0.1021 | 0 | 1 | 0.5896 | 1.318 | 1.318 | 0.2205 | 0.3871 | 0.3139 | 2054 | 248 | 130.5 | 203.7 | MI CD | 0.0014 | 0.9775 | 0.1894 | 0.198 | 0.1669
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | camera 2396x3200 |  | 90 clockwise (user 90) | user | grid (fit rms 0.245 mm, 51x37 rules, 8.0028 px/mm, inliers 0.891) | ticks 8.0480 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6 | 0.9263 | -0.0047 | 7 | 0 | 0.9829 | 8.1915 | 13.351 | 0.08505 | 0.9017 | 0.5644 | 21466 | 19653 | 505.3 | 549.6 | MI CD | 0.0005 | 0.9898 | 0.133 | 0.2298 | 0.2363
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | check | LEAD_COVERAGE/warn/V4 | 0.43055 | 0.0373 | 0 | 0 | 0.50345 | 0.579 | 0.713 | 0.1946 | 0.17465 | 0.4221 | 1494 | 191 | 119.8 | 182.4 | MI CD | 0.0219 | 0.9144 | 0.2255 | 0.3021 | 0.0627
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | camera 2396x3200 |  |  | automatic (refined=false) | grid (fit rms 0.222 mm, 55x39 rules, 6.9340 px/mm, inliers 0.896) | ticks 6.9389 px/mm, 9/9 ticks, t0 from pulses | ok |  | 0.9679 | 0.716 | 10 | 0 | 0.98775 | 12.067 | 13.373 | 0.0531 | 0.91865 | 0.8723 | 20954 | 19615 | 501.6 | 540.4 | MI CD | 0.0001 | 0.9464 | 0.1147 | 0.615 | 0.4451

## Per sheet
set_name | ecg_id | ptbxl_label | tracer_kind | photos | leads | leads_unavailable | pct_pcc_ge_0_9 | pcc_median | pcc_q1 | pcc_q3 | pcc_ci95_lo | pcc_ci95_hi | pcc_aligned_median | snr_db_asis_median | shifted_snr_db_asis_median | rmse_mv_asis_median | gain_ratio_median | r_ratio_median | coverage_median | pcc_aligned_not_measurable
C (robustness) | 9 | NORM | classical | 1 | 12 | 1 | 0 | 0.5919 | 0.56505 | 0.7486 |  |  | 0.642 | 1.663 | 1.692 | 0.1154 | 0.3949 | 0.1759 | 0.9368 | 0
C (robustness) | 9 | NORM | learned | 2 | 24 | 4 | 37.5 | 0.8698 | 0.01275 | 0.9505 | -0.05795 | 0.9408 | 0.9688 | 6.1435 | 10.0505 | 0.0757 | 0.85245 | 0.9186 | 0.9944 | 0
C (robustness) | 40 | NORM | classical | 6 | 72 | 8 | 1.4 | 0.2105 | 0.065425 | 0.56793 | 0.0667 | 0.54385 | 0.3898 | -0.1585 | -0.0705 | 0.1544 | 0.18655 | 0.04375 | 0.9576 | 0
C (robustness) | 40 | NORM | learned | 8 | 96 | 9 | 24 | 0.1885 | 0.002 | 0.92265 | 0.01437 | 0.932 | 0.6746 | -1.309 | -0.871 | 0.1605 | 0.2115 | 0.1662 | 0.9976 | 0
C (robustness) | 1737 | MI | classical | 1 | 12 | 4 | 0 | 0.60085 | 0.38927 | 0.71108 |  |  | 0.6408 | 0.966 | 1.0235 | 0.143 | 0.41385 | 0.21945 | 0.4372 | 0
C (robustness) | 1737 | MI | learned | 1 | 12 | 0 | 75 | 0.97275 | 0.93578 | 0.97685 |  |  | 0.98815 | 12.8285 | 16.159 | 0.03635 | 0.9248 | 0.9148 | 0.9964 | 0
C (robustness) | 4016 | CD | classical | 6 | 72 | 1 | 8.3 | 0.453 | 0.26905 | 0.7177 | 0.2513 | 0.7169 | 0.5637 | 0.377 | 0.829 | 0.1995 | 0.3473 | 0.2785 | 0.9616 | 0
C (robustness) | 4016 | CD | learned | 6 | 72 | 0 | 56.9 | 0.93115 | 0.52295 | 0.9855 | 0.302 | 0.9691 | 0.97615 | 8.768 | 10.729 | 0.0715 | 0.90635 | 0.954 | 0.9976 | 0

## PCC by lead
set_name | tracer_kind | I | II | III | aVR | aVL | aVF | V1 | V2 | V3 | V4 | V5 | V6 | II rhythm
C (robustness) | classical | 0.2489 | 0.46805 | 0.4104 | 0.4733 | 0.44685 | 0.6193 | 0.3657 | 0.41345 | 0.5014 | 0.2587 | 0.5654 | 0.53705 | 0.476
C (robustness) | learned | 0.7113 | 0.5153 | 0.3585 | 0.9449 | 0.8812 | 0.9133 | 0.95175 | 0.9731 | 0.9151 | 0.8794 | 0.793 | 0.7305 | 0.4935

## SNR by lead
set_name | tracer_kind | I | II | III | aVR | aVL | aVF | V1 | V2 | V3 | V4 | V5 | V6 | II rhythm
C (robustness) | classical | 0.024 | 0.513 | 0.605 | 1.02 | -0.0015 | 0.567 | -0.284 | 0.283 | 1.2725 | -0.4005 | 1.5735 | 1.422 | 0.033
C (robustness) | learned | 2.6 | -0.698 | -1.111 | 9.69 | 6.561 | 7.797 | 10.328 | 13.0715 | 8.452 | 6.408 | 3.941 | 2.8115 | -0.369

## Learned vs classical
photo_id | leads | learned_pcc_median | classical_pcc_median | diff_median | learned_better_leads | set
C_crumple_20261015_124711 | 12 | 0.8283 | 0.3828 | 0.373 | 12 | C (robustness)
C_crumple_20261015_124807 | 12 | 0.1264 | 0.0824 | 0.01205 | 7 | C (robustness)
C_crumple_20261015_124910 | 12 | -0.04845 | 0.0264 | -0.0239 | 6 | C (robustness)
C_crumple_20261015_125142 | 4 | 0.0783 | 0.03985 | -0.08085 | 1 | C (robustness)
C_crumple_20261015_125332 | 12 | 0.9318 | 0.4392 | 0.44755 | 11 | C (robustness)
C_crumple_20261015_125428 | 12 | 0.01365 | 0.02545 | -0.0016 | 6 | C (robustness)
C_crumple_20261015_125638 | 12 | 0.9281 | 0.88605 | 0.00965 | 8 | C (robustness)
C_shadow_20261014_231732 | 12 | 0.9742 | 0.54875 | 0.41915 | 10 | C (robustness)
C_shadow_20261014_232541 | 11 | 0.9426 | 0.5498 | 0.2389 | 8 | C (robustness)
C_shadow_20261014_235356 | 12 | 0.93575 | 0.74595 | 0.17645 | 9 | C (robustness)
C_tilt_20261014_124026 | 11 | 0.9483 | 0.5919 | 0.3182 | 11 | C (robustness)
C_tilt_20261014_124720 | 8 | 0.97385 | 0.60085 | 0.3851 | 8 | C (robustness)
C_tilt_20261014_124927 | 12 | 0.96065 | 0.58555 | 0.3772 | 12 | C (robustness)
C_tilt_20261014_125027 | 12 | 0.9679 | 0.43055 | 0.53855 | 11 | C (robustness)

## Timing & memory
set | ecg_id | photo_id | tracer | ms_corners | ms_rectify | ms_grid | ms_locate | ms_calibrate+segment | ms_trace | ms_cut | ms_quality | ms_score* | ms_classify | native_mb_corners | native_mb_rectify | native_mb_grid | native_mb_locate | native_mb_calibrate+segment | native_mb_trace | native_mb_cut | native_mb_quality | native_mb_score* | native_mb_classify
C (robustness) | 9 | C_tilt_20261014_122954 | learned | 0 | 73 | 875 | 409 | 24 | 20345 | 29 | 9 | 19 | 101 | 61.8 | 100.9 | 116.6 | 106.1 | 79 | 501 | 439.4 | 439 | 439 | 447.8
C (robustness) | 9 | C_tilt_20261014_124026 | classical | 0 | 134 | 1104 | 114 | 2 | 182 | 6 | 0 | 9 | 58 | 72.2 | 114.3 | 119.9 | 109.7 | 80.7 | 83 | 81.2 | 81.2 | 81.3 | 89.9
C (robustness) | 9 | C_tilt_20261014_124026 | learned | 0 | 109 | 1071 | 116 | 2 | 19433 | 8 | 1 | 8 | 57 | 69.2 | 111.2 | 121.1 | 107 | 78.1 | 500 | 436.6 | 436.6 | 436.7 | 445.9
C (robustness) | 40 | C_crumple_20261015_124711 | classical | 0 | 234 | 1811 | 276 | 4 | 347 | 8 | 1 | 7 | 53 | 67.8 | 120 | 134.6 | 114 | 86 | 90.5 | 87.7 | 87.7 | 87.7 | 96.1
C (robustness) | 40 | C_crumple_20261015_124711 | learned | 0 | 194 | 1825 | 278 | 13 | 20295 | 47 | 10 | 65 | 90 | 57.6 | 109.6 | 129.6 | 108.9 | 80.3 | 502.4 | 439.6 | 439.6 | 440.9 | 448.9
C (robustness) | 40 | C_crumple_20261015_124807 | classical | 0 | 259 | 1974 | 245 | 4 | 374 | 8 | 1 | 6 | 64 | 73.2 | 125.3 | 139.9 | 119.1 | 91 | 96.2 | 91.3 | 91.3 | 91.3 | 99.4
C (robustness) | 40 | C_crumple_20261015_124807 | learned | 0 | 155 | 1867 | 253 | 4 | 20003 | 11 | 1 | 9 | 63 | 66.7 | 118.7 | 134.9 | 114.4 | 86.2 | 508 | 445.7 | 445.4 | 445.4 | 453.6
C (robustness) | 40 | C_crumple_20261015_124910 | classical | 0 | 202 | 2057 | 309 | 4 | 427 | 10 | 1 | 7 | 61 | 74.4 | 126.5 | 145.6 | 125.3 | 96.9 | 104.1 | 97.7 | 97.7 | 97.7 | 106.6
C (robustness) | 40 | C_crumple_20261015_124910 | learned | 0 | 189 | 2062 | 302 | 4 | 19553 | 11 | 1 | 6 | 59 | 67.5 | 119.6 | 139.2 | 118.9 | 90.5 | 512.6 | 450.4 | 450.1 | 450.1 | 458.8
C (robustness) | 40 | C_crumple_20261015_125142 | classical | 0 | 241 | 2350 | 599 | 4 | 444 | 1 | 0 | 2 | 50 | 72.5 | 128.7 | 147.7 | 123 | 94.6 | 100.1 | 95.3 | 95.3 | 95.3 | 103.8
C (robustness) | 40 | C_crumple_20261015_125142 | learned | 0 | 210 | 2222 | 592 | 4 | 20110 | 16 | 1 | 9 | 69 | 65.6 | 121.6 | 141.6 | 117 | 88.8 | 510.6 | 448 | 447.9 | 447.9 | 457
C (robustness) | 40 | C_shadow_20261014_234838 | learned | 0 | 188 | 1724 | 208 | 5 | 19734 | 7 | 0 | 5 | 57 | 104.9 | 156.9 | 176.4 | 115.9 | 87.5 | 509.3 | 446.9 | 446.9 | 447.1 | 456.1
C (robustness) | 40 | C_shadow_20261014_235356 | classical | 0 | 217 | 1451 | 269 | 4 | 337 | 7 | 1 | 6 | 55 | 69.1 | 115.3 | 124 | 109.5 | 81.1 | 85.4 | 82.5 | 82.5 | 82.5 | 91.4
C (robustness) | 40 | C_shadow_20261014_235356 | learned | 0 | 157 | 1431 | 255 | 3 | 19940 | 35 | 10 | 28 | 94 | 58.9 | 105.1 | 118.7 | 104.6 | 76.1 | 498 | 435.5 | 435.7 | 436.9 | 444.2
C (robustness) | 40 | C_tilt_20261014_124846 | learned | 0 | 69 | 866 | 124 | 2 | 19428 | 5 | 1 | 4 | 56 | 76.5 | 116.5 | 124.4 | 115.9 | 87.6 | 509.6 | 447.1 | 447.1 | 447.2 | 455.9
C (robustness) | 40 | C_tilt_20261014_124927 | classical | 0 | 105 | 910 | 156 | 2 | 188 | 7 | 0 | 7 | 59 | 71 | 111.1 | 118.6 | 110.3 | 81.8 | 84.8 | 83.1 | 83.1 | 83.3 | 91.7
C (robustness) | 40 | C_tilt_20261014_124927 | learned | 0 | 73 | 944 | 152 | 2 | 19422 | 8 | 0 | 7 | 52 | 66.9 | 106.9 | 115.2 | 106.7 | 78.3 | 500.2 | 436.6 | 436.4 | 436.4 | 445.2
C (robustness) | 1737 | C_tilt_20261014_124720 | classical | 0 | 110 | 948 | 120 | 2 | 184 | 4 | 0 | 4 | 58 | 70.7 | 110.8 | 115.2 | 107.2 | 78.8 | 80.7 | 79.3 | 79.3 | 79.3 | 87.7
C (robustness) | 1737 | C_tilt_20261014_124720 | learned | 0 | 88 | 974 | 122 | 2 | 19434 | 8 | 0 | 7 | 57 | 63.3 | 103.3 | 109.7 | 101.4 | 73.3 | 495 | 433 | 433.1 | 433.1 | 442.2
C (robustness) | 4016 | C_crumple_20261015_125332 | classical | 0 | 202 | 2541 | 335 | 4 | 369 | 11 | 0 | 6 | 54 | 73.3 | 129.4 | 151.8 | 123.6 | 95.3 | 100.7 | 95.9 | 95.9 | 95.9 | 104
C (robustness) | 4016 | C_crumple_20261015_125332 | learned | 0 | 192 | 2562 | 340 | 4 | 21244 | 15 | 0 | 15 | 79 | 67.9 | 124 | 144.1 | 119.3 | 91.3 | 513.3 | 450.4 | 450.1 | 450.4 | 459.1
C (robustness) | 4016 | C_crumple_20261015_125428 | classical | 0 | 229 | 2586 | 354 | 4 | 455 | 9 | 0 | 8 | 63 | 73.6 | 129.8 | 148.6 | 124.1 | 95.9 | 102.4 | 96.8 | 96.8 | 96.7 | 106.1
C (robustness) | 4016 | C_crumple_20261015_125428 | learned | 0 | 207 | 2591 | 349 | 4 | 19790 | 10 | 0 | 7 | 54 | 66.7 | 122.6 | 142.8 | 118.1 | 89.9 | 512 | 449.5 | 449.5 | 449.5 | 460.4
C (robustness) | 4016 | C_crumple_20261015_125638 | classical | 0 | 222 | 2723 | 384 | 4 | 393 | 11 | 1 | 7 | 57 | 71.8 | 127.8 | 146.7 | 121.8 | 93.6 | 99.7 | 94.7 | 94.7 | 94.7 | 103.6
C (robustness) | 4016 | C_crumple_20261015_125638 | learned | 0 | 207 | 2661 | 385 | 4 | 19753 | 10 | 0 | 9 | 55 | 102.8 | 158.8 | 179.3 | 117.8 | 89.8 | 511.8 | 449.3 | 449.4 | 449.4 | 459.8
C (robustness) | 4016 | C_shadow_20261014_231732 | classical | 0 | 142 | 1334 | 232 | 3 | 225 | 7 | 0 | 5 | 57 | 74.1 | 120.2 | 133.9 | 115.5 | 87.1 | 91 | 87.9 | 88 | 88 | 96.3
C (robustness) | 4016 | C_shadow_20261014_231732 | learned | 0 | 142 | 1377 | 237 | 3 | 19474 | 9 | 0 | 7 | 70 | 71 | 117.1 | 126.6 | 112.3 | 83.9 | 505.5 | 442.3 | 442.3 | 442.3 | 451
C (robustness) | 4016 | C_shadow_20261014_232541 | classical | 0 | 159 | 1358 | 222 | 3 | 248 | 8 | 0 | 5 | 56 | 75.3 | 121.3 | 130.5 | 114.9 | 86.6 | 91 | 88.3 | 88.3 | 88.3 | 96.8
C (robustness) | 4016 | C_shadow_20261014_232541 | learned | 0 | 139 | 1380 | 232 | 3 | 19653 | 7 | 0 | 6 | 52 | 106.1 | 152.1 | 158.5 | 111.9 | 83.4 | 505.3 | 443.7 | 443.8 | 443.8 | 451.8
C (robustness) | 4016 | C_tilt_20261014_125027 | classical | 0 | 112 | 966 | 160 | 2 | 191 | 7 | 0 | 7 | 56 | 73.2 | 113.2 | 119.8 | 111.7 | 83.3 | 86 | 83.5 | 83.5 | 83.5 | 91.6
C (robustness) | 4016 | C_tilt_20261014_125027 | learned | 0 | 94 | 983 | 160 | 2 | 19615 | 8 | 5 | 9 | 87 | 68.7 | 108.7 | 116.8 | 108.1 | 79.7 | 501.6 | 438.6 | 438.7 | 438.8 | 447.5

## Per lead (all)
set_name | ecg_id | ptbxl_label | photo_id | tracer_kind | export | lead_label | row | col | coverage | pcc_asis | pcc_aligned | snr_db_asis | shifted_snr_db_asis | rmse_mv_asis | gain_ratio | r_ratio | r_ratio_aligned | align_scale_pct | align_lag_ms | align_overlap_s | shift_ms | shift_status | quality | warnings
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | I | 0 | 0 | 0.1696 | 0.1883 | 0.2807 | -0.321 | -0.176 | 0.0851 | 0.1027 | 0 | 0 | 0 | 80 | 2.42 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | aVR | 0 | 1 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | V1 | 0 | 2 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | V4 | 0 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | II | 1 | 0 | 0.908 | -0.213 | 0.066 | -15.873 | -15.577 | 0.657 | -1.2731 | 1.5724 | 1.1742 | 10 | -100 | 2.172 | 184 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | aVL | 1 | 1 | 0.9984 | 0.1511 | 0.2385 | -20.017 | -19.774 | 0.5695 | 1.523 | 0.2802 | 0.1535 | 10 | -6 | 2.266 | 200 | FAILED | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | V2 | 1 | 2 | 0.6704 | -0.0784 | -0.0195 | -6.957 | -6.786 | 0.7167 | -0.1408 | 0.3243 | 0.397 | 10 | -40 | 2.232 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | V5 | 1 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | III | 2 | 0 | 0.9928 | 0.0295 | 0.2335 | -20.376 | -19.732 | 0.6831 | 0.3094 | 1.1097 | 1.5429 | -9.5 | -100 | 2.5 | -76 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | aVF | 2 | 1 | 0.9984 | -0.0859 | 0.1323 | -12.442 | -12.178 | 0.3648 | -0.3533 | 1.1296 | 1.2085 | 9.75 | 98 | 2.278 | -194 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | V3 | 2 | 2 | 0.9984 | -0.0375 | 0.0493 | -6.948 | -6.85 | 0.5721 | -0.0747 | 0.1015 | 0.3513 | 4.75 | -100 | 2.286 | 140 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | V6 | 2 | 3 | 0.3616 | -0.1811 | -0.0053 | -8.744 | -8.27 | 0.3356 | -0.4066 | 0 | 0.2523 | 10 | -100 | 2.172 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_122954 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_122954.csv | II rhythm | 3 | 0 | 0.8354 | -0.3106 |  | -16.358 | -16.328 | 0.7823 | -2.0123 | 0.118 |  | -10 | 7064 | 2.936 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | I | 0 | 0 | 0.976 | 0.8325 | 0.8467 | 5.025 | 5.399 | 0.046 | 0.7945 | 0.9243 | 0.9243 | 0.25 | -2 | 2.492 | 6 | ok | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | aVR | 0 | 1 | 0.9624 | 0.8292 | 0.8395 | 4.99 | 5.2 | 0.0541 | 0.7228 | 0.7818 | 0.7795 | -0.5 | -10 | 2.5 | 4 | ok | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | V1 | 0 | 2 | 0.956 | 0.5898 | 0.6009 | 1.359 | 1.387 | 0.1352 | 0.19 | 0.1467 | 0.1454 | 0.5 | 2 | 2.488 | 4 | ok | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | V4 | 0 | 3 | 0.9696 | 0.5571 | 0.642 | 1.663 | 1.692 | 0.1892 | 0.3796 | -0.0537 | -0.0548 | 1.25 | 14 | 2.47 | 2 | ok | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | II | 1 | 0 | 0.504 | 0.4913 | 0.5093 | -0.765 | -0.729 | 0.1154 | 0.4257 | 0.097 | 0.1 | 5.25 | 84 | 2.376 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | aVL | 1 | 1 | 0.9584 | 0.782 | 0.8049 | 4.012 | 4.012 | 0.0358 | 0.5931 | 0.1759 | 0.1756 | -0.75 | -14 | 2.5 | 0 | ok | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | V2 | 1 | 2 | 0.8992 | 0.5919 | 0.5964 | 1.645 | 1.666 | 0.2662 | 0.2375 | 0.1898 | 0.189 | 0.25 | 6 | 2.494 | -2 | ok | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | V5 | 1 | 3 | 0.8704 | 0.573 | 0.5769 | 1.663 | 1.663 | 0.1411 | 0.3436 | 0.039 | 0.0383 | 1.25 | 14 | 2.47 | 0 | ok | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | III | 2 | 0 | 0 | [12 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | aVF | 2 | 1 | 0.728 | 0.6348 | 0.6585 | 2.486 | 2.699 | 0.0654 | 0.3949 | 0.222 | 0.2158 | 0.25 | 10 | 2.49 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | V3 | 2 | 2 | 0.9176 | 0.1893 | 0.2162 | -5.528 | -5.433 | 0.4858 | 0.3472 | -0.0261 | -0.0227 | 1.25 | 28 | 2.47 | -10 | UNRELIABLE | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | V6 | 2 | 3 | 0.96 | 0.7152 | 0.7665 | 3.226 | 3.573 | 0.0846 | 0.5763 | 0.4893 | 0.4886 | 0.75 | 12 | 2.482 | -6 | ok | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | classical | eacg_sheet_ptbxl9_kit_500hz_300dpi_classical_20261014_124029.csv | II rhythm | 3 | 0 | 0.7834 | 0.6937 | 0.6937 | 2.893 | 2.893 | 0.0853 | 0.5965 | 0.1036 | 0.1036 | 0 | 0 | 10 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | I | 0 | 0 | 0.996 | 0.9234 | 0.9762 | 8.193 | 13.203 | 0.0319 | 0.9415 | 0.9366 | 0.9379 | 0 | -6 | 2.494 | 6 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | aVR | 0 | 1 | 0.9992 | 0.9758 | 0.9856 | 13.28 | 14.73 | 0.0208 | 0.9469 | 0.9092 | 0.9027 | -0.25 | -6 | 2.5 | 2 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | V1 | 0 | 2 | 0.996 | 0.908 | 0.9969 | 7.55 | 20.314 | 0.0663 | 0.8581 | 0.9493 | 0.9489 | 0 | 6 | 2.494 | -6 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | V4 | 0 | 3 | 0.9976 | 0.8674 | 0.9614 | 6.046 | 6.046 | 0.1142 | 0.8427 | 0.9041 | 0.9006 | 0.75 | 8 | 2.482 | 0 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | II | 1 | 0 | 0.9928 | 0.8722 | 0.9879 | 6.241 | 16.37 | 0.0515 | 0.8492 | 0.9399 | 0.9403 | 0 | -10 | 2.49 | 10 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | aVL | 1 | 1 | 0.996 | 0.9333 | 0.9564 | 8.536 | 9.817 | 0.0213 | 0.8557 | 0.9158 | 0.9015 | -0.25 | -6 | 2.5 | 2 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | V2 | 1 | 2 | 0.9976 | 0.9571 | 0.9841 | 10.702 | 14.491 | 0.0938 | 0.8805 | 0.8818 | 0.8818 | 0 | 2 | 2.498 | -2 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | V5 | 1 | 3 | 0.9976 | 0.9483 | 0.9853 | 10.037 | 10.284 | 0.0538 | 0.9416 | 0.957 | 0.9514 | 0.5 | 4 | 2.488 | 2 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | III | 2 | 0 | 0.9784 | 0.8099 | 0.9766 | 4.42 | 10.315 | 0.0393 | 0.7774 | 0.9214 | 0.9122 | -0.75 | -20 | 2.498 | 10 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | aVF | 2 | 1 | 0.996 | 0.9847 | 0.9847 | 15.379 | 15.379 | 0.0148 | 0.9906 | 0.9671 | 0.9671 | 0 | 0 | 2.5 | 0 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | V3 | 2 | 2 | 0.9976 | 0.9679 | 0.9892 | 11.999 | 16.032 | 0.0646 | 0.8941 | 0.8804 | 0.8804 | 0 | 2 | 2.498 | -2 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | V6 | 2 | 3 | 0.9928 | 0.9593 | 0.9937 | 11.049 | 18.755 | 0.0344 | 0.946 | 0.9524 | 0.9524 | 0 | 4 | 2.496 | -4 | ok | ok
C (robustness) | 9 | NORM | C_tilt_20261014_124026 | learned | eacg_sheet_ptbxl9_kit_500hz_300dpi_model_20261014_124026.csv | II rhythm | 3 | 0 | 0.9964 | 0.9057 | 0.908 | 7.353 | 7.461 | 0.051 | 0.9526 | 0.9397 | 0.9397 | 0 | 2 | 9.998 | -2 | ok | ok
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | I | 0 | 0 | 0.9752 | 0.1889 | 0.3488 | 0.024 | 0.037 | 0.1635 | 0.0569 | -0.0114 | 0.0625 | 8.25 | -100 | 2.21 | -166 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | aVR | 0 | 1 | 0.9608 | 0.5256 | 0.5256 | 1.58 | 1.58 | 0.1112 | 0.248 | 0.1047 | 0.1047 | 0 | 0 | 2.5 | 0 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | V1 | 0 | 2 | 0.9616 | 0.494 | 0.5403 | 1.078 | 1.123 | 0.2482 | 0.1837 | 0.0336 | 0.0288 | -4 | -20 | 2.5 | -14 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | V4 | 0 | 3 | 0.7488 | 0.1115 | 0.2676 | -2.958 | -2.922 | 0.2518 | 0.1288 | -0.0537 | 0.2366 | 10 | -100 | 2.172 |  | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | II | 1 | 0 | 0.9608 | 0.6892 | 0.8098 | 2.896 | 3.004 | 0.0768 | 0.5318 | 0.0985 | 0.11 | -4.25 | -68 | 2.5 | 8 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | aVL | 1 | 1 | 0.9624 | 0.1392 | 0.2948 | 0.013 | 0.018 | 0.1553 | 0.0254 | 0.0155 | 0.0435 | -5.75 | -96 | 2.5 | -196 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | V2 | 1 | 2 | 0.9368 | 0.355 | 0.4232 | 0.399 | 0.402 | 0.378 | 0.2035 | 0.1368 | 0.1159 | 10 | 90 | 2.272 | 8 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | V5 | 1 | 3 | 0.972 | 0.4106 | 0.4277 | 0.948 | 1.007 | 0.1452 | 0.1688 | 0.0194 | 0.0201 | 0.75 | -2 | 2.48 | 12 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | III | 2 | 0 | 0.9784 | 0.2933 | 0.3416 | 0.369 | 0.507 | 0.111 | 0.1099 | 0.0347 | 0.0406 | 1.5 | 0 | 2.464 | -200 | FAILED | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | aVF | 2 | 1 | 0.964 | 0.765 | 0.8286 | 3.843 | 4.055 | 0.0527 | 0.6669 | 0.3468 | 0.3408 | 1.25 | 10 | 2.47 | 6 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | V3 | 2 | 2 | 0.9592 | 0.2221 | 0.3187 | -1.617 | -1.315 | 0.2899 | 0.1912 | 0.0016 | -0.0367 | 10 | 12 | 2.272 | 28 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | V6 | 2 | 3 | 0.976 | 0.5165 | 0.5727 | 1.526 | 1.655 | 0.1138 | 0.2468 | 0.1007 | 0.1199 | -2.75 | -12 | 2.5 | -18 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124715.csv | II rhythm | 3 | 0 | 0.9986 | 0.476 | 0.4964 | -1.529 | -1.445 | 0.1253 | 0.628 | 0.6044 | 0.603 | 0.25 | 4 | 9.976 | 10 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | I | 0 | 0 | 0.9864 | 0.5733 | 0.9674 | 1.249 | 11.808 | 0.142 | 0.5216 | 0.8767 | 0.8979 | 0.25 | -18 | 2.476 | 22 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | aVR | 0 | 1 | 0.9968 | 0.8872 | 0.8958 | 6.965 | 6.967 | 0.0598 | 0.7435 | 0.8959 | 0.8949 | 0.5 | 6 | 2.488 | 2 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | V1 | 0 | 2 | 0.9984 | 0.9409 | 0.9872 | 9.36 | 14.379 | 0.0956 | 0.8998 | 0.9512 | 0.9456 | -0.25 | 2 | 2.498 | -6 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | V4 | 0 | 3 | 0.9984 | 0.963 | 0.9784 | 11.676 | 13.229 | 0.0467 | 0.8976 | 0.888 | 0.8831 | 0.25 | 6 | 2.494 | -4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | II | 1 | 0 | 0.976 | 0.7186 | 0.9804 | 2.94 | 3.975 | 0.0765 | 0.6669 | 0.5796 | 0.792 | -3.75 | -72 | 2.5 | 10 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | aVL | 1 | 1 | 0.9984 | 0.7694 | 0.8607 | 3.803 | 5.53 | 0.1004 | 0.7097 | 0.8872 | 0.8851 | -0.5 | 0 | 2.5 | -8 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | V2 | 1 | 2 | 0.9976 | 0.9659 | 0.9916 | 11.958 | 11.958 | 0.0999 | 0.9566 | 0.9189 | 0.917 | -1 | -12 | 2.5 | 0 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | V5 | 1 | 3 | 0.9992 | 0.6928 | 0.969 | 2.637 | 6.896 | 0.1196 | 0.6627 | 0.9059 | 0.8958 | 1.75 | 8 | 2.458 | 14 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | III | 2 | 0 | 0.9968 | 0.3585 | 0.9166 | -1.242 | 5.291 | 0.1337 | 0.3854 | 0.7124 | 0.9272 | 1.5 | 50 | 2.45 | -30 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | aVF | 2 | 1 | 0.9992 | 0.9133 | 0.9807 | 7.797 | 14.269 | 0.0334 | 0.9008 | 0.9377 | 0.938 | 0 | 6 | 2.494 | -6 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | V3 | 2 | 2 | 0.9976 | 0.8929 | 0.9826 | 7.316 | 7.419 | 0.1037 | 0.8565 | 0.884 | 0.8755 | -1 | -14 | 2.5 | 2 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | V6 | 2 | 3 | 0.9872 | 0.745 | 0.9901 | 3.386 | 11.014 | 0.0918 | 0.7099 | 0.9254 | 0.9258 | -1 | 2 | 2.498 | -14 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124711 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124711.csv | II rhythm | 3 | 0 | 0.9978 | 0.7321 | 0.7974 | 2.698 | 2.698 | 0.077 | 0.7649 | 0.8155 | 0.7501 | 0.5 | 22 | 9.95 | 0 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | I | 0 | 0 | 0.98 | -0.0463 | 0.0998 | -5.445 | -4.642 | 0.3068 | -0.0728 | 0.017 | 0.0601 | -0.25 | -100 | 2.406 | 130 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | aVR | 0 | 1 | 0.9288 | 0.0709 | 0.161 | -6.601 | -6.43 | 0.2852 | 0.1336 | 0.1504 | 0.1527 | -10 | 6 | 2.494 | -8 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | V1 | 0 | 2 | 0.9632 | -0.0226 | 0.2131 | -4.347 | -4.347 | 0.4634 | -0.026 | 0.0086 | 0.2131 | 10 | -94 | 2.178 | 200 | FAILED | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | V4 | 0 | 3 | 0.868 | 0.1967 | 0.2183 | -1.537 | -1.489 | 0.2139 | 0.1894 | 0.0437 | 0.0437 | -2.25 | -2 | 2.5 | -12 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | II | 1 | 0 | 0.8584 | 0.4641 | 0.6162 | 0.937 | 1.234 | 0.0963 | 0.3126 | 0.0642 | 0.2854 | 10 | -62 | 2.21 | 142 | ok | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | aVL | 1 | 1 | 0.9632 | 0.1398 | 0.8029 | -2.247 | -2.008 | 0.2015 | 0.1334 | 0.0125 | 0.4905 | -10 | 28 | 2.472 | -126 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | V2 | 1 | 2 | 0.7408 | -0.0647 | -0.0052 | -4.623 | -4.428 | 0.6739 | -0.0869 | 0.035 | -0.0195 | -10 | 68 | 2.432 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | V5 | 1 | 3 | 0.7976 | 0.1763 | 0.2171 | -1.125 | -1.1 | 0.1844 | 0.1433 | 0.0042 | 0.0033 | -4 | -38 | 2.5 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | III | 2 | 0 | 0.9392 | 0.1747 | 0.279 | -0.145 | -0.121 | 0.1178 | 0.0744 | 0.0164 | 0.0335 | 2.25 | -100 | 2.344 | 114 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | aVF | 2 | 1 | 0.964 | -0.0758 | 0.6438 | -4.845 | -4.138 | 0.1433 | -0.1052 | 0.0438 | 0.2205 | -10 | 2 | 2.498 | -200 | FAILED | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | V3 | 2 | 2 | 0.9632 | -0.1755 | 0.062 | -2.37 | -2.091 | 0.3161 | -0.1405 | -0.0002 | 0.4102 | -10 | 80 | 2.42 | 200 | FAILED | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | V6 | 2 | 3 | 0.8448 | 0.0939 | 0.1863 | -1.755 | -1.747 | 0.166 | 0.0817 | 0.0015 | 0.0027 | -8.5 | -72 | 2.5 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124811.csv | II rhythm | 3 | 0 | 0.8704 | 0.1401 | 0.5943 | -3.028 | -2.261 | 0.149 | 0.165 | 0.0205 | 0.3684 | -10 | -224 | 10 | 44 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | I | 0 | 0 | 0.9864 | 0.003 | 0.3719 | -4.046 | -3.479 | 0.2612 | 0.0039 | 0.013 | 0.3984 | 5.25 | -100 | 2.276 | 146 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | aVR | 0 | 1 | 0.9984 | 0.2218 | 0.5965 | -0.846 | -0.594 | 0.147 | 0.1801 | 0.2213 | 0.7822 | -10 | 12 | 2.488 | -8 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | V1 | 0 | 2 | 0.9984 | 0.001 | 0.033 | -2.252 | -2.248 | 0.3641 | 0.0008 | 0.0164 | 0.0382 | 10 | -90 | 2.182 | -72 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | V4 | 0 | 3 | 0.6904 | 0.1193 | 0.2152 | -1.435 | -1.423 | 0.2114 | 0.0969 | 0.0067 | -0.0065 | -7.25 | -40 | 2.5 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | II | 1 | 0 | 0.9808 | 0.2034 | 0.5615 | -2.063 | -0.269 | 0.136 | 0.2117 | 0.0298 | 0.3596 | -1.75 | -100 | 2.444 | 164 | ok | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | aVL | 1 | 1 | 0.9984 | 0.2482 | 0.7313 | -1.024 | -0.871 | 0.175 | 0.2115 | 0.0146 | 0.8031 | -10 | 22 | 2.478 | -124 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | V2 | 1 | 2 | 0.9984 | -0.1265 | 0.0535 | -2.47 | -2.427 | 0.5259 | -0.1034 | 0.0212 | 0.078 | 10 | -100 | 2.172 | -76 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | V5 | 1 | 3 | 0.6888 | 0.1386 | 0.2144 | -1.473 | -1.372 | 0.1919 | 0.115 | 0.0048 | 0.0056 | -4.75 | -46 | 2.5 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | III | 2 | 0 | 0.9984 | 0.0605 | 0.9013 | -5.612 | -2.137 | 0.2211 | 0.1034 | 0.1513 | 0.7642 | -3.5 | -100 | 2.49 | 68 | ok | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | aVF | 2 | 1 | 0.9984 | 0.241 | 0.9403 | -1.559 | -0.87 | 0.0981 | 0.2337 | 0.0711 | 0.8408 | -10 | 0 | 2.5 | -102 | ok | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | V3 | 2 | 2 | 0.9984 | -0.175 | 0.0624 | -2.501 | -2.203 | 0.3209 | -0.1454 | 0.0007 | 0.4163 | -10 | 80 | 2.42 | -92 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | V6 | 2 | 3 | 0.6936 | 0.1335 | 0.2153 | -1.462 | -1.462 | 0.1605 | 0.1118 | 0.0179 | 0.0038 | -8.5 | -74 | 2.5 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124807 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124807.csv | II rhythm | 3 | 0 | 0.9106 | 0.129 | 0.5899 | -2.404 | -1.874 | 0.1386 | 0.1325 | 0.0313 | 0.7498 | -10 | -222 | 10 | 30 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | I | 0 | 0 | 0.9808 | -0.1185 | -0.0638 | -6.424 | -5.92 | 0.3435 | -0.2062 | 0.022 | 0.0201 | -3.75 | -100 | 2.498 | 200 | FAILED | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | aVR | 0 | 1 | 0.952 | -0.0168 | 0.194 | -0.652 | -0.539 | 0.1438 | -0.0076 | -0.0093 | -0.0078 | -5.75 | 2 | 2.498 | -116 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | V1 | 0 | 2 | 0.964 | 0.0667 | 0.2787 | -0.318 | -0.317 | 0.2915 | 0.023 | 0.0288 | 0.0387 | 10 | -100 | 2.172 | 200 | FAILED | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | V4 | 0 | 3 | 0.6776 | -0.0297 | 0.2951 | -0.43 | -0.228 | 0.1883 | -0.0109 | 0.0203 | 0.1291 | 10 | -30 | 2.242 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | II | 1 | 0 | 0.9808 | 0.1099 | 0.3564 | -0.559 | -0.331 | 0.1144 | 0.0581 | -0.0049 | 0.0652 | 10 | -100 | 2.172 | 200 | FAILED | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | aVL | 1 | 1 | 0.956 | 0.1299 | 0.1949 | -1.645 | -1.408 | 0.188 | 0.0957 | 0.0033 | 0.0472 | 0.25 | -58 | 2.436 | 60 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | V2 | 1 | 2 | 0.9048 | 0.0202 | 0.1509 | -0.086 | -0.072 | 0.3997 | 0.0044 | 0.0562 | 0.0391 | 6 | -100 | 2.258 | 178 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | V5 | 1 | 3 | 0.7144 | -0.0701 | 0.3397 | -0.908 | -0.768 | 0.1799 | -0.0209 | -0.0321 | 0.0239 | -10 | 26 | 2.474 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | III | 2 | 0 | 0.956 | 0.1194 | 0.2317 | -0.305 | 0.116 | 0.12 | 0.0506 | 0.0414 | 0.0414 | -0.5 | -36 | 2.476 | 34 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | aVF | 2 | 1 | 0.956 | -0.198 | 0.1668 | -4.737 | -4.217 | 0.1415 | -0.2451 | 0.0258 | -0.0345 | -10 | 100 | 2.4 | 162 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | V3 | 2 | 2 | 0.9648 | 0.047 | 0.2817 | -2.393 | -1.525 | 0.317 | 0.0461 | 0.2907 | 0.3349 | -2.5 | -100 | 2.464 | 40 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | V6 | 2 | 3 | 0.6832 | 0.0326 | 0.2621 | -5.428 | -5.301 | 0.2533 | 0.0448 | -0.1398 | 0.2309 | -1 | 100 | 2.4 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_124915.csv | II rhythm | 3 | 0 | 0.8852 | -0.0303 | 0.266 | -4.696 | -3.946 | 0.1805 | -0.043 | -0.0002 | 0.0669 | -3.25 | 794 | 9.206 | 158 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | I | 0 | 0 | 0.8728 | -0.0878 | 0.1085 | -7.063 | -6.623 | 0.3697 | -0.1687 | 0.0221 | 0.3705 | 1.25 | -98 | 2.372 | 120 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | aVR | 0 | 1 | 0.9984 | 0.0132 | 0.2389 | -3.039 | -2.971 | 0.1893 | 0.0144 | 0.1194 | 0.1705 | 9.5 | 26 | 2.284 | -162 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | V1 | 0 | 2 | 0.9984 | -0.1302 | 0.3249 | -2.318 | -1.678 | 0.3669 | -0.0933 | 0.0274 | 0.0461 | 1.75 | -100 | 2.358 | 126 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | V4 | 0 | 3 | 0.5448 | -0.0056 | 0.6841 | -1.58 | -1.459 | 0.2149 | -0.0033 | -0.0192 | 0.3706 | -10 | 48 | 2.452 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | II | 1 | 0 | 0.8744 | -0.2347 | -0.018 | -9.321 | -8.719 | 0.3136 | -0.5801 | -0.028 | 0.3799 | 0.25 | -100 | 2.394 | 106 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | aVL | 1 | 1 | 0.9984 | 0.1893 | 0.322 | -2.06 | -1.683 | 0.1972 | 0.1909 | 0.0367 | -0.016 | -7.75 | 100 | 2.4 | -172 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | V2 | 1 | 2 | 0.9984 | -0.0894 | 0.2721 | -1.892 | -1.338 | 0.4921 | -0.0642 | 0.0634 | 0.0688 | 1.5 | -100 | 2.364 | 126 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | V5 | 1 | 3 | 0.5464 | -0.05 | 0.6845 | -1.309 | -1.234 | 0.1883 | -0.0283 | -0.022 | 0.3938 | -10 | 28 | 2.472 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | III | 2 | 0 | 0.9568 | -0.0469 | 0.2426 | -7.17 | -6.952 | 0.2645 | -0.0932 | 0.0491 | 0.3705 | 8.25 | -100 | 2.21 | 176 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | aVF | 2 | 1 | 0.9984 | 0.1885 | 0.5595 | -3.658 | -1.386 | 0.125 | 0.2585 | -0.0325 | 0.3733 | -10 | 54 | 2.446 | -156 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | V3 | 2 | 2 | 0.9984 | -0.1108 | 0.2638 | -1.818 | -0.65 | 0.2967 | -0.0804 | 0.0217 | 0.3188 | 10 | -12 | 2.26 | 148 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | V6 | 2 | 3 | 0.5432 | -0.0353 | 0.6943 | -1.169 | -1.005 | 0.1551 | -0.0198 | -0.0109 | 0.4155 | -10 | 36 | 2.464 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_124910 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_124910.csv | II rhythm | 3 | 0 | 0.8584 | 0.0712 |  | -2.021 | -1.346 | 0.1326 | 0.0631 | 0 |  | 8.75 | 6410 | 3.59 | 44 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | I | 0 | 0 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | aVR | 0 | 1 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | V1 | 0 | 2 | 0.0896 | 0.1637 | 0.1851 | -8.36 | -8.249 | 0.7356 | 0.3222 | 0 | 0 | -3.75 | -98 | 2.5 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | V4 | 0 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | II | 1 | 0 | 0.8616 | -0.1183 | 0.116 | -3.113 | -2.679 | 0.1535 | -0.1123 | -0.0134 | 0.3061 | 1.25 | -100 | 2.37 | 122 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | aVL | 1 | 1 | 0.9552 | 0.196 | 0.2968 | -4.377 | -3.381 | 0.2574 | 0.292 | 0.0376 | 0.0526 | -10 | 94 | 2.406 | -190 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | V2 | 1 | 2 | 0.3256 | -0.084 | 0.0916 | -4.668 | -4.632 | 0.6773 | -0.0881 | 0 | 0.3594 | 10 | -100 | 2.172 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | V5 | 1 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | III | 2 | 0 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | aVF | 2 | 1 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | V3 | 2 | 2 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | V6 | 2 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261015_125147.csv | II rhythm | 3 | 0 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | I | 0 | 0 | 0.896 | -0.0934 | 0.205 | -5.004 | -4.587 | 0.2916 | -0.1293 | 0.0155 | 0.3912 | 2.5 | -100 | 2.34 | 142 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | aVR | 0 | 1 | 0.9992 | 0.0156 | 0.3059 | -3.354 | -3.013 | 0.1963 | 0.0182 | 0.1662 | 0.366 | -8.25 | 100 | 2.4 | -176 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | V1 | 0 | 2 | 1 | 0.0287 | 0.3398 | -1.73 | -1.6 | 0.3429 | 0.0207 | 0.0269 | 0.0293 | 0.25 | -100 | 2.394 | 106 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | V4 | 0 | 3 | 0.5048 | 0.0367 | 0.6854 | -1.073 | -0.627 | 0.2027 | 0.0201 | 0.0714 | 0.3923 | -10 | 94 | 2.406 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | II | 1 | 0 | 0.8632 | -0.3068 | -0.222 | -12.746 | -12.472 | 0.4652 | -1.1603 | 0.0043 | 0.3543 | 1.25 | -100 | 2.37 | 126 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | aVL | 1 | 1 | 0.9984 | 0.1693 | 0.3235 | -2.344 | -1.866 | 0.2037 | 0.1776 | 0.0345 | 0.0549 | -10 | 92 | 2.408 | -188 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | V2 | 1 | 2 | 0.9984 | 0.1279 | 0.2342 | -0.914 | -0.871 | 0.4397 | 0.09 | 0.0571 | 0.0663 | -0.25 | -100 | 2.406 | 184 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | V5 | 1 | 3 | 0.5128 | -0.0161 | 0.6746 | -1.093 | -0.462 | 0.1837 | -0.0086 | 0.0339 | 0.3857 | -10 | 90 | 2.41 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | III | 2 | 0 | 0.9264 | -0.0227 | 0.0981 | -9.659 | -9.257 | 0.3523 | -0.066 | 0.0302 | 0.5746 | 0.25 | -96 | 2.398 | 100 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | aVF | 2 | 1 | 0.9984 | 0.057 | 0.4866 | -4.022 | -2.376 | 0.1303 | 0.0755 | -0.0022 | 0.2866 | -10 | 72 | 2.428 | -172 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | V3 | 2 | 2 | 0.9984 | 0.1104 | 0.2776 | -0.846 | -0.678 | 0.2653 | 0.0782 | -0.0094 | 0.1196 | 9.5 | 22 | 2.284 | 106 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | V6 | 2 | 3 | 0.5168 | -0.1061 | 0.0123 | -15.143 | -15.12 | 0.7752 | -0.6011 | 0.2193 | 1.191 | -10 | 100 | 2.4 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_crumple_20261015_125142 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261015_125142.csv | II rhythm | 3 | 0 | 0.8376 | 0.0347 | 0.3095 | -4.233 | -4.056 | 0.1711 | 0.0476 | 0.1405 | 0.0769 | 0 | 696 | 9.304 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | I | 0 | 0 | 0.0432 | 0.1036 | 0.1234 | -4.464 | -4.419 | 0.2741 | 0.1454 | 0 | 0 | -10 | 74 | 2.426 |  | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | aVR | 0 | 1 | 0 | [12 empty] | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | V1 | 0 | 2 | 0 | [12 empty] | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | V4 | 0 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | II | 1 | 0 | 0.988 | 0.141 | 0.29 | -4.843 | -4.821 | 0.1873 | 0.2283 | -0.0337 | 0.099 | -10 | 14 | 2.486 | -132 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | aVL | 1 | 1 | 0.9992 | -0.0316 | 0.2591 | -1.804 | -1.276 | 0.1914 | -0.0213 | -0.0073 | 0.0744 | -7.75 | 76 | 2.424 | -164 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | V2 | 1 | 2 | 0.9976 | -0.0151 | 0.4302 | -1.589 | -1.579 | 0.4752 | -0.0097 | -0.0102 | 0.0755 | -10 | 44 | 2.456 | -200 | FAILED | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | V5 | 1 | 3 | 0.9976 | 0.1717 | 0.4896 | -3.853 | -3.837 | 0.2524 | 0.2506 | 0.3607 | 0.5941 | -8.25 | 100 | 2.4 | 6 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | III | 2 | 0 | 0.9808 | 0.0655 | 0.1418 | -3.483 | -2.952 | 0.173 | 0.0753 | 0.0942 | 0.0661 | 2.75 | -100 | 2.334 | 152 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | aVF | 2 | 1 | 0.9992 | -0.0746 | 0.0349 | -5.384 | -5.263 | 0.1524 | -0.1093 | 0.0032 | 0.0609 | -9.75 | -100 | 2.5 | 200 | FAILED | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | V3 | 2 | 2 | 0.9976 | 0.1023 | 0.3496 | -4.213 | -4.132 | 0.3909 | 0.1535 | 0.8649 | 0.0815 | -10 | 62 | 2.438 | -196 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | V6 | 2 | 3 | 0.9976 | 0.0997 | 0.3743 | -6.272 | -5.713 | 0.2792 | 0.1966 | 0.7365 | 0.7415 | 7.75 | 98 | 2.32 | -44 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_234838 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_234838.csv | II rhythm | 3 | 0 | 0.9424 | 0.0477 | 0.5926 | -4.62 | -4.601 | 0.1789 | 0.0707 | 0.1787 | 0.3124 | 2.25 | -524 | 9.256 | -32 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | I | 0 | 0 | 0.3008 | 0.0616 | 0.3267 | -0.001 | 0.06 | 0.1639 | 0.0123 | 0 | 0.0672 | 10 | -88 | 2.184 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | aVR | 0 | 1 | 0.8744 | 0.1989 | 0.227 | -6.248 | -6.243 | 0.2739 | 0.3791 | 0.3936 | 0.3861 | -3.25 | -56 | 2.5 | -4 | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | V1 | 0 | 2 | 0.9624 | 0.9164 | 0.9738 | 7.909 | 12.069 | 0.113 | 0.8021 | 0.8703 | 0.8703 | 0 | 8 | 2.492 | -8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | V4 | 0 | 3 | 0.9784 | 0.8611 | 0.9799 | 6.13 | 9.657 | 0.0885 | 0.8104 | 0.7534 | 0.7531 | 1.25 | 6 | 2.47 | 8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | II | 1 | 0 | 0.444 | 0.472 | 0.4916 | 0.712 | 0.749 | 0.0988 | 0.2299 | 0.0538 | 0.0538 | 5.5 | 98 | 2.37 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | aVL | 1 | 1 | 0.9624 | 0.7217 | 0.7372 | 3.234 | 3.27 | 0.1072 | 0.4771 | 0.2815 | 0.2818 | 1 | 20 | 2.476 | -2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | V2 | 1 | 2 | 0.9008 | 0.5117 | 0.6107 | -0.172 | -0.069 | 0.4037 | 0.5513 | 0.7623 | 0.6922 | 4 | 30 | 2.404 | -6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | V5 | 1 | 3 | 0.9784 | 0.8031 | 0.989 | 4.604 | 7.38 | 0.0954 | 0.756 | 0.8405 | 0.8374 | 1.75 | 12 | 2.458 | 10 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | III | 2 | 0 | 0.9792 | 0.3653 | 0.4957 | 0.682 | 0.831 | 0.1071 | 0.1558 | 0.1273 | 0.1647 | 5 | 100 | 2.38 | -12 | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | aVF | 2 | 1 | 0.8944 | 0.7702 | 0.8053 | 3.881 | 4.473 | 0.0525 | 0.6844 | 0.8172 | 0.8139 | 0.5 | 14 | 2.486 | -6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | V3 | 2 | 2 | 0.9608 | 0.7999 | 0.8532 | 4.873 | 5.913 | 0.1373 | 0.6312 | 0.3224 | 0.3595 | 0.5 | 14 | 2.486 | -8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | V6 | 2 | 3 | 0.9784 | 0.8994 | 0.9848 | 7.355 | 9.893 | 0.0581 | 0.8673 | 0.8456 | 0.8438 | 1.25 | 8 | 2.47 | 6 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_235400.csv | II rhythm | 3 | 0 | 0.9486 | 0.7455 | 0.8076 | 2.796 | 2.829 | 0.0762 | 0.7705 | 0.5973 | 0.5879 | 0.5 | 24 | 9.95 | 2 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | I | 0 | 0 | 0.9968 | 0.9784 | 0.9973 | 13.864 | 16.256 | 0.0332 | 0.9486 | 0.9422 | 0.932 | 0.5 | 4 | 2.488 | 2 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | aVR | 0 | 1 | 0.9984 | 0.8532 | 0.9918 | 5.781 | 11.743 | 0.0686 | 0.817 | 0.92 | 0.9171 | -0.5 | -14 | 2.498 | 8 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | V1 | 0 | 2 | 0.9992 | 0.9342 | 0.9945 | 8.874 | 19.399 | 0.1011 | 0.9179 | 0.9543 | 0.9543 | 0 | 6 | 2.494 | -6 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | V4 | 0 | 3 | 0.9976 | 0.8225 | 0.996 | 4.851 | 9.931 | 0.1025 | 0.8249 | 0.9273 | 0.9152 | 1.25 | 6 | 2.47 | 10 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | II | 1 | 0 | 0.9984 | 0.9521 | 0.9933 | 10.431 | 16.033 | 0.0323 | 0.9221 | 0.8728 | 0.8697 | 0.25 | -2 | 2.492 | 6 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | aVL | 1 | 1 | 0.9992 | 0.8812 | 0.995 | 6.561 | 13.094 | 0.0731 | 0.8358 | 0.8969 | 0.8823 | -0.5 | -14 | 2.498 | 8 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | V2 | 1 | 2 | 0.9976 | 0.9812 | 0.9968 | 14.563 | 22.073 | 0.074 | 0.9668 | 0.9202 | 0.9198 | 0 | 4 | 2.496 | -4 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | V5 | 1 | 3 | 0.9976 | 0.793 | 0.9779 | 3.941 | 8.24 | 0.1029 | 0.8318 | 0.9406 | 0.9358 | 1.5 | 8 | 2.464 | 10 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | III | 2 | 0 | 0.9968 | 0.9894 | 0.9943 | 16.713 | 19.116 | 0.0169 | 0.9481 | 0.8847 | 0.8842 | 0 | -2 | 2.498 | 2 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | aVF | 2 | 1 | 0.9992 | 0.9636 | 0.9854 | 11.596 | 12.893 | 0.0216 | 0.9327 | 0.9114 | 0.9117 | -0.25 | -4 | 2.5 | 2 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | V3 | 2 | 2 | 0.9864 | 0.9373 | 0.992 | 9.588 | 18.253 | 0.0798 | 0.9101 | 0.9028 | 0.903 | 0 | 6 | 2.494 | -6 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | V6 | 2 | 3 | 0.9976 | 0.8729 | 0.9933 | 6.083 | 9.942 | 0.0673 | 0.8959 | 0.9695 | 0.9685 | 1.25 | 6 | 2.47 | 8 | ok | ok
C (robustness) | 40 | NORM | C_shadow_20261014_235356 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_235356.csv | II rhythm | 3 | 0 | 0.9988 | 0.8204 | 0.8634 | 4.158 | 4.203 | 0.0651 | 0.8949 | 0.8779 | 0.8616 | 0.25 | 8 | 9.976 | 2 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | I | 0 | 0 | 0.9976 | 0.0891 | 0.1192 | -4.283 | -4.256 | 0.2684 | 0.1275 | 0.2033 | 0.1559 | 4.5 | 58 | 2.392 | 200 | FAILED | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | aVR | 0 | 1 | 0.8256 | 0.0789 | 0.0951 | -3.205 | -3.133 | 0.1929 | 0.0898 | 0.1122 | 0.1104 | -1.5 | 10 | 2.49 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | V1 | 0 | 2 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | V4 | 0 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | II | 1 | 0 | 0.9984 | 0.2482 | 0.2988 | -8.726 | -8.509 | 0.2929 | 0.6991 | 0.635 | 0.8788 | -5.25 | -78 | 2.5 | 28 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | aVL | 1 | 1 | 0.8296 | -0.1518 | -0.0411 | -2.765 | -2.666 | 0.2138 | -0.127 | -0.0644 | 0.1033 | 10 | 100 | 2.272 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | V2 | 1 | 2 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | V5 | 1 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | III | 2 | 0 | 0.9992 | -0.0441 | 0.2775 | -5.642 | -4.673 | 0.2218 | -0.0671 | -0.0493 | 0.2514 | -10 | -44 | 2.5 | -124 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | aVF | 2 | 1 | 0.972 | 0.2626 | 0.3989 | -12.387 | -12.213 | 0.3414 | 1.0438 | 0.1344 | 0.1361 | 5.25 | -20 | 2.356 | 94 | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | V3 | 2 | 2 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | V6 | 2 | 3 | 0 | [12 empty] | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124846 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124846.csv | II rhythm | 3 | 0 | 0.4574 | -0.1681 |  | -5.717 | -5.524 | 0.203 | -0.2023 | 0 |  | 2.25 | 8178 | 1.822 |  | UNRELIABLE | unreliable | FEW_TICKS/warn; SCALE_DISAGREE/severe; NO_PULSE/warn; LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6; LEAD_COVERAGE/warn/II_rhythm
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | I | 0 | 0 | 0.98 | 0.5836 | 0.6422 | 1.871 | 2.01 | 0.1322 | 0.391 | 0.2491 | 0.2489 | 2.75 | 36 | 2.434 | 4 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | aVR | 0 | 1 | 0.9624 | 0.5875 | 0.6165 | 1.775 | 2.064 | 0.1087 | 0.4851 | 0.211 | 0.2075 | 0.25 | -4 | 2.49 | 6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | V1 | 0 | 2 | 0.9624 | 0.4424 | 0.4626 | -0.25 | -0.132 | 0.2892 | 0.406 | 0.0592 | 0.0533 | 0.75 | 18 | 2.482 | -10 | UNRELIABLE | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | V4 | 0 | 3 | 0.9792 | 0.525 | 0.5335 | 1.691 | 1.743 | 0.1475 | 0.3155 | 0.0432 | 0.0481 | 0 | 8 | 2.492 | -8 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | II | 1 | 0 | 0.98 | 0.7679 | 0.7956 | 3.92 | 4.054 | 0.0683 | 0.6693 | 0.2957 | 0.2928 | -1.5 | -16 | 2.5 | -4 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | aVL | 1 | 1 | 0.9392 | 0.5057 | 0.5788 | 1.348 | 1.485 | 0.1332 | 0.3062 | 0.2712 | 0.266 | 1.75 | 18 | 2.458 | 6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | V2 | 1 | 2 | 0.9624 | 0.8529 | 0.8643 | 5.669 | 5.8 | 0.2061 | 0.6256 | 0.0714 | 0.0713 | 0.75 | 12 | 2.482 | -6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | V5 | 1 | 3 | 0.9792 | 0.5627 | 0.566 | 1.721 | 1.731 | 0.1329 | 0.2608 | 0.0685 | 0.0685 | 0 | -4 | 2.496 | 6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | III | 2 | 0 | 0.9544 | 0.8352 | 0.8585 | 4.341 | 4.59 | 0.0703 | 0.4946 | 0.333 | 0.3326 | 0.75 | 14 | 2.482 | -6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | aVF | 2 | 1 | 0.9624 | 0.6193 | 0.6239 | -3.953 | -3.631 | 0.1293 | 1.1989 | 0.9319 | 0.9319 | 0.25 | 6 | 2.494 | -2 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | V3 | 2 | 2 | 0.9624 | 0.7661 | 0.8065 | 4.238 | 4.768 | 0.1477 | 0.5829 | 0.2771 | 0.2769 | 0.5 | 14 | 2.486 | -6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | V6 | 2 | 3 | 0.9792 | 0.4895 | 0.5279 | -2.107 | -2.092 | 0.1728 | 0.6681 | 0.6785 | 0.6733 | 2 | 24 | 2.45 | -4 | UNRELIABLE | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | classical | eacg_sheet_ptbxl40_kit_500hz_300dpi_classical_20261014_124930.csv | II rhythm | 3 | 0 | 0.9312 | 0.7643 | 0.841 | 3.725 | 3.725 | 0.0685 | 0.7001 | 0.6475 | 0.6435 | 0.5 | 30 | 9.95 | 0 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | I | 0 | 0 | 0.9936 | 0.945 | 0.9899 | 9.621 | 16.054 | 0.0541 | 0.937 | 0.9208 | 0.9181 | -0.25 | -8 | 2.498 | 6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | aVR | 0 | 1 | 0.9984 | 0.9805 | 0.9922 | 14.259 | 17.705 | 0.0258 | 0.9222 | 0.8989 | 0.8989 | 0 | -2 | 2.498 | 2 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | V1 | 0 | 2 | 0.9984 | 0.9626 | 0.9964 | 11.296 | 20.631 | 0.0765 | 0.939 | 0.9361 | 0.936 | 0 | 6 | 2.494 | -6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | V4 | 0 | 3 | 0.9984 | 0.9778 | 0.9906 | 13.718 | 13.718 | 0.0369 | 0.9557 | 0.9144 | 0.9012 | 0.5 | 6 | 2.488 | 0 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | II | 1 | 0 | 0.996 | 0.9419 | 0.9894 | 9.618 | 16.242 | 0.0354 | 0.8913 | 0.8782 | 0.8746 | -0.25 | -8 | 2.498 | 6 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | aVL | 1 | 1 | 0.9984 | 0.9884 | 0.9937 | 16.193 | 18.441 | 0.0241 | 0.9371 | 0.8995 | 0.8995 | 0 | -2 | 2.498 | 2 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | V2 | 1 | 2 | 0.9984 | 0.9803 | 0.9965 | 14.375 | 21.705 | 0.0756 | 0.966 | 0.9109 | 0.9103 | 0 | 4 | 2.496 | -4 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | V5 | 1 | 3 | 0.9984 | 0.9587 | 0.9883 | 11.115 | 12.129 | 0.0451 | 0.9543 | 0.9462 | 0.9342 | 1 | 10 | 2.476 | 2 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | III | 2 | 0 | 0.9984 | 0.9462 | 0.9851 | 9.788 | 10.207 | 0.0375 | 0.9534 | 0.8883 | 0.8829 | -0.5 | -4 | 2.5 | 2 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | aVF | 2 | 1 | 0.9984 | 0.932 | 0.9585 | 7.802 | 8.578 | 0.0334 | 1.0088 | 0.9433 | 0.9338 | 0.25 | 6 | 2.494 | -2 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | V3 | 2 | 2 | 0.9984 | 0.946 | 0.9899 | 10.207 | 17.032 | 0.0743 | 0.912 | 0.9107 | 0.9107 | 0 | 4 | 2.496 | -4 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | V6 | 2 | 3 | 0.9984 | 0.9741 | 0.9971 | 13.144 | 13.975 | 0.0299 | 0.9618 | 0.9526 | 0.9493 | 0.75 | 8 | 2.482 | 2 | ok | ok
C (robustness) | 40 | NORM | C_tilt_20261014_124927 | learned | eacg_sheet_ptbxl40_kit_500hz_300dpi_model_20261014_124927.csv | II rhythm | 3 | 0 | 0.9994 | 0.8603 | 0.9561 | 5.861 | 6 | 0.0535 | 0.84 | 0.8681 | 0.8544 | 0.25 | 12 | 9.976 | -2 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | I | 0 | 0 | 0.9464 | 0.5148 | 0.5785 | 1.327 | 1.364 | 0.1741 | 0.3857 | 0.3254 | 0.3262 | 2.25 | 24 | 2.444 | 0 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | aVR | 0 | 1 | 0.9448 | 0.7419 | 0.7756 | 3.648 | 4.202 | 0.0677 | 0.6076 | 0.4886 | 0.4883 | 0.25 | -2 | 2.492 | 4 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | V1 | 0 | 2 | 0.048 | -0.2094 | 0.0924 | -0.692 | -0.269 | 0.1731 | -0.0501 | 0 | 0 | -10 | 100 | 2.4 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | V4 | 0 | 3 | 0 | [12 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | II | 1 | 0 | 0.9784 | 0.7008 | 0.7587 | 2.957 | 3.501 | 0.0409 | 0.5175 | 0.0465 | 0.0448 | 0.5 | 6 | 2.488 | 4 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | aVL | 1 | 1 | 0.96 | 0.7717 | 0.7733 | 3.718 | 3.732 | 0.1131 | 0.442 | 0.2609 | 0.2609 | 0 | 2 | 2.498 | -2 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | V2 | 1 | 2 | 0.1304 | 0.3157 | 0.3359 | 0.167 | 0.195 | 0.2219 | 0.1438 | 0 | 0 | 10 | 30 | 2.272 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | V5 | 1 | 3 | 0 | [12 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | III | 2 | 0 | 0.9472 | 0.4138 | 0.4399 | 0.605 | 0.677 | 0.1729 | 0.3186 | 0.178 | 0.1781 | 2 | 26 | 2.45 | 0 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | aVF | 2 | 1 | 0.744 | 0.6869 | 0.7031 | 0.567 | 0.683 | 0.079 | 0.8983 | 0.6706 | 0.6687 | -0.75 | -16 | 2.5 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | V3 | 2 | 2 | 0 | [12 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | V6 | 2 | 3 | 0 | [12 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | classical | eacg_sheet_ptbxl1737_kit_500hz_300dpi_classical_20261014_124723.csv | II rhythm | 3 | 0 | 0.9456 | 0.6483 | 0.6958 | 1.742 | 1.782 | 0.0475 | 0.6063 | 0.5703 | 0.5654 | 0.25 | 18 | 9.976 | -2 | ok | check | LEAD_COVERAGE/warn/V1; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/aVF; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | I | 0 | 0 | 0.996 | 0.9877 | 0.9885 | 16.035 | 16.302 | 0.032 | 0.9312 | 0.9119 | 0.9119 | 0 | -2 | 2.498 | 2 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | aVR | 0 | 1 | 0.9992 | 0.9721 | 0.987 | 12.815 | 16.016 | 0.0236 | 0.9418 | 0.899 | 0.899 | 0 | -2 | 2.498 | 2 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | V1 | 0 | 2 | 0.9872 | 0.9743 | 0.9965 | 13.063 | 19.388 | 0.0355 | 0.9236 | 0.9403 | 0.9397 | -0.25 | 2 | 2.498 | -4 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | V4 | 0 | 3 | 0.9984 | 0.7566 | 0.9702 | 3.471 | 4.298 | 0.1183 | 0.7202 | 0.9177 | 0.9146 | 1 | 6 | 2.476 | 4 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | II | 1 | 0 | 0.9984 | 0.8811 | 0.9541 | 6.472 | 10.116 | 0.0273 | 0.8014 | 0.8632 | 0.8568 | 0.25 | 0 | 2.494 | 4 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | aVL | 1 | 1 | 0.9976 | 0.9734 | 0.9934 | 13.183 | 18.241 | 0.038 | 0.9387 | 0.8947 | 0.8926 | -0.25 | -6 | 2.5 | 4 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | V2 | 1 | 2 | 0.9912 | 0.9871 | 0.9957 | 15.671 | 19.808 | 0.0372 | 0.9639 | 0.9542 | 0.9543 | 0 | 4 | 2.496 | -4 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | V5 | 1 | 3 | 0.9936 | 0.8298 | 0.98 | 5.052 | 6.535 | 0.0727 | 0.7863 | 0.8637 | 0.8582 | 0.75 | 4 | 2.482 | 4 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | III | 2 | 0 | 0.9936 | 0.9579 | 0.9964 | 11.245 | 16.958 | 0.0508 | 0.926 | 0.9517 | 0.9485 | -0.25 | -8 | 2.498 | 4 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | aVF | 2 | 1 | 0.9992 | 0.9842 | 0.9869 | 15.187 | 15.93 | 0.0147 | 0.983 | 0.9645 | 0.9612 | -0.25 | -4 | 2.5 | 2 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | V3 | 2 | 2 | 0.996 | 0.954 | 0.989 | 10.806 | 16.754 | 0.0666 | 0.9199 | 0.9245 | 0.9244 | 0 | 4 | 2.496 | -4 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | V6 | 2 | 3 | 0.9968 | 0.9744 | 0.9878 | 12.842 | 12.842 | 0.0202 | 0.8958 | 0.8893 | 0.872 | 0.25 | 4 | 2.494 | 0 | ok | ok
C (robustness) | 1737 | MI | C_tilt_20261014_124720 | learned | eacg_sheet_ptbxl1737_kit_500hz_300dpi_model_20261014_124720.csv | II rhythm | 3 | 0 | 0.9976 | 0.8937 | 0.8937 | 6.916 | 6.916 | 0.0262 | 0.8626 | 0.885 | 0.885 | 0 | 0 | 10 | 0 | ok | ok
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | I | 0 | 0 | 0.8664 | 0.2489 | 0.3995 | 0.099 | 0.38 | 0.2116 | 0.0999 | -0.0381 | 0.1306 | -5.5 | 66 | 2.434 | -184 | UNRELIABLE | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | aVR | 0 | 1 | 0.8664 | 0.069 | 0.147 | -1.507 | -1.342 | 0.2004 | 0.0505 | -0.0109 | -0.0385 | 4.25 | 100 | 2.398 | -42 | UNRELIABLE | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | V1 | 0 | 2 | 0.9632 | 0.3215 | 0.3335 | -1.343 | -1.325 | 0.4354 | 0.2965 | 0.4751 | 0.4734 | -2 | -14 | 2.5 | -6 | UNRELIABLE | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | V4 | 0 | 3 | 0.9248 | 0.073 | 0.1229 | -5.25 | -5.06 | 0.4983 | 0.0883 | 0.3596 | 0.3584 | -2.75 | -58 | 2.5 | 24 | UNRELIABLE | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | II | 1 | 0 | 0.9744 | 0.2934 | 0.306 | -5.682 | -5.578 | 0.2385 | 0.515 | -0.2541 | 0.0367 | -2.5 | -100 | 2.464 | -6 | UNRELIABLE | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | aVL | 1 | 1 | 0.9264 | 0.3928 | 0.5183 | 0.417 | 0.534 | 0.1962 | 0.1368 | 0.0409 | 0.0594 | 5.5 | 100 | 2.37 | -22 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | V2 | 1 | 2 | 0.952 | 0.4977 | 0.5192 | 1.024 | 1.178 | 0.4068 | 0.3274 | 0.0144 | 0.0252 | 0.25 | 22 | 2.478 | -6 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | V5 | 1 | 3 | 0.98 | 0.8894 | 0.9213 | 6.769 | 8.223 | 0.0522 | 0.8642 | 0.9397 | 0.9533 | 0.25 | -4 | 2.49 | 8 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | III | 2 | 0 | 0.9392 | 0.4856 | 0.5868 | 1.089 | 1.112 | 0.1566 | 0.3311 | -0.0191 | 0.3479 | 7.5 | 60 | 2.326 | 2 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | aVF | 2 | 1 | 0.964 | 0.5216 | 0.6302 | -1.09 | -0.938 | 0.1289 | 0.6733 | 0.868 | 0.8696 | 4 | 72 | 2.404 | -6 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | V3 | 2 | 2 | 0.96 | 0.8874 | 0.9146 | 6.741 | 7.694 | 0.1599 | 0.7223 | 0.8856 | 0.8854 | 0 | 10 | 2.49 | -10 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | V6 | 2 | 3 | 0.972 | 0.7645 | 0.7858 | 2.485 | 2.792 | 0.0696 | 0.8697 | 0.8821 | 0.8773 | -0.5 | 2 | 2.498 | -4 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125336.csv | II rhythm | 3 | 0 | 0.9954 | 0.4088 |  | -0.386 | 0.382 | 0.1403 | 0.3758 | 0.5099 |  | -0.5 | 4272 | 5.728 | -20 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | I | 0 | 0 | 0.9976 | 0.7113 | 0.9755 | 2.6 | 3.94 | 0.1586 | 0.6827 | 0.8382 | 0.9825 | 4.5 | 72 | 2.392 | -20 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | aVR | 0 | 1 | 0.9976 | 0.9449 | 0.9676 | 9.69 | 10.604 | 0.0552 | 0.9422 | 0.943 | 0.9427 | 0.75 | 4 | 2.482 | 4 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | V1 | 0 | 2 | 0.9992 | 0.9893 | 0.9943 | 16.68 | 18.163 | 0.0547 | 0.9642 | 0.983 | 0.9818 | -0.25 | 0 | 2.5 | -4 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | V4 | 0 | 3 | 0.9992 | 0.8914 | 0.959 | 6.77 | 10.43 | 0.1249 | 0.8873 | 0.9914 | 0.9906 | -0.5 | -20 | 2.492 | 14 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | II | 1 | 0 | 0.9984 | 0.6871 | 0.9425 | 2.214 | 6.192 | 0.0961 | 0.6654 | 0.9555 | 0.948 | 1.25 | 0 | 2.47 | 16 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | aVL | 1 | 1 | 0.9992 | 0.9879 | 0.9954 | 16.062 | 19.614 | 0.0324 | 0.9468 | 0.9609 | 0.9603 | 0 | -4 | 2.496 | 4 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | V2 | 1 | 2 | 0.9984 | 0.9873 | 0.9971 | 16.023 | 17.451 | 0.0723 | 0.9761 | 0.9858 | 0.9849 | 0.5 | 2 | 2.488 | 4 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | V5 | 1 | 3 | 0.9968 | 0.8769 | 0.9219 | 6.23 | 8.211 | 0.0555 | 0.8703 | 0.9986 | 0.9984 | 0 | -8 | 2.492 | 8 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | III | 2 | 0 | 0.9912 | 0.5255 | 0.95 | 0.363 | 6.034 | 0.1703 | 0.5351 | 0.9423 | 0.9519 | -2.25 | -60 | 2.498 | 40 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | aVF | 2 | 1 | 1 | 0.9543 | 0.9665 | 10.554 | 11.857 | 0.0337 | 0.9174 | 0.9554 | 0.9553 | 0 | 2 | 2.498 | -2 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | V3 | 2 | 2 | 0.9984 | 0.9744 | 0.9933 | 13.135 | 18.595 | 0.0766 | 0.9801 | 0.9889 | 0.9879 | -0.25 | 4 | 2.496 | -8 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | V6 | 2 | 3 | 0.9912 | 0.9187 | 0.9292 | 7.528 | 7.936 | 0.0389 | 0.9743 | 0.9595 | 0.9428 | -0.25 | 0 | 2.5 | -2 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125332 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125332.csv | II rhythm | 3 | 0 | 0.995 | 0.4935 |  | -0.132 | 0.146 | 0.1362 | 0.502 | 1.0331 |  | -0.5 | 4264 | 5.736 | -10 | ok | check | NO_PULSE/warn
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | I | 0 | 0 | 0.8456 | 0.2617 | 0.7337 | -0.809 | 1.353 | 0.2349 | 0.209 | 0.0322 | 0.2838 | 2.75 | -52 | 2.382 |  | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | aVR | 0 | 1 | 0.5328 | 0.0449 | 0.4784 | -5.288 | -3.172 | 0.3097 | 0.0683 | 0 | 0.4671 | -1.25 | 74 | 2.426 |  | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | V1 | 0 | 2 | 0.9336 | -0.2346 | 0.328 | -2.736 | -2.538 | 0.5111 | -0.1735 | -0.1738 | 0.0601 | -10 | 100 | 2.4 | -200 | FAILED | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | V4 | 0 | 3 | 0.9808 | 0.0252 | 0.0579 | -0.866 | -0.824 | 0.3008 | 0.0128 | 0.0211 | 0.038 | -6 | -74 | 2.5 | 26 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | II | 1 | 0 | 0.9816 | -0.1602 | 0.3019 | -7.54 | -6.054 | 0.2954 | -0.3183 | 0.1631 | 0.3479 | 3.5 | -100 | 2.316 | 136 | ok | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | aVL | 1 | 1 | 0.9616 | 0.1021 | 0.546 | -5.446 | -4.391 | 0.3854 | 0.1751 | 0.7251 | 0.805 | -7 | -34 | 2.5 | -38 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | V2 | 1 | 2 | 0.8624 | -0.1511 | 0.3765 | -2.171 | -2.032 | 0.5876 | -0.0974 | -0.0896 | 0.0104 | -10 | 100 | 2.4 | -200 | FAILED | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | V5 | 1 | 3 | 0.9808 | 0.2212 | 0.2487 | -7.465 | -7.366 | 0.2686 | 0.458 | 0 | 0.005 | 2 | -42 | 2.408 | 60 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | III | 2 | 0 | 0.8824 | 0.0257 | 0.571 | -2.13 | -0.842 | 0.2269 | 0.0224 | 0.1497 | 0.6687 | 0.25 | -100 | 2.394 | 158 | ok | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | aVF | 2 | 1 | 0.9648 | 0.0108 | 0.9551 | -4.885 | -1.227 | 0.1995 | 0.016 | 0.3843 | 0.8405 | -7.5 | -36 | 2.5 | -94 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | V3 | 2 | 2 | 0.9664 | -0.0588 | 0.7298 | -2.39 | -2.277 | 0.4574 | -0.0503 | -0.0701 | 0.9131 | -10 | 100 | 2.4 | -200 | FAILED | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | V6 | 2 | 3 | 0.9792 | 0.2513 | 0.2776 | -1.23 | -1.156 | 0.1068 | 0.219 | 0.0628 | 0.044 | -6.5 | -32 | 2.5 | 50 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125433.csv | II rhythm | 3 | 0 | 0.9836 | 0.0844 |  | -2.308 | -2.308 | 0.175 | 0.0782 | 0.1507 |  | -6.25 | 1526 | 8.474 | 0 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | I | 0 | 0 | 0.9552 | 0.0328 | 0.9519 | -2.453 | 1.595 | 0.2838 | 0.0301 | 0.0315 | 0.8836 | -1.5 | -100 | 2.438 | 80 | ok | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | aVR | 0 | 1 | 0.9976 | 0.302 | 0.9693 | -2.885 | -2.17 | 0.2349 | 0.4021 | 0.7083 | 0.8935 | -6.75 | -32 | 2.5 | -82 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | V1 | 0 | 2 | 0.9992 | -0.1944 | 0.512 | -3 | -2.83 | 0.5269 | -0.1613 | -0.1679 | 0.8008 | -10 | 100 | 2.4 | -200 | FAILED | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | V4 | 0 | 3 | 0.812 | 0.015 | 0.0385 | -1.776 | -1.72 | 0.334 | 0.0112 | 0.0257 | 0.0354 | -5.25 | -82 | 2.5 |  | UNRELIABLE | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | II | 1 | 0 | 0.9536 | -0.1599 | 0.2793 | -7.816 | -6.556 | 0.305 | -0.3312 | 0.1743 | 0.7138 | -0.25 | -100 | 2.406 | 138 | ok | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | aVL | 1 | 1 | 0.9976 | 0.2177 | 0.9972 | -3.523 | -1.047 | 0.3089 | 0.2981 | 0.6408 | 0.8912 | -7.25 | -36 | 2.5 | -80 | ok | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | V2 | 1 | 2 | 0.9992 | -0.2393 | 0.6897 | -2.862 | -2.662 | 0.6363 | -0.1831 | -0.139 | 0.8973 | -10 | 100 | 2.4 | -200 | FAILED | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | V5 | 1 | 3 | 0.9192 | 0.1225 | 0.1593 | -1.419 | -1.259 | 0.1339 | 0.096 | 0.0227 | 0.0416 | -6.5 | -88 | 2.5 | 42 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | III | 2 | 0 | 0.916 | -0.003 | 0.5075 | -2.832 | -1.683 | 0.246 | -0.0031 | 0.1672 | 0.7977 | 0.25 | -100 | 2.394 | 166 | ok | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | aVF | 2 | 1 | 0.9976 | 0.0123 | 0.9587 | -4.524 | -1.254 | 0.1913 | 0.0171 | 0.409 | 0.8856 | -7.5 | -38 | 2.5 | -96 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | V3 | 2 | 2 | 0.9992 | -0.0623 | 0.7513 | -2.402 | -2.296 | 0.458 | -0.0534 | -0.0719 | 0.9262 | -10 | 100 | 2.4 | -200 | FAILED | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | V6 | 2 | 3 | 0.9952 | 0.3281 | 0.3689 | -6.58 | -6.047 | 0.1977 | 0.6739 | 0.0569 | 0.0448 | -6.75 | -36 | 2.5 | -188 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125428 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125428.csv | II rhythm | 3 | 0 | 0.9328 | 0.0774 |  | -2.549 | -2.542 | 0.1799 | 0.0754 | 0.1289 |  | -6.25 | 1524 | 8.476 | 2 | UNRELIABLE | unreliable | NO_TICK_SCALE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | I | 0 | 0 | 0.8424 | 0.8368 | 0.9625 | 5.088 | 5.506 | 0.1191 | 0.7912 | 0.8823 | 0.9226 | 3 | 42 | 2.428 |  | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | aVR | 0 | 1 | 0.6712 | 0.7185 | 0.7534 | 2.919 | 3.046 | 0.1204 | 0.4524 | 0.4559 | 0.6562 | 4.25 | 60 | 2.398 |  | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | V1 | 0 | 2 | 0.9632 | 0.982 | 0.995 | 14.417 | 17.751 | 0.0709 | 0.9426 | 0.972 | 0.9708 | -0.5 | 0 | 2.5 | -6 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | V4 | 0 | 3 | 0.9808 | 0.8747 | 0.883 | 6.354 | 6.354 | 0.131 | 0.7363 | 0.9687 | 0.9687 | -1 | -14 | 2.5 | 0 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | II | 1 | 0 | 0.9784 | 0.3228 | 0.9326 | -1.387 | 8.879 | 0.1455 | 0.3312 | 0.5565 | 0.8729 | 0 | -34 | 2.466 | 34 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | aVL | 1 | 1 | 0.9624 | 0.9545 | 0.9694 | 10.283 | 10.427 | 0.063 | 0.9904 | 0.906 | 0.9047 | 1 | 12 | 2.476 | 0 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | V2 | 1 | 2 | 0.9632 | 0.9912 | 0.9957 | 17.336 | 17.336 | 0.0622 | 0.953 | 0.9654 | 0.9634 | 0.5 | 4 | 2.488 | 0 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | V5 | 1 | 3 | 0.9808 | 0.8977 | 0.9178 | 7.154 | 7.154 | 0.0499 | 0.8574 | 0.9515 | 0.9515 | -0.75 | -8 | 2.5 | 0 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | III | 2 | 0 | 0.9304 | 0.4571 | 0.9358 | 0.329 | 4.235 | 0.171 | 0.4051 | 0.7811 | 0.9261 | -3.5 | -96 | 2.494 | 68 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | aVF | 2 | 1 | 0.964 | 0.8974 | 0.9476 | 7.128 | 9.887 | 0.05 | 0.8377 | 0.8966 | 0.8966 | 0 | 6 | 2.494 | -6 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | V3 | 2 | 2 | 0.9576 | 0.958 | 0.9932 | 11.047 | 17.928 | 0.0974 | 0.9484 | 0.9729 | 0.9729 | -0.25 | 8 | 2.492 | -10 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | V6 | 2 | 3 | 0.9656 | 0.7517 | 0.8601 | 3.243 | 5.164 | 0.0638 | 0.7094 | 0.8839 | 0.8792 | -1 | 4 | 2.496 | -10 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261015_125643.csv | II rhythm | 3 | 0 | 0.9914 | 0.4966 |  | 0.043 | 0.054 | 0.1335 | 0.4884 | 0.976 |  | -4 | 6776 | 3.224 | 2 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | I | 0 | 0 | 0.9984 | 0.8483 | 0.9639 | 5.378 | 5.655 | 0.1152 | 0.8164 | 0.9475 | 0.9782 | 3 | 42 | 2.428 | -6 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | aVR | 0 | 1 | 0.9976 | 0.9559 | 0.9849 | 10.74 | 10.91 | 0.0489 | 0.9051 | 0.9507 | 0.9476 | 1 | 8 | 2.476 | 2 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | V1 | 0 | 2 | 0.9992 | 0.9858 | 0.9942 | 15.477 | 17.651 | 0.0628 | 0.9557 | 0.9791 | 0.978 | -0.25 | 2 | 2.498 | -4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | V4 | 0 | 3 | 0.9992 | 0.9302 | 0.9633 | 8.709 | 9.526 | 0.0999 | 0.9214 | 0.9877 | 0.9868 | -1 | -16 | 2.5 | 6 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | II | 1 | 0 | 0.9984 | 0.296 | 0.9324 | -1.645 | 8.824 | 0.1499 | 0.3097 | 0.4705 | 0.9495 | 0 | -34 | 2.466 | 34 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | aVL | 1 | 1 | 0.9992 | 0.9944 | 0.9973 | 19.417 | 22.081 | 0.022 | 0.967 | 0.9526 | 0.949 | 0.25 | 0 | 2.494 | 2 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | V2 | 1 | 2 | 1 | 0.9898 | 0.9966 | 16.937 | 18 | 0.0651 | 0.9708 | 0.978 | 0.9778 | 0.5 | 2 | 2.488 | 2 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | V5 | 1 | 3 | 0.988 | 0.8965 | 0.918 | 7.032 | 7.088 | 0.0506 | 0.8777 | 0.9944 | 0.9923 | -0.75 | -8 | 2.5 | 2 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | III | 2 | 0 | 0.9664 | 0.2883 | 0.9351 | -1.111 | 3.122 | 0.2018 | 0.2747 | 0.8663 | 0.962 | -3.5 | -98 | 2.492 | 68 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | aVF | 2 | 1 | 1 | 0.926 | 0.9573 | 8.482 | 10.839 | 0.0428 | 0.9057 | 0.9821 | 0.9821 | 0 | 4 | 2.496 | -4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | V3 | 2 | 2 | 0.9976 | 0.9658 | 0.9925 | 11.907 | 17.349 | 0.0882 | 0.9611 | 0.9801 | 0.9789 | -0.25 | 6 | 2.494 | -8 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | V6 | 2 | 3 | 0.9784 | 0.8083 | 0.9326 | 3.943 | 6.512 | 0.0589 | 0.8456 | 0.9629 | 0.956 | -1 | 2 | 2.498 | -10 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_crumple_20261015_125638 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261015_125638.csv | II rhythm | 3 | 0 | 0.9902 | 0.4656 |  | -0.392 | -0.306 | 0.1404 | 0.4763 | 1.0264 |  | -3.75 | 6780 | 3.22 | 4 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | I | 0 | 0 | 0.98 | 0.5556 | 0.5767 | -1.56 | -1.433 | 0.2561 | 0.7105 | 0.8588 | 0.8586 | 1 | 6 | 2.476 | 6 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | aVR | 0 | 1 | 0.6928 | 0.282 | 0.2986 | -3.981 | -3.953 | 0.2664 | 0.3772 | 0.369 | 0.3663 | 1 | 2 | 2.476 |  | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | V1 | 0 | 2 | 0.964 | 0.3826 | 0.4085 | -1.729 | -1.593 | 0.4552 | 0.3674 | 0.1389 | 0.1389 | -2.25 | -26 | 2.5 | -12 | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | V4 | 0 | 3 | 0.4184 | 0.4388 | 0.4794 | -0.371 | -0.366 | 0.2841 | 0.3017 | 0 | 0 | 7.75 | 18 | 2.32 |  | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | II | 1 | 0 | 0.9776 | 0.6581 | 0.9438 | 1.577 | 9.364 | 0.1034 | 0.6721 | 0.8571 | 0.8536 | 0.25 | -14 | 2.48 | 16 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | aVL | 1 | 1 | 0.9624 | 0.5419 | 0.5637 | -3.29 | -3.253 | 0.3007 | 0.8101 | 0.8733 | 0.8704 | 1.75 | 30 | 2.458 | -4 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | V2 | 1 | 2 | 0.7216 | 0.3346 | 0.3372 | -2.284 | -2.274 | 0.5953 | 0.3661 | 0.5335 | 0.5323 | 0.5 | 6 | 2.488 |  | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | V5 | 1 | 3 | 0.48 | 0.5802 | 0.6068 | 1.705 | 1.875 | 0.0935 | 0.3065 | 0.0679 | 0.0703 | -1.5 | 4 | 2.496 |  | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | III | 2 | 0 | 0.98 | 0.9569 | 0.9678 | 10.811 | 11.557 | 0.0511 | 0.8897 | 0.8597 | 0.8582 | -0.5 | -2 | 2.5 | -4 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | aVF | 2 | 1 | 0.9624 | 0.6948 | 0.7204 | 0.98 | 1.125 | 0.1015 | 0.8578 | 0.8491 | 0.8483 | 1.25 | 16 | 2.47 | 4 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | V3 | 2 | 2 | 0.964 | 0.9844 | 0.9953 | 15.304 | 19.683 | 0.0596 | 0.949 | 0.9697 | 0.9697 | 0.25 | 8 | 2.492 | -6 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | V6 | 2 | 3 | 0.48 | 0.4374 | 0.4951 | 0.735 | 0.979 | 0.0851 | 0.2676 | 0 | 0.6016 | 1.5 | 14 | 2.464 |  | UNRELIABLE | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_231734.csv | II rhythm | 3 | 0 | 0.8988 | 0.5824 | 0.8313 | 1.202 | 1.325 | 0.1168 | 0.5138 | 0.8221 | 0.835 | 1 | 56 | 9.9 | -8 | ok | unreliable | SCALE_DISAGREE/severe; LEAD_COVERAGE/warn/aVR; LEAD_COVERAGE/warn/V4; LEAD_COVERAGE/warn/V2; LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | I | 0 | 0 | 0.9984 | 0.9321 | 0.9864 | 8.827 | 14.513 | 0.0775 | 0.8974 | 0.9709 | 0.9425 | 0.5 | -4 | 2.484 | 12 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | aVR | 0 | 1 | 0.9976 | 0.9838 | 0.9969 | 14.966 | 21.434 | 0.0301 | 0.9473 | 0.9621 | 0.9622 | 0 | -4 | 2.496 | 4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | V1 | 0 | 2 | 0.9984 | 0.9863 | 0.9921 | 15.582 | 17.62 | 0.062 | 0.9688 | 0.9772 | 0.9761 | -0.25 | 2 | 2.498 | -4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | V4 | 0 | 3 | 0.9928 | 0.9746 | 0.9965 | 13.12 | 19.163 | 0.0601 | 0.9539 | 0.9931 | 0.9923 | -0.25 | 4 | 2.496 | -6 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | II | 1 | 0 | 0.9976 | 0.6152 | 0.8652 | -0.525 | 1.507 | 0.1317 | 0.8216 | 0.9513 | 0.9543 | -2.25 | -52 | 2.5 | 10 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | aVL | 1 | 1 | 0.9976 | 0.9913 | 0.9975 | 17.581 | 22.4 | 0.0272 | 0.9685 | 0.9479 | 0.9476 | 0 | -4 | 2.496 | 4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | V2 | 1 | 2 | 0.9984 | 0.9941 | 0.9988 | 19.222 | 24.463 | 0.05 | 0.9707 | 0.9693 | 0.9687 | -0.25 | 2 | 2.498 | -4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | V5 | 1 | 3 | 0.9992 | 0.9691 | 0.9931 | 12.267 | 18.618 | 0.0277 | 0.9406 | 0.9801 | 0.9799 | 0 | 6 | 2.494 | -6 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | III | 2 | 0 | 0.9984 | 0.8712 | 0.9246 | 5.734 | 5.883 | 0.0918 | 0.9329 | 0.9868 | 0.9831 | -1 | -10 | 2.5 | 2 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | aVF | 2 | 1 | 0.9976 | 0.9586 | 0.9801 | 10.991 | 14.135 | 0.0321 | 0.9392 | 0.9426 | 0.9426 | 0 | -4 | 2.496 | 4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | V3 | 2 | 2 | 0.9984 | 0.9932 | 0.999 | 18.865 | 26.671 | 0.0396 | 0.9735 | 0.9781 | 0.9778 | 0 | 4 | 2.496 | -4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | V6 | 2 | 3 | 0.9992 | 0.9738 | 0.9929 | 12.66 | 16.776 | 0.0216 | 0.9193 | 0.8878 | 0.8853 | 0.25 | 6 | 2.494 | -4 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_231732 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_231732.csv | II rhythm | 3 | 0 | 0.9996 | 0.6662 | 0.9378 | 1.788 | 1.827 | 0.1092 | 0.6608 | 0.9532 | 0.9586 | 0.75 | 48 | 9.926 | -2 | ok | unreliable | SCALE_DISAGREE/severe
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | I | 0 | 0 | 0.0416 | 0.1021 | 0.1134 | 0.047 | 0.049 | 0.2128 | 0.008 | 0 | 0 | 2.25 | 48 | 2.444 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | aVR | 0 | 1 | 0.8464 | 0.7169 | 0.8267 | 3.111 | 3.13 | 0.1178 | 0.4881 | 0.4089 | 0.6667 | 4.25 | 68 | 2.398 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | V1 | 0 | 2 | 0.96 | 0.7662 | 0.7816 | 3.794 | 4.066 | 0.241 | 0.6528 | 0.9085 | 0.907 | 0.25 | 10 | 2.49 | -8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | V4 | 0 | 3 | 0.9784 | 0.3207 | 0.3425 | 0.377 | 0.442 | 0.2607 | 0.1632 | 0.2785 | 0.2763 | 0.75 | 0 | 2.482 | 8 | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | II | 1 | 0 | 0.12 | 0.2764 | 0.294 | 0.314 | 0.319 | 0.1196 | 0.0467 | 0 | 0 | 2.75 | 60 | 2.434 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | aVL | 1 | 1 | 0.9552 | 0.5009 | 0.5254 | -3.496 | -3.477 | 0.3079 | 0.7538 | 0.8622 | 0.8611 | 3.5 | 58 | 2.416 | -4 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | V2 | 1 | 2 | 0.9592 | 0.6525 | 0.6813 | 2.176 | 2.267 | 0.3562 | 0.3871 | 0.0282 | 0.0528 | -2.5 | 2 | 2.498 | -8 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | V5 | 1 | 3 | 0.98 | 0.5681 | 0.5896 | 1.484 | 1.522 | 0.0959 | 0.2517 | 0.1206 | 0.1441 | 1.75 | 10 | 2.458 | 12 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | III | 2 | 0 | 0 | [12 empty] | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | aVF | 2 | 1 | 0.6096 | 0.3851 | 0.3955 | -5.756 | -5.749 | 0.2205 | 0.6995 | 0.8045 | 0.8031 | 1.25 | 22 | 2.47 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | V3 | 2 | 2 | 0.9568 | 0.5498 | 0.5899 | 1.644 | 1.949 | 0.2874 | 0.3473 | 0.1739 | 0.2035 | -1 | 16 | 2.484 | -20 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | V6 | 2 | 3 | 0.9784 | 0.5576 | 0.6093 | 1.318 | 1.318 | 0.0796 | 0.4349 | 0.1463 | 0.1354 | 3.25 | 44 | 2.422 | 0 | ok | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_232544.csv | II rhythm | 3 | 0 | 0.849 | 0.3139 |  | -0.602 | -0.485 | 0.1438 | 0.2534 | 0.2084 |  | 1.5 | -1616 | 8.236 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/I; LEAD_COVERAGE/warn/II; LEAD_COVERAGE/warn/III; LEAD_COVERAGE/warn/aVF
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | I | 0 | 0 | 0.9984 | 0.9426 | 0.9921 | 9.534 | 17.483 | 0.0714 | 0.9215 | 0.9709 | 0.9665 | 0.25 | -8 | 2.486 | 10 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | aVR | 0 | 1 | 0.9968 | 0.9904 | 0.9967 | 17.085 | 21.101 | 0.0236 | 0.9544 | 0.9694 | 0.9694 | 0 | -2 | 2.498 | 2 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | V1 | 0 | 2 | 0.9984 | 0.9892 | 0.9987 | 16.631 | 21.976 | 0.055 | 0.9671 | 0.9803 | 0.9801 | -0.25 | 2 | 2.498 | -4 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | V4 | 0 | 3 | 0.9904 | 0.9805 | 0.9954 | 14.247 | 16.473 | 0.0528 | 0.9602 | 0.9913 | 0.9902 | 0.5 | 0 | 2.488 | 4 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | II | 1 | 0 | 0.9928 | 0.5153 | 0.9547 | -0.698 | 8.541 | 0.1344 | 0.6108 | 0.9253 | 0.9658 | -0.5 | -26 | 2.486 | 20 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | aVL | 1 | 1 | 0.9968 | 0.9948 | 0.998 | 19.402 | 22.156 | 0.0221 | 0.9736 | 0.9491 | 0.9426 | 0.25 | 0 | 2.494 | 2 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | V2 | 1 | 2 | 0.9912 | 0.9854 | 0.9897 | 14.185 | 15.11 | 0.0894 | 0.8819 | 0.9203 | 0.9195 | -0.25 | 2 | 2.498 | -4 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | V5 | 1 | 3 | 0.5544 | 0.6748 | 0.6977 | 1.344 | 1.534 | 0.0974 | 0.6258 | 0.801 | 0.7997 | 0.75 | 12 | 2.482 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | III | 2 | 0 | 0.9776 | 0.91 | 0.9761 | 6.849 | 11.592 | 0.0807 | 1.0147 | 0.9647 | 0.959 | -0.5 | -16 | 2.496 | 10 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | aVF | 2 | 1 | 0.912 | -0.0047 | 0.1901 | -13.874 | -13.71 | 0.5615 | -0.0224 | 1.4207 | 1.5961 | -4.75 | -6 | 2.5 | -44 | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | V3 | 2 | 2 | 0.7688 | 0.1927 | 0.2008 | -3.224 | -3.208 | 0.5035 | 0.2453 | 0.8152 | 0.8136 | -1.5 | -18 | 2.5 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | V6 | 2 | 3 | 0.7656 | 0.1917 | 0.2585 | -6.49 | -6.194 | 0.1956 | 0.3944 | 0.6144 | 0.6221 | 0 | 12 | 2.488 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_shadow_20261014_232541 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_232541.csv | II rhythm | 3 | 0 | 0.942 | 0.5644 | 0.5991 | -0.369 | -0.029 | 0.14 | 0.6789 | 0.6584 | 0.6584 | 0 | 8 | 9.992 | -8 | ok | check | LEAD_COVERAGE/warn/V5; LEAD_COVERAGE/warn/V3; LEAD_COVERAGE/warn/V6
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | I | 0 | 0 | 0.976 | 0.1025 | 0.1527 | -2.146 | -1.992 | 0.274 | 0.0924 | -0.0323 | 0.1971 | 7 | 52 | 2.336 | -4 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | aVR | 0 | 1 | 0.9448 | 0.4733 | 0.4916 | 1.02 | 1.038 | 0.1498 | 0.2231 | -0.0182 | -0.0182 | -2.5 | -32 | 2.5 | 2 | ok | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | V1 | 0 | 2 | 0.9608 | 0.3488 | 0.4634 | 0.369 | 0.525 | 0.3575 | 0.1012 | -0.0027 | 0.0314 | 5.25 | 100 | 2.376 | -28 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | V4 | 0 | 3 | 0.4176 | 0.0373 | 0.0718 | -5.57 | -5.371 | 0.5169 | 0.0448 | -0.0129 | 0 | 5.75 | 0 | 2.364 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | II | 1 | 0 | 0.9808 | 0.7725 | 0.7868 | 3.606 | 3.65 | 0.0819 | 0.4482 | 0.3251 | 0.325 | -0.5 | -6 | 2.5 | 4 | ok | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | aVL | 1 | 1 | 0.9264 | 0.2839 | 0.5471 | -0.016 | 0.412 | 0.2063 | 0.107 | 0.0289 | 0.1684 | -2.5 | 100 | 2.4 | -134 | ok | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | V2 | 1 | 2 | 0.9616 | 0.4719 | 0.5033 | 0.547 | 0.597 | 0.4297 | 0.1485 | 0.0118 | 0.0122 | 2.25 | 38 | 2.444 | -18 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | V5 | 1 | 3 | 0.98 | 0.4303 | 0.5036 | -4.126 | -3.739 | 0.1829 | 0.6966 | 0.907 | 0.9054 | 1 | -6 | 2.47 | 22 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | III | 2 | 0 | 0.9776 | 0.4104 | 0.4738 | 0.611 | 0.829 | 0.1655 | 0.1321 | 0.0386 | 0.0603 | -5.5 | -30 | 2.5 | -62 | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | aVF | 2 | 1 | 0.964 | 0.4308 | 0.5599 | 0.828 | 0.962 | 0.1033 | 0.2135 | 0.0388 | 0.2518 | 3.25 | 74 | 2.422 | -40 | ok | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | V3 | 2 | 2 | 0.9616 | 0.453 | 0.5208 | 0.901 | 1.103 | 0.3131 | 0.2008 | 0.0608 | 0.0875 | 0.25 | 100 | 2.4 | -108 | ok | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | V6 | 2 | 3 | 0.9384 | 0.7413 | 0.9518 | 2.451 | 3.185 | 0.0699 | 0.7996 | 0.7931 | 0.9248 | 2 | 10 | 2.45 | 8 | ok | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | classical | eacg_sheet_ptbxl4016_kit_500hz_300dpi_classical_20261014_125029.csv | II rhythm | 3 | 0 | 0.8366 | 0.4221 | 0.6171 | 0.033 | 0.033 | 0.1337 | 0.3512 | -0.0697 | 0.1236 | 1.75 | -734 | 9.094 |  | UNRELIABLE | check | LEAD_COVERAGE/warn/V4
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | I | 0 | 0 | 0.9992 | 0.9635 | 0.9755 | 11.427 | 11.427 | 0.0574 | 0.9062 | 0.9464 | 0.9425 | 0.75 | 8 | 2.482 | 2 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | aVR | 0 | 1 | 0.9976 | 0.9645 | 0.9762 | 11.646 | 13.376 | 0.0441 | 0.9416 | 0.9375 | 0.9376 | 0 | -4 | 2.496 | 4 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | V1 | 0 | 2 | 0.9976 | 0.988 | 0.9958 | 16.199 | 19.655 | 0.0578 | 0.9713 | 0.9857 | 0.9857 | -0.25 | 2 | 2.498 | -4 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | V4 | 0 | 3 | 0.996 | 0.9191 | 0.9939 | 8.03 | 10.619 | 0.108 | 0.9194 | 1.0081 | 1.0073 | 1.5 | 4 | 2.464 | 10 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | II | 1 | 0 | 0.9992 | 0.9333 | 0.9697 | 8.879 | 8.879 | 0.0446 | 0.8667 | 0.9344 | 0.9246 | 0.75 | 8 | 2.482 | 0 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | aVL | 1 | 1 | 0.9976 | 0.9894 | 0.9921 | 15.71 | 16.588 | 0.0337 | 0.9179 | 0.9172 | 0.9172 | 0 | -2 | 2.498 | 2 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | V2 | 1 | 2 | 0.9968 | 0.9863 | 0.9974 | 15.7 | 20.62 | 0.0751 | 0.9705 | 0.977 | 0.9765 | -0.25 | 2 | 2.498 | -6 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | V5 | 1 | 3 | 0.9992 | 0.8696 | 0.9834 | 5.917 | 7.069 | 0.0575 | 0.8738 | 1.0004 | 0.9946 | 1.5 | 6 | 2.464 | 8 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | III | 2 | 0 | 0.9912 | 0.9894 | 0.9952 | 16.862 | 19.294 | 0.0255 | 0.9551 | 0.9668 | 0.9634 | -0.25 | -6 | 2.5 | 2 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | aVF | 2 | 1 | 0.9992 | 0.9713 | 0.9771 | 12.488 | 13.37 | 0.027 | 0.9065 | 0.932 | 0.932 | 0 | -2 | 2.498 | 2 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | V3 | 2 | 2 | 0.9968 | 0.9896 | 0.9972 | 17.05 | 21.334 | 0.0488 | 0.9845 | 0.9854 | 0.985 | -0.25 | 2 | 2.498 | -4 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | V6 | 2 | 3 | 0.9992 | 0.716 | 0.9768 | 2.237 | 3.757 | 0.0716 | 0.7454 | 0.8788 | 0.9278 | 1.75 | 6 | 2.458 | 10 | ok | ok
C (robustness) | 4016 | CD | C_tilt_20261014_125027 | learned | eacg_sheet_ptbxl4016_kit_500hz_300dpi_model_20261014_125027.csv | II rhythm | 3 | 0 | 0.9998 | 0.8723 | 0.9113 | 5.938 | 5.938 | 0.0677 | 0.8703 | 0.9661 | 0.9862 | -0.25 | 844 | 9.156 | 0 | ok | ok